IAEA-TEL-2021-03 World Wide Open Proficiency Test Exercise, Pie-charts, S-Shapes and Reported Results with Scores

## A. Mauring1, S. Patterson1, B. Seslak1, S. Tarjan1, and A. Trinkl1

1Terrestrial Environment Laboratory, NA Environment Laboratories November 17, 2021

Contents

1  Introduction	1

1.1  Evaluation criteria	1

| 2 | Data Evaluation |  | 3 |
| --- | --- | --- | --- |
| 3 | Proficiency Test Parameters |  | 3 |
| 4 | Intercomparison Parameters |  | 84 |
| 5 | References |  | 123 |
| 6 | Contributors to drafting and review | Abstract | 124 |

This report contains the results and evaluation for the proficiency test conducted within the IAEA-TEL-2021-03 world wide open proficiency test exercise. The data is evaluated by the Terrestrial Environment Laboratory (TEL) of the NA Environment Laboratories using its standard approach for proficiency test evaluations.

# Introduction

The following figures show the results (displayed in pie charts, characteristic S-shape graphs and tables) for the ALMERA Proficiency Test. In the S-shape graphs, the vertical axis displays the activity in [Bq/kg] and the horizontal axis shows the labcodes for the participating laboratories. The range on the vertical axis is limited to (± 2 x Maximum acceptable relative bias(MARB)). The red line is the target value for the exercise, whereas the green lines are the as- signed uncertainty for each analyte. The results are grouped by samples, then analytes. The data points are coloured in blue, yellow and red, corresponding to a result of "Accepted", "Warning", and "Not Accepted", respectively.

Disclaimer: This report has been generated automatically to provide information on-line to the participants. A more detailed evaluation will be published in the final report. If you find that any information provided in this report might be incorrect please contact us as soon as possible.

## Evaluation criteria

The data is evaluated according to the following steps:

    - The relative bias between the reported and the target value (the best estimation of the true value) is expressed by the following equation:

Biasrelative  =

V a l u e r e p o r t e d − V a l u e t a r g e t V aluetar get

∗ 100%

    - The relative bias is compared to the Maximum Acceptable Relative Bias (MARB) which has been determined for each analyte, considering the radioanalytical methods, the level of radioactivity and the complexity of the analysis.

If the |Biasrelati ve | ≤ M ARB , the result will be "Accepted" for accuracy.

    - Based on fit-for-purpose and good laboratory practice principles, the expanded relative combined uncertainty should cover the relative bias:

「

a utar get t2

P =

a urepor ted t2

+

∗ 100

｜ Atar get

Arepor ted

|Biasrelati ve | ≤ k ∗ P

where k is the coverage factor, for the 99% confidential level, k = 2.58. 4. If the result is between the ±M ARB values, but it is not overlapping the target value within the uncertainty, this equation is used to decide if they are significantly different or not.

- The P value is compared to the MARB also. If both the:

and

P ≤ M ARB

|Biasrelati ve | ≤ k ∗ P

are fulfilled, the reported results will be "Accepted" for precision. The result will be assigned "Not accepted" for precision if either conditions are not fulfilled.

- The final score is assigned according to the detailed evaluation described above. The possible scores are listed below:

- "Accepted" when both, accuracy and precision achieved "Accepted" status

- "Not Accepted" when the accuracy is "Not accepted"

- "Warning" when accuracy is "Accepted" but the precision is "Not accepted"

- As additional information, a z-score parameter is shown in the evaluation tables that is calculated by using the robust standard deviation described in [2]:

V aluerepor ted

V aluetar get

−

z =

robustsd

- If the analyte is included in the proficiency test evaluation scheme, the stated target value is used to calculate the z-score. For those analytes, which are subject of an intercomparison only, the robust mean of the values reported is used instead.

# Data Evaluation

A total of 326 laboratories from 78 different countries successfully reported data in the IAEA-TEL-2021-03 proficiency test exercise.

For each of the analytes evaluated, the target value, target uncertainty and the MARB are listed. A pie chart showing the distribution of accepted, warning, not acceptable and not reported results, the s-shape graph, and a detailed table listing all evaluation results for the analyte (sorted by the laboratory code) are provided.

For analytes evaluated by intercomparison (IC) parameters, the target value table lists the robust mean, robust standard deviation (SD) and MARB. A detailed table listing all evaluation results for the analyte (sorted by the labora- tory code) are provided, where the Z-Scores are calculated using the robust mean as the V aluetar get in the equation listed in 1.1 Evaluation criteria, point 7.

# Proficiency Test Parameters

## Reported Results for Co-60 Determination in Sample 1

TABLE 1. Target values

| Parameter | Value |
| --- | --- |
| Target Value, Bq/kg | 74.2 |
| UNC, Bq/kg | 3.3 |
| MARB, % | 20 |

FIGURE 1. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

FIGURE 2. Reported results and evaluation parameters for Co-60. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

TABLE 2. Evaluation Results for Co-60 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 73.00 | 2.29 | -1.62 | 5.44 | A | A | A | 0.36 |
| 2 | 74.6 | 3.5 | 0.54 | 6.46 | A | A | A | 0.12 |
| 3 | 74.3 | 3.8 | 0.13 | 6.78 | A | A | A | 0.03 |
| 4 | 70.5 | 5.0 | -4.99 | 8.37 | A | A | A | 1.12 |
| 5 | 71.3 | 6.98 | -3.91 | 10.75 | A | A | A | 0.88 |
| 6 | 65.5 | 1.3 | -11.73 | 4.87 | A | A | A | 2.63 |
| 7 | 73.34 | 1.05 | -1.16 | 4.67 | A | A | A | 0.26 |
| 8 | 82.37 | 1.10 | 11.01 | 4.64 | A | A | A | 2.47 |
| 9 | 77 | 9 | 3.77 | 12.51 | A | A | A | 0.85 |
| 10 | 7.15 | 0.97 | -90.36 | 14.28 | N | N | N | 20.26 |
| 11 | 78.1 | 2.4 | 5.26 | 5.41 | A | A | A | 1.18 |
| 12 | 81.5 | 1.6 | 9.84 | 4.86 | A | A | A | 2.21 |
| 13 | 76.03 | 2.70 | 2.47 | 5.69 | A | A | A | 0.55 |
| 14 | 74.8 | 3.4 | 0.81 | 6.36 | A | A | A | 0.18 |
| 15 | 75.2 | 3.7 | 1.35 | 6.63 | A | A | A | 0.30 |
| 16 | 75.8 | 1.3 | 2.16 | 4.77 | A | A | A | 0.48 |
| 17 | 75.86 | 2.50 | 2.24 | 5.54 | A | A | A | 0.50 |
| 18 | 76.5 | 6.2 | 3.10 | 9.24 | A | A | A | 0.69 |
| 19 | 74.4 | 3.6 | 0.27 | 6.57 | A | A | A | 0.06 |
| 20 | 80.6 | 1.6 | 8.63 | 4.87 | A | A | A | 1.93 |
| 21 | 73.1 | 5.2 | -1.48 | 8.39 | A | A | A | 0.33 |
| 22 | 77.9 | 2.0 | 4.99 | 5.14 | A | A | A | 1.12 |
| 25 | 73 | 3 | -1.62 | 6.06 | A | A | A | 0.36 |
| 26 | 83.77 | 4.51 | 12.90 | 6.98 | A | A | A | 2.89 |
| 27 | 70.5 | 7.3 | -4.99 | 11.27 | A | A | A | 1.12 |
| 29 | 80.57 | 2.29 | 8.58 | 5.28 | A | A | A | 1.92 |
| 31 | 77.70 | 2.81 | 4.72 | 5.73 | A | A | A | 1.06 |
| 32 | 77.4 | 8.5 | 4.31 | 11.85 | A | A | A | 0.97 |
| 33 | 71.5 | 3.2 | -3.64 | 6.31 | A | A | A | 0.82 |
| 34 | 75.1 | 3.2 | 1.21 | 6.16 | A | A | A | 0.27 |

TABLE 2. Evaluation Results for Co-60 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 35 | 79.5 | 6.2 | 7.14 | 8.98 | A | A | A | 1.60 |
| 36 | 74.7 | 3.4 | 0.67 | 6.36 | A | A | A | 0.15 |
| 38 | 73.2 | 4.0 | -1.35 | 7.05 | A | A | A | 0.30 |
| 39 | 67.8 | 1.2 | -8.63 | 4.79 | A | A | A | 1.93 |
| 41 | 75.4 | 3.0 | 1.62 | 5.97 | A | A | A | 0.36 |
| 42 | 59.71 | 1.25 | -19.53 | 4.92 | A | N | W | 4.38 |
| 43 | 76.59 | 2.09 | 3.22 | 5.22 | A | A | A | 0.72 |
| 46 | 130.33 | 2.97 | 75.65 | 5.00 | N | N | N | 16.96 |
| 47 | 64.7 | 1.4 | -12.80 | 4.95 | A | N | W | 2.87 |
| 48 | 73.8 | 2.6 | -0.54 | 5.67 | A | A | A | 0.12 |
| 50 | 83.13 | 2.30 | 12.04 | 5.24 | A | A | A | 2.70 |
| 51 | 80 | 5 | 7.82 | 7.67 | A | A | A | 1.75 |
| 52 | 76.1 | 2.7 | 2.56 | 5.69 | A | A | A | 0.57 |
| 53 | 77.1 | 3.5 | 3.91 | 6.36 | A | A | A | 0.88 |
| 55 | 73.6 | 1.7 | -0.81 | 5.01 | A | A | A | 0.18 |
| 56 | 75.26 | 3.56 | 1.43 | 6.49 | A | A | A | 0.32 |
| 59 | 74.7 | 5.2 | 0.67 | 8.26 | A | A | A | 0.15 |
| 61 | 73.72 | 2.09 | -0.65 | 5.27 | A | A | A | 0.15 |
| 63 | 77.15 | 1.87 | 3.98 | 5.07 | A | A | A | 0.89 |
| 66 | 78.9 | 2.3 | 6.33 | 5.32 | A | A | A | 1.42 |
| 68 | 75.4 | 5.9 | 1.62 | 9.00 | A | A | A | 0.36 |
| 69 | 77.21 | 1.44 | 4.06 | 4.82 | A | A | A | 0.91 |
| 70 | 74.44 | 0.68 | 0.32 | 4.54 | A | A | A | 0.07 |
| 72 | 73.302 | 4.040 | -1.21 | 7.08 | A | A | A | 0.27 |
| 74 | 78.9 | 3.3 | 6.33 | 6.11 | A | A | A | 1.42 |
| 76 | 72.00 | 0.44 | -2.96 | 4.49 | A | A | A | 0.66 |
| 77 | 75.2 | 0.9 | 1.35 | 4.61 | A | A | A | 0.30 |
| 78 | 74.8 | 0.89 | 0.81 | 4.60 | A | A | A | 0.18 |
| 80 | 68.9 | 4.9 | -7.14 | 8.39 | A | A | A | 1.60 |
| 81 | 82.4 | 6.0 | 11.05 | 8.53 | A | A | A | 2.48 |
| 84 | 76.29 | 0.87 | 2.82 | 4.59 | A | A | A | 0.63 |
| 85 | 79.69 | 6.48 | 7.40 | 9.27 | A | A | A | 1.66 |
| 86 | 76.4 | 2.85 | 2.96 | 5.80 | A | A | A | 0.66 |
| 87 | 80 | 3 | 7.82 | 5.82 | A | A | A | 1.75 |
| 88 | 73.4 | 5.2 | -1.08 | 8.36 | A | A | A | 0.24 |
| 89 | 208 | 10 | 180.32 | 6.55 | N | N | N | 40.42 |
| 92 | 72.4 | 5.6 | -2.43 | 8.92 | A | A | A | 0.54 |
| 93 | 75.72 | 0.65 | 2.05 | 4.53 | A | A | A | 0.46 |
| 94 | 72.22 | 0.3 | -2.67 | 4.47 | A | A | A | 0.60 |
| 95 | 72.40 | 0.29 | -2.43 | 4.47 | A | A | A | 0.54 |
| 96 | 78.3 | 4.7 | 5.53 | 7.47 | A | A | A | 1.24 |
| 98 | 77.9 | 4.1 | 4.99 | 6.89 | A | A | A | 1.12 |
| 99 | 72 | 2.2 | -2.96 | 5.40 | A | A | A | 0.66 |
| 101 | 76.6 | 3.4 | 3.23 | 6.28 | A | A | A | 0.73 |
| 102 | 78.6 | 2.5 | 5.93 | 5.47 | A | A | A | 1.33 |
| 103 | 75.44 | 1.62 | 1.67 | 4.94 | A | A | A | 0.37 |
| 105 | 76.89 | 1.04 | 3.63 | 4.65 | A | A | A | 0.81 |
| 106 | 76.5 | 0.3 | 3.10 | 4.46 | A | A | A | 0.69 |
| 107 | 89.7 | 1.6 | 20.89 | 4.79 | N | N | N | 4.68 |
| 108 | 75.8646 | 0.6531 | 2.24 | 4.53 | A | A | A | 0.50 |
| 109 | 79.1 | 3.5 | 6.60 | 6.27 | A | A | A | 1.48 |
| 111 | 76.9 | 5.1 | 3.64 | 7.99 | A | A | A | 0.82 |
| 113 | 73.7 | 3.7 | -0.67 | 6.71 | A | A | A | 0.15 |
| 116 | 75.7 | 1.2 | 2.02 | 4.72 | A | A | A | 0.45 |

TABLE 2. Evaluation Results for Co-60 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 118 | 62.5 | 5.0 | -15.77 | 9.15 | A | A | A | 3.53 |
| 120 | 74.6 | 3.1 | 0.54 | 6.09 | A | A | A | 0.12 |
| 121 | 74.0 | 2.8 | -0.27 | 5.84 | A | A | A | 0.06 |
| 123 | 86 | 6 | 15.90 | 8.27 | A | A | A | 3.56 |
| 124 | 74.4 | 1.4 | 0.27 | 4.83 | A | A | A | 0.06 |
| 125 | 75.0 | 2.4 | 1.08 | 5.48 | A | A | A | 0.24 |
| 126 | 66.4 | 2.8 | -10.51 | 6.13 | A | A | A | 2.36 |
| 129 | 75.5 | 3.0 | 1.75 | 5.96 | A | A | A | 0.39 |
| 130 | 76 | 3 | 2.43 | 5.95 | A | A | A | 0.54 |
| 132 | 72.9 | 4.4 | -1.75 | 7.50 | A | A | A | 0.39 |
| 133 | 73.9 | 3.8 | -0.40 | 6.80 | A | A | A | 0.09 |
| 134 | 76.89 | 2.29 | 3.63 | 5.35 | A | A | A | 0.81 |
| 135 | 75.07 | 2.5 | 1.17 | 5.56 | A | A | A | 0.26 |
| 137 | 67.813 | 1.176 | -8.61 | 4.77 | A | A | A | 1.93 |
| 139 | 75.322 | 1.359 | 1.51 | 4.80 | A | A | A | 0.34 |
| 141 | 77.9 | 4.5 | 4.99 | 7.29 | A | A | A | 1.12 |
| 143 | 72.4 | 4.3 | -2.43 | 7.42 | A | A | A | 0.54 |
| 144 | 72.8 | 4.5 | -1.89 | 7.62 | A | A | A | 0.42 |
| 146 | 73.7 | 2.5 | -0.67 | 5.59 | A | A | A | 0.15 |
| 147 | 72.5 | 6.9 | -2.29 | 10.51 | A | A | A | 0.51 |
| 148 | 76.28 | 5.45 | 2.80 | 8.42 | A | A | A | 0.63 |
| 149 | 76.2 | 1.6 | 2.70 | 4.92 | A | A | A | 0.60 |
| 150 | 76.0 | 2.4 | 2.43 | 5.45 | A | A | A | 0.54 |
| 151 | 77.0 | 4.0 | 3.77 | 6.84 | A | A | A | 0.85 |
| 152 | 72.8 | 2.9 | -1.89 | 5.97 | A | A | A | 0.42 |
| 153 | 79.1 | 5.6 | 6.60 | 8.36 | A | A | A | 1.48 |
| 154 | 73.7 | 1.5 | -0.67 | 4.89 | A | A | A | 0.15 |
| 155 | 78.3 | 4.1 | 5.53 | 6.87 | A | A | A | 1.24 |
| 156 | 75.8 | 4.8 | 2.16 | 7.74 | A | A | A | 0.48 |
| 157 | 70.89 | 3.47 | -4.46 | 6.61 | A | A | A | 1.00 |
| 158 | 76.7 | 7.0 | 3.37 | 10.15 | A | A | A | 0.76 |
| 159 | 68.6 | 0.9 | -7.55 | 4.64 | A | A | A | 1.69 |
| 160 | 74.2 | 2.5 | 0.00 | 5.58 | A | A | A | 0.00 |
| 161 | 72.9 | 2.1 | -1.75 | 5.30 | A | A | A | 0.39 |
| 162 | 72.8 | 2.1 | -1.89 | 5.30 | A | A | A | 0.42 |
| 163 | 75.7 | 2.2 | 2.02 | 5.31 | A | A | A | 0.45 |
| 164 | 77 | 4 | 3.77 | 6.84 | A | A | A | 0.85 |
| 165 | 84.6 | 5.1 | 14.02 | 7.49 | A | A | A | 3.14 |
| 166 | 75.3 | 4.0 | 1.48 | 6.93 | A | A | A | 0.33 |
| 168 | 71.35 | 2.06 | -3.84 | 5.30 | A | A | A | 0.86 |
| 169 | 74.45 | 1.04 | 0.34 | 4.66 | A | A | A | 0.08 |
| 171 | 70.30 | 6.40 | -5.26 | 10.13 | A | A | A | 1.18 |
| 172 | 84.1 | 11.8 | 13.34 | 14.72 | A | A | A | 2.99 |
| 173 | 85.7 | 3.0 | 15.50 | 5.66 | A | N | W | 3.47 |
| 176 | 74.3 | 4.34 | 0.13 | 7.34 | A | A | A | 0.03 |
| 179 | 72.7 | 1.2 | -2.02 | 4.74 | A | A | A | 0.45 |
| 180 | 76.05 | 5.60 | 2.49 | 8.60 | A | A | A | 0.56 |
| 181 | 72.468 | 1.113 | -2.33 | 4.71 | A | A | A | 0.52 |
| 182 | 76.71 | 4.82 | 3.38 | 7.70 | A | A | A | 0.76 |
| 183 | 76.2 | 2.1 | 2.70 | 5.23 | A | A | A | 0.60 |
| 184 | 73.69 | 0.75 | -0.69 | 4.56 | A | A | A | 0.15 |
| 185 | 75.5 | 5.3 | 1.75 | 8.31 | A | A | A | 0.39 |
| 186 | 76.7 | 4.7 | 3.37 | 7.57 | A | A | A | 0.76 |
| 187 | 481 | 11 | 548.25 | 5.00 | N | N | N | 122.90 |

TABLE 2. Evaluation Results for Co-60 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 189 | 76.3 | 2.5 | 2.83 | 5.52 | A | A | A | 0.63 |
| 191 | 74.9 | 2.8 | 0.94 | 5.81 | A | A | A | 0.21 |
| 192 | 0.039 | 0.002 | -99.95 | 6.79 | N | N | N | 22.41 |
| 193 | 74.5 | 7.45 | 0.40 | 10.94 | A | A | A | 0.09 |
| 195 | 77.1 | 2.5 | 3.91 | 5.50 | A | A | A | 0.88 |
| 196 | 73.0 | 2.8 | -1.62 | 5.87 | A | A | A | 0.36 |
| 197 | 74.7 | 2.4 | 0.67 | 5.49 | A | A | A | 0.15 |
| 198 | 73.10 | 3.25 | -1.48 | 6.29 | A | A | A | 0.33 |
| 199 | 67.06 | 1.68 | -9.62 | 5.10 | A | A | A | 2.16 |
| 200 | 76.04 | 1.39 | 2.48 | 4.81 | A | A | A | 0.56 |
| 201 | 72.0 | 7.2 | -2.96 | 10.94 | A | A | A | 0.66 |
| 202 | 73.47 | 4.31 | -0.98 | 7.36 | A | A | A | 0.22 |
| 205 | 74.3 | 28 | 0.13 | 37.95 | A | N | W | 0.03 |
| 206 | 73.4 | 2.6 | -1.08 | 5.69 | A | A | A | 0.24 |
| 207 | 74 | 3.9 | -0.27 | 6.90 | A | A | A | 0.06 |
| 208 | 68.0 | 3.13 | -8.36 | 6.40 | A | A | A | 1.87 |
| 209 | 75.8 | 4.5 | 2.16 | 7.42 | A | A | A | 0.48 |
| 210 | 77.2 | 5.0 | 4.04 | 7.86 | A | A | A | 0.91 |
| 211 | 73.1 | 6.4 | -1.48 | 9.82 | A | A | A | 0.33 |
| 212 | 77.1 | 1.9 | 3.91 | 5.08 | A | A | A | 0.88 |
| 213 | 75.5 | 4.6 | 1.75 | 7.54 | A | A | A | 0.39 |
| 214 | 73.9 | 5.2 | -0.40 | 8.32 | A | A | A | 0.09 |
| 215 | 73 | 3.84 | -1.62 | 6.89 | A | A | A | 0.36 |
| 216 | 73.56 | 3.07 | -0.86 | 6.10 | A | A | A | 0.19 |
| 219 | 74.8 | 5.4 | 0.81 | 8.48 | A | A | A | 0.18 |
| 221 | 73.14 | 2.76 | -1.43 | 5.83 | A | A | A | 0.32 |
| 222 | 70.542 | 1.46 | -4.93 | 4.91 | A | A | A | 1.11 |
| 223 | 73.8 | 2.8 | -0.54 | 5.85 | A | A | A | 0.12 |
| 224 | 79.8 | 3.2 | 7.55 | 5.99 | A | A | A | 1.69 |
| 225 | 73 | 3 | -1.62 | 6.06 | A | A | A | 0.36 |
| 227 | 74.13 | 1.07 | -0.09 | 4.68 | A | A | A | 0.02 |
| 228 | 72.1 | 2.9 | -2.83 | 6.00 | A | A | A | 0.63 |
| 229 | 74.3 | 2.34 | 0.13 | 5.45 | A | A | A | 0.03 |
| 230 | 73.95 | 2.8 | -0.34 | 5.84 | A | A | A | 0.08 |
| 231 | 73.8 | 11.70 | -0.54 | 16.47 | A | A | A | 0.12 |
| 232 | 72.945 | 2.9487 | -1.69 | 6.01 | A | A | A | 0.38 |
| 233 | 68.83 | 0.45 | -7.24 | 4.50 | A | A | A | 1.62 |
| 234 | 79.54 | 2.30 | 7.20 | 5.30 | A | A | A | 1.61 |
| 235 | 71.75 | 1.41 | -3.30 | 4.86 | A | A | A | 0.74 |
| 237 | 75.0 | 2.2 | 1.08 | 5.33 | A | A | A | 0.24 |
| 238 | 75.889 | 9.865 | 2.28 | 13.74 | A | A | A | 0.51 |
| 239 | 74.2 | 2.6 | 0.00 | 5.66 | A | A | A | 0.00 |
| 240 | 82.2 | 2.4 | 10.78 | 5.32 | A | A | A | 2.42 |
| 241 | 74.2 | 1.8 | 0.00 | 5.07 | A | A | A | 0.00 |
| 242 | 73.8 | 3.8 | -0.54 | 6.80 | A | A | A | 0.12 |
| 245 | 74.7 | 6.2 | 0.67 | 9.42 | A | A | A | 0.15 |
| 246 | 64.28 | 1.13 | -13.37 | 4.78 | A | N | W | 3.00 |
| 248 | 66.55 | 1.66 | -10.31 | 5.10 | A | A | A | 2.31 |
| 249 | 78.0 | 2.4 | 5.12 | 5.41 | A | A | A | 1.15 |
| 250 | 67.4 | 1.1 | -9.16 | 4.74 | A | A | A | 2.05 |
| 251 | 76.0 | 1.3 | 2.43 | 4.77 | A | A | A | 0.54 |
| 252 | 69.38059 | 2.02 | -6.50 | 5.32 | A | A | A | 1.46 |
| 254 | 74.1 | 3.3 | -0.13 | 6.29 | A | A | A | 0.03 |
| 255 | 74.4 | 2.2 | 0.27 | 5.34 | A | A | A | 0.06 |

TABLE 2. Evaluation Results for Co-60 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 256 | 70.32 | 0.39 | -5.23 | 4.48 | A | A | A | 1.17 |
| 257 | 81 | 2.5 | 9.16 | 5.41 | A | A | A | 2.05 |
| 258 | 73.8 | 3.34 | -0.54 | 6.35 | A | A | A | 0.12 |
| 259 | 69.54 | 0.20 | -6.28 | 4.46 | A | A | A | 1.41 |
| 260 | 71.7 | 5.3 | -3.37 | 8.63 | A | A | A | 0.76 |
| 261 | 75.2 | 3.9 | 1.35 | 6.83 | A | A | A | 0.30 |
| 262 | 75.4 | 4.3 | 1.62 | 7.23 | A | A | A | 0.36 |
| 263 | 71.87 | 2.15 | -3.14 | 5.36 | A | A | A | 0.70 |
| 266 | 67.5 | 1.9 | -9.03 | 5.26 | A | A | A | 2.02 |
| 267 | 126.84 | 7.87 | 70.94 | 7.63 | N | N | N | 15.90 |
| 268 | 72.4 | 2.6 | -2.43 | 5.72 | A | A | A | 0.54 |
| 270 | 73.2 | 7.92 | -1.35 | 11.70 | A | A | A | 0.30 |
| 274 | 71.4 | 3.8 | -3.77 | 6.94 | A | A | A | 0.85 |
| 275 | 83.2 | 5.09 | 12.13 | 7.56 | A | A | A | 2.72 |
| 279 | 74.4 | 2.1 | 0.27 | 5.27 | A | A | A | 0.06 |
| 281 | 67 | 2 | -9.70 | 5.36 | A | A | A | 2.18 |
| 283 | 75.3 | 2.43 | 1.48 | 5.49 | A | A | A | 0.33 |
| 286 | 76.0 | 3.2 | 2.43 | 6.12 | A | A | A | 0.54 |
| 287 | 72.2 | 1.5 | -2.70 | 4.91 | A | A | A | 0.60 |
| 289 | 77.0 | 4.2 | 3.77 | 7.04 | A | A | A | 0.85 |
| 290 | 67.24 | 2.94 | -9.38 | 6.24 | A | A | A | 2.10 |
| 291 | 76.3 | 3.1 | 2.83 | 6.02 | A | A | A | 0.63 |
| 295 | 77.0 | 5.2 | 3.77 | 8.09 | A | A | A | 0.85 |
| 297 | 67.1 | 2.7 | -9.57 | 6.00 | A | A | A | 2.15 |
| 298 | 73.3 | 1.5 | -1.21 | 4.90 | A | A | A | 0.27 |
| 299 | 77.8 | 4.7 | 4.85 | 7.50 | A | A | A | 1.09 |
| 304 | 5.83 | 0.44 | -92.14 | 8.76 | N | N | N | 20.66 |
| 307 | 56.31 | 2.2 | -24.11 | 5.92 | N | N | N | 5.40 |
| 310 | 75.2 | 1.7 | 1.35 | 4.99 | A | A | A | 0.30 |
| 312 | 61.05 | 1.08 | -17.72 | 4.79 | A | N | W | 3.97 |
| 313 | 77.2 | 2.0 | 4.04 | 5.15 | A | A | A | 0.91 |
| 314 | 74.63 | 3.40 | 0.58 | 6.37 | A | A | A | 0.13 |
| 316 | 74.931 | 2.997 | 0.99 | 5.98 | A | A | A | 0.22 |
| 317 | 73.3 |  | -1.21 |  | A |  |  | 0.27 |
| 320 | 74.74 | 2.31 | 0.73 | 5.42 | A | A | A | 0.16 |
| 321 | 69.14 | 0.21 | -6.82 | 4.46 | A | A | A | 1.53 |
| 322 | 74.8 | 7.4 | 0.81 | 10.85 | A | A | A | 0.18 |
| 323 | 72.9 | 2.0 | -1.75 | 5.23 | A | A | A | 0.39 |
| 324 | 74.49 | 1.17 | 0.39 | 4.72 | A | A | A | 0.09 |
| 325 | 73.0 | 6.2 | -1.62 | 9.59 | A | A | A | 0.36 |

## Reported Results for Cs-134 Determination in Sample 1

TABLE 3. Target values

| Parameter | Value |
| --- | --- |
| Target Value, Bq/kg | 113.2 |
| UNC, Bq/kg | 5.1 |
| MARB, % | 20 |

FIGURE 3. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

FIGURE 4. Reported results and evaluation parameters for Cs-134. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

TABLE 4. Evaluation Results for Cs-134 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 103.98 | 2.85 | -8.14 | 5.27 | A | A | A | 1.51 |
| 2 | 110.0 | 3.6 | -2.83 | 5.57 | A | A | A | 0.52 |
| 3 | 110.2 | 6.3 | -2.65 | 7.28 | A | A | A | 0.49 |
| 4 | 106.7 | 5.4 | -5.74 | 6.78 | A | A | A | 1.06 |
| 5 | 109 | 8.87 | -3.71 | 9.30 | A | A | A | 0.69 |
| 6 | 95.4 | 2 | -15.72 | 4.97 | A | N | W | 2.91 |
| 7 | 106.12 | 2.28 | -6.25 | 4.99 | A | A | A | 1.16 |
| 8 | 94.15 | 1.29 | -16.83 | 4.71 | A | N | W | 3.11 |
| 9 | 117 | 13 | 3.36 | 11.99 | A | A | A | 0.62 |
| 10 | 111.9 | 1.5 | -1.15 | 4.70 | A | A | A | 0.21 |
| 11 | 115.1 | 3.1 | 1.68 | 5.25 | A | A | A | 0.31 |
| 12 | 118.1 | 2.7 | 4.33 | 5.05 | A | A | A | 0.80 |
| 13 | 107.17 | 3.87 | -5.33 | 5.77 | A | A | A | 0.99 |
| 14 | 111.1 | 6.1 | -1.86 | 7.10 | A | A | A | 0.34 |
| 15 | 119.3 | 5.9 | 5.39 | 6.69 | A | A | A | 1.00 |
| 16 | 120.7 | 1.7 | 6.63 | 4.72 | A | A | A | 1.23 |
| 17 | 109.05 | 3.48 | -3.67 | 5.52 | A | A | A | 0.68 |
| 18 | 115.7 | 9.3 | 2.21 | 9.21 | A | A | A | 0.41 |
| 19 | 112.3 | 11.4 | -0.80 | 11.11 | A | A | A | 0.15 |
| 20 | 115 | 3 | 1.59 | 5.21 | A | A | A | 0.29 |
| 21 | 109.1 | 7.7 | -3.62 | 8.37 | A | A | A | 0.67 |
| 22 | 115.7 | 2.2 | 2.21 | 4.89 | A | A | A | 0.41 |
| 25 | 112 | 5 | -1.06 | 6.34 | A | A | A | 0.20 |
| 26 | 129.87 | 7.00 | 14.73 | 7.02 | A | A | A | 2.72 |
| 27 | 112 | 12 | -1.06 | 11.62 | A | A | A | 0.20 |
| 29 | 111.77 | 2.64 | -1.26 | 5.09 | A | A | A | 0.23 |
| 31 | 102.60 | 3.71 | -9.36 | 5.78 | A | A | A | 1.73 |
| 32 | 116.7 | 12.8 | 3.09 | 11.86 | A | A | A | 0.57 |
| 33 | 109.2 | 4.5 | -3.53 | 6.11 | A | A | A | 0.65 |
| 34 | 110.8 | 5.2 | -2.12 | 6.51 | A | A | A | 0.39 |

TABLE 4. Evaluation Results for Cs-134 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 35 | 108 | 8.8 | -4.59 | 9.31 | A | A | A | 0.85 |
| 36 | 114.3 | 5.2 | 0.97 | 6.40 | A | A | A | 0.18 |
| 38 | 109 | 6 | -3.71 | 7.11 | A | A | A | 0.69 |
| 39 | 113 | 5 | -0.18 | 6.31 | A | A | A | 0.03 |
| 41 | 113.6 | 4.5 | 0.35 | 6.00 | A | A | A | 0.07 |
| 42 | 85.95 | 1.45 | -24.07 | 4.81 | N | N | N | 4.45 |
| 43 | 113.86 | 3.42 | 0.58 | 5.41 | A | A | A | 0.11 |
| 46 | 237.81 | 5.23 | 110.08 | 5.01 | N | N | N | 20.36 |
| 47 | 80.1 | 5.1 | -29.24 | 7.80 | N | N | N | 5.41 |
| 48 | 113 | 4 | -0.18 | 5.73 | A | A | A | 0.03 |
| 50 | 104.79 | 4.26 | -7.43 | 6.07 | A | A | A | 1.37 |
| 51 | 111 | 6 | -1.94 | 7.04 | A | A | A | 0.36 |
| 52 | 93.4 | 5.2 | -17.49 | 7.16 | A | A | A | 3.24 |
| 53 | 113.9 | 5.1 | 0.62 | 6.35 | A | A | A | 0.11 |
| 55 | 113.0 | 2.3 | -0.18 | 4.94 | A | A | A | 0.03 |
| 56 | 116.7 | 4.34 | 3.09 | 5.84 | A | A | A | 0.57 |
| 59 | 113 | 8.0 | -0.18 | 8.39 | A | A | A | 0.03 |
| 61 | 101.86 | 2.49 | -10.02 | 5.13 | A | A | A | 1.85 |
| 63 | 108.51 | 2.00 | -4.14 | 4.87 | A | A | A | 0.77 |
| 64 | 109.0 | 5.2 | -3.71 | 6.56 | A | A | A | 0.69 |
| 66 | 118 | 2.6 | 4.24 | 5.02 | A | A | A | 0.78 |
| 68 | 115.4 | 8.9 | 1.94 | 8.93 | A | A | A | 0.36 |
| 69 | 105.08 | 2.00 | -7.17 | 4.89 | A | A | A | 1.33 |
| 70 | 113.29 | 1.03 | 0.08 | 4.60 | A | A | A | 0.01 |
| 72 | 112.558 | 5.534 | -0.57 | 6.67 | A | A | A | 0.10 |
| 74 | 113.5 | 5.3 | 0.27 | 6.49 | A | A | A | 0.05 |
| 75 | 98.9 | 1.9 | -12.63 | 4.90 | A | A | A | 2.34 |
| 76 | 99.00 | 0.54 | -12.54 | 4.54 | A | N | W | 2.32 |
| 77 | 132.9 | 1.5 | 17.40 | 4.64 | A | N | W | 3.22 |
| 78 | 106 | 1.1 | -6.36 | 4.62 | A | A | A | 1.18 |
| 80 | 107.4 | 7.9 | -5.12 | 8.63 | A | A | A | 0.95 |
| 81 | 115.0 | 8.2 | 1.59 | 8.43 | A | A | A | 0.29 |
| 84 | 116.34 | 1.4 | 2.77 | 4.66 | A | A | A | 0.51 |
| 85 | 118.09 | 9.60 | 4.32 | 9.29 | A | A | A | 0.80 |
| 86 | 110 | 2.58 | -2.83 | 5.08 | A | A | A | 0.52 |
| 87 | 115 | 5 | 1.59 | 6.26 | A | A | A | 0.29 |
| 88 | 110.0 | 6.5 | -2.83 | 7.43 | A | A | A | 0.52 |
| 92 | 110 | 7 | -2.83 | 7.80 | A | A | A | 0.52 |
| 93 | 98.23 | 0.75 | -13.22 | 4.57 | A | N | W | 2.45 |
| 94 | 107.71 | 0.37 | -4.85 | 4.52 | A | A | A | 0.90 |
| 95 | 108.47 | 0.35 | -4.18 | 4.52 | A | A | A | 0.77 |
| 96 | 117.3 | 7.0 | 3.62 | 7.48 | A | A | A | 0.67 |
| 98 | 108.2 | 5.0 | -4.42 | 6.45 | A | A | A | 0.82 |
| 99 | 114 | 3.8 | 0.71 | 5.60 | A | A | A | 0.13 |
| 101 | 115.8 | 3.0 | 2.30 | 5.20 | A | A | A | 0.42 |
| 102 | 120.8 | 4.3 | 6.71 | 5.74 | A | A | A | 1.24 |
| 103 | 116.57 | 2.45 | 2.98 | 4.97 | A | A | A | 0.55 |
| 105 | 118.04 | 1.46 | 4.28 | 4.67 | A | A | A | 0.79 |
| 106 | 112.5 | 0.5 | -0.62 | 4.53 | A | A | A | 0.11 |
| 107 | 173 | 8 | 52.83 | 6.46 | N | N | N | 9.77 |
| 108 | 112.0638 | 0.5731 | -1.00 | 4.53 | A | A | A | 0.19 |
| 109 | 113.7 | 5.3 | 0.44 | 6.48 | A | A | A | 0.08 |
| 110 | 105.387 | 1.694 | -6.90 | 4.78 | A | A | A | 1.28 |
| 111 | 118.0 | 7.8 | 4.24 | 8.00 | A | A | A | 0.78 |

TABLE 4. Evaluation Results for Cs-134 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 112 | 70.6188 | 2.475 | -37.62 | 5.71 | N | N | N | 6.96 |
| 113 | 108.8 | 5.5 | -3.89 | 6.77 | A | A | A | 0.72 |
| 116 | 114.6 | 1.1 | 1.24 | 4.61 | A | A | A | 0.23 |
| 118 | 88.7 | 6.5 | -21.64 | 8.60 | N | A | N | 4.00 |
| 120 | 110 | 4 | -2.83 | 5.79 | A | A | A | 0.52 |
| 121 | 110 | 11 | -2.83 | 10.97 | A | A | A | 0.52 |
| 123 | 121 | 8 | 6.89 | 8.00 | A | A | A | 1.27 |
| 124 | 113.0 | 1.6 | -0.18 | 4.72 | A | A | A | 0.03 |
| 125 | 97.2 | 2.6 | -14.13 | 5.24 | A | N | W | 2.61 |
| 126 | 100 | 6.1 | -11.66 | 7.58 | A | A | A | 2.16 |
| 129 | 113 | 7 | -0.18 | 7.66 | A | A | A | 0.03 |
| 130 | 108 | 4 | -4.59 | 5.83 | A | A | A | 0.85 |
| 132 | 101 | 6.0 | -10.78 | 7.46 | A | A | A | 1.99 |
| 133 | 112.0 | 5.7 | -1.06 | 6.80 | A | A | A | 0.20 |
| 134 | 112.57 | 3.65 | -0.56 | 5.55 | A | A | A | 0.10 |
| 135 | 111.6 | 3.7 | -1.41 | 5.59 | A | A | A | 0.26 |
| 137 | 106.519 | 1.709 | -5.90 | 4.78 | A | A | A | 1.09 |
| 139 | 94.819 | 1.1534 | -16.24 | 4.67 | A | N | W | 3.00 |
| 140 | 109.2 | 3.4 | -3.53 | 5.48 | A | A | A | 0.65 |
| 141 | 112 | 6.4 | -1.06 | 7.28 | A | A | A | 0.20 |
| 143 | 105 | 6.2 | -7.24 | 7.43 | A | A | A | 1.34 |
| 144 | 112.5 | 5.89 | -0.62 | 6.91 | A | A | A | 0.11 |
| 146 | 111.7 | 4.1 | -1.33 | 5.81 | A | A | A | 0.25 |
| 147 | 111 | 10 | -1.94 | 10.07 | A | A | A | 0.36 |
| 148 | 116.2 | 6.95 | 2.65 | 7.49 | A | A | A | 0.49 |
| 149 | 112.9 | 2.4 | -0.27 | 4.98 | A | A | A | 0.05 |
| 150 | 117.8 | 4.2 | 4.06 | 5.75 | A | A | A | 0.75 |
| 151 | 111.0 | 7.0 | -1.94 | 7.75 | A | A | A | 0.36 |
| 152 | 113 | 4 | -0.18 | 5.73 | A | A | A | 0.03 |
| 153 | 117.1 | 8.2 | 3.45 | 8.33 | A | A | A | 0.64 |
| 154 | 104.5 | 1.8 | -7.69 | 4.82 | A | A | A | 1.42 |
| 155 | 112.5 | 5.0 | -0.62 | 6.33 | A | A | A | 0.11 |
| 156 | 113 | 8 | -0.18 | 8.39 | A | A | A | 0.03 |
| 157 | 98.35 | 4.83 | -13.12 | 6.66 | A | A | A | 2.43 |
| 158 | 99.1 | 12.4 | -12.46 | 13.30 | A | A | A | 2.30 |
| 159 | 92.7 | 1.2 | -18.11 | 4.69 | A | N | W | 3.35 |
| 160 | 104 | 4 | -8.13 | 5.92 | A | A | A | 1.50 |
| 161 | 110.5 | 3.1 | -2.39 | 5.31 | A | A | A | 0.44 |
| 162 | 112.9 | 2.7 | -0.27 | 5.10 | A | A | A | 0.05 |
| 163 | 114.0 | 4.6 | 0.71 | 6.05 | A | A | A | 0.13 |
| 164 | 113 | 6 | -0.18 | 6.96 | A | A | A | 0.03 |
| 165 | 115.8 | 5.8 | 2.30 | 6.74 | A | A | A | 0.42 |
| 166 | 114.8 | 6.5 | 1.41 | 7.24 | A | A | A | 0.26 |
| 168 | 106.66 | 3.15 | -5.78 | 5.39 | A | A | A | 1.07 |
| 169 | 112.36 | 1.63 | -0.74 | 4.73 | A | A | A | 0.14 |
| 171 | 111.34 | 11.13 | -1.64 | 10.96 | A | A | A | 0.30 |
| 172 | 111.5 | 14.2 | -1.50 | 13.51 | A | A | A | 0.28 |
| 173 | 109.3 | 4.3 | -3.45 | 5.98 | A | A | A | 0.64 |
| 176 | 107 | 6.15 | -5.48 | 7.30 | A | A | A | 1.01 |
| 177 | 95 | 7 | -16.08 | 8.64 | A | A | A | 2.97 |
| 179 | 100.2 | 1.1 | -11.48 | 4.64 | A | A | A | 2.12 |
| 180 | 114.05 | 5.28 | 0.75 | 6.46 | A | A | A | 0.14 |
| 181 | 96.606 | 1.287 | -14.66 | 4.70 | A | N | W | 2.71 |
| 182 | 110.20 | 4.60 | -2.65 | 6.14 | A | A | A | 0.49 |

TABLE 4. Evaluation Results for Cs-134 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 183 | 112.62 | 3.84 | -0.51 | 5.65 | A | A | A | 0.09 |
| 184 | 98.47 | 0.68 | -13.01 | 4.56 | A | N | W | 2.41 |
| 185 | 110.8 | 7.2 | -2.12 | 7.91 | A | A | A | 0.39 |
| 186 | 115.4 | 7.1 | 1.94 | 7.63 | A | A | A | 0.36 |
| 187 | 690.7 | 13.3 | 510.16 | 4.90 | N | N | N | 94.36 |
| 189 | 112.1 | 3.9 | -0.97 | 5.69 | A | A | A | 0.18 |
| 190 | 103 | 12 | -9.01 | 12.49 | A | A | A | 1.67 |
| 191 | 113 | 5.6 | -0.18 | 6.70 | A | A | A | 0.03 |
| 193 | 110.3 | 11.03 | -2.56 | 10.97 | A | A | A | 0.47 |
| 195 | 107.3 | 3.90 | -5.21 | 5.79 | A | A | A | 0.96 |
| 196 | 115.2 | 8.8 | 1.77 | 8.87 | A | A | A | 0.33 |
| 197 | 113 | 4 | -0.18 | 5.73 | A | A | A | 0.03 |
| 198 | 107.57 | 5.76 | -4.97 | 7.00 | A | A | A | 0.92 |
| 199 | 97.01 | 2.33 | -14.30 | 5.11 | A | N | W | 2.65 |
| 200 | 115.16 | 1.54 | 1.73 | 4.70 | A | A | A | 0.32 |
| 201 | 102 | 10 | -9.89 | 10.79 | A | A | A | 1.83 |
| 202 | 112.77 | 6.55 | -0.38 | 7.35 | A | A | A | 0.07 |
| 205 | 113.6 | 6 | 0.35 | 6.94 | A | A | A | 0.07 |
| 206 | 110.7 | 4.2 | -2.21 | 5.89 | A | A | A | 0.41 |
| 207 | 109 | 5.7 | -3.71 | 6.90 | A | A | A | 0.69 |
| 208 | 86.4 | 3.89 | -23.67 | 6.37 | N | N | N | 4.38 |
| 209 | 106 | 6.7 | -6.36 | 7.76 | A | A | A | 1.18 |
| 210 | 112 | 8 | -1.06 | 8.45 | A | A | A | 0.20 |
| 211 | 115 | 11 | 1.59 | 10.57 | A | A | A | 0.29 |
| 212 | 112 | 2 | -1.06 | 4.85 | A | A | A | 0.20 |
| 213 | 114 | 7 | 0.71 | 7.62 | A | A | A | 0.13 |
| 214 | 101.7 | 7.1 | -10.16 | 8.31 | A | A | A | 1.88 |
| 215 | 109 | 6.8 | -3.71 | 7.70 | A | A | A | 0.69 |
| 216 | 110.7 | 5.04 | -2.21 | 6.41 | A | A | A | 0.41 |
| 219 | 111 | 7.1 | -1.94 | 7.82 | A | A | A | 0.36 |
| 221 | 109.62 | 4.24 | -3.16 | 5.94 | A | A | A | 0.58 |
| 222 | 91.78 | 1.49 | -18.92 | 4.79 | A | N | W | 3.50 |
| 223 | 107.3 | 3.0 | -5.21 | 5.30 | A | A | A | 0.96 |
| 224 | 116.2 | 4.2 | 2.65 | 5.78 | A | A | A | 0.49 |
| 225 | 111 | 5 | -1.94 | 6.37 | A | A | A | 0.36 |
| 227 | 112.51 | 1.3 | -0.61 | 4.65 | A | A | A | 0.11 |
| 228 | 110.9 | 4.4 | -2.03 | 6.00 | A | A | A | 0.38 |
| 229 | 109 | 3.35 | -3.71 | 5.45 | A | A | A | 0.69 |
| 230 | 113.0 | 4.0 | -0.18 | 5.73 | A | A | A | 0.03 |
| 231 | 117.1 | 18.36 | 3.45 | 16.31 | A | A | A | 0.64 |
| 232 | 106.5404 | 4.5378 | -5.88 | 6.20 | A | A | A | 1.09 |
| 233 | 105.86 | 0.82 | -6.48 | 4.57 | A | A | A | 1.20 |
| 234 | 114.73 | 2.86 | 1.35 | 5.15 | A | A | A | 0.25 |
| 235 | 105.16 | 1.55 | -7.10 | 4.74 | A | A | A | 1.31 |
| 237 | 110.2 | 3.9 | -2.65 | 5.73 | A | A | A | 0.49 |
| 238 | 114.42 | 14.875 | 1.08 | 13.76 | A | A | A | 0.20 |
| 239 | 111 | 4 | -1.94 | 5.77 | A | A | A | 0.36 |
| 240 | 116.2 | 3.3 | 2.65 | 5.33 | A | A | A | 0.49 |
| 241 | 120 | 3 | 6.01 | 5.15 | A | A | A | 1.11 |
| 242 | 109.1 | 5.6 | -3.62 | 6.83 | A | A | A | 0.67 |
| 243 | 102.8 | 5.2 | -9.19 | 6.77 | A | A | A | 1.70 |
| 245 | 96.3 | 0.95 | -14.93 | 4.61 | A | N | W | 2.76 |
| 246 | 90.16 | 1.22 | -20.35 | 4.70 | N | N | N | 3.76 |
| 248 | 97.51 | 2.44 | -13.86 | 5.15 | A | N | W | 2.56 |

TABLE 4. Evaluation Results for Cs-134 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 249 | 112 | 3 | -1.06 | 5.24 | A | A | A | 0.20 |
| 250 | 90.4 | 2.4 | -20.14 | 5.23 | N | N | N | 3.73 |
| 251 | 107.9 | 0.9 | -4.68 | 4.58 | A | A | A | 0.87 |
| 252 | 99.61 | 1.87 | -12.01 | 4.88 | A | A | A | 2.22 |
| 254 | 111.5 | 4.9 | -1.50 | 6.29 | A | A | A | 0.28 |
| 255 | 114.6 | 3.3 | 1.24 | 5.35 | A | A | A | 0.23 |
| 256 | 102.83 | 0.47 | -9.16 | 4.53 | A | A | A | 1.69 |
| 257 | 113 | 3.5 | -0.18 | 5.47 | A | A | A | 0.03 |
| 258 | 108 | 4.81 | -4.59 | 6.34 | A | A | A | 0.85 |
| 259 | 113 | 0.52 | -0.18 | 4.53 | A | A | A | 0.03 |
| 260 | 128.7 | 6.6 | 13.69 | 6.83 | A | A | A | 2.53 |
| 261 | 108 | 5.5 | -4.59 | 6.80 | A | A | A | 0.85 |
| 262 | 113.9 | 6.5 | 0.62 | 7.27 | A | A | A | 0.11 |
| 263 | 94.37 | 2.88 | -16.63 | 5.44 | A | N | W | 3.08 |
| 266 | 80.0 | 2.0 | -29.33 | 5.15 | N | N | N | 5.42 |
| 267 | 183.89 | 12.10 | 62.45 | 7.97 | N | N | N | 11.55 |
| 268 | 102.9 | 3.1 | -9.10 | 5.42 | A | A | A | 1.68 |
| 270 | 109 | 11.8 | -3.71 | 11.73 | A | A | A | 0.69 |
| 273 | 110.7 | 2.15 | -2.21 | 4.91 | A | A | A | 0.41 |
| 274 | 105.2 | 5.7 | -7.07 | 7.05 | A | A | A | 1.31 |
| 275 | 116 | 7.22 | 2.47 | 7.68 | A | A | A | 0.46 |
| 279 | 112.8 | 4.8 | -0.35 | 6.20 | A | A | A | 0.07 |
| 281 | 97 | 4 | -14.31 | 6.11 | A | A | A | 2.65 |
| 283 | 111 | 2.7 | -1.94 | 5.12 | A | A | A | 0.36 |
| 286 | 113.5 | 4.6 | 0.27 | 6.06 | A | A | A | 0.05 |
| 287 | 110.9 | 1.2 | -2.03 | 4.63 | A | A | A | 0.38 |
| 289 | 115 | 7 | 1.59 | 7.57 | A | A | A | 0.29 |
| 290 | 64.45 | 0.52 | -43.07 | 4.58 | N | N | N | 7.97 |
| 291 | 106 | 4 | -6.36 | 5.88 | A | A | A | 1.18 |
| 295 | 95.0 | 5.7 | -16.08 | 7.50 | A | A | A | 2.97 |
| 297 | 98.3 | 6.2 | -13.16 | 7.75 | A | A | A | 2.43 |
| 299 | 122.5 | 7.3 | 8.22 | 7.47 | A | A | A | 1.52 |
| 307 | 83.3 | 1.7 | -26.41 | 4.95 | N | N | N | 4.89 |
| 308 | 79.33 | 2.42 | -29.92 | 5.44 | N | N | N | 5.53 |
| 310 | 105.0 | 2.3 | -7.24 | 5.01 | A | A | A | 1.34 |
| 312 | 109.89 | 2.35 | -2.92 | 4.99 | A | A | A | 0.54 |
| 313 | 109.3 | 2.8 | -3.45 | 5.18 | A | A | A | 0.64 |
| 314 | 112.35 | 5.11 | -0.75 | 6.40 | A | A | A | 0.14 |
| 316 | 104.865 | 7.341 | -7.36 | 8.32 | A | A | A | 1.36 |
| 317 | 103 |  | -9.01 |  | A |  |  | 1.67 |
| 320 | 115.40 | 3.42 | 1.94 | 5.39 | A | A | A | 0.36 |
| 321 | 86.26 | 0.61 | -23.80 | 4.56 | N | N | N | 4.40 |
| 322 | 106.0 | 9.9 | -6.36 | 10.37 | A | A | A | 1.18 |
| 323 | 111.9 | 2.5 | -1.15 | 5.03 | A | A | A | 0.21 |
| 324 | 103.03 | 1.34 | -8.98 | 4.69 | A | A | A | 1.66 |
| 325 | 110.9 | 8.9 | -2.03 | 9.20 | A | A | A | 0.38 |

## Reported Results for Cs-137 Determination in Sample 1

TABLE 5. Target values

| Parameter | Value |
| --- | --- |
| Target Value, Bq/kg | 69.2 |
| UNC, Bq/kg | 3.1 |
| MARB, % | 20 |

FIGURE 5. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

FIGURE 6. Reported results and evaluation parameters for Cs-137. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

TABLE 6. Evaluation Results for Cs-137 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 71.41 | 2.09 | 3.19 | 5.35 | A | A | A | 0.68 |
| 2 | 68.7 | 2.6 | -0.72 | 5.86 | A | A | A | 0.15 |
| 3 | 67.5 | 3.7 | -2.46 | 7.08 | A | A | A | 0.53 |
| 4 | 66.6 | 3.3 | -3.76 | 6.68 | A | A | A | 0.80 |
| 5 | 66.7 | 7 | -3.61 | 11.41 | A | A | A | 0.77 |
| 6 | 64.6 | 2.5 | -6.65 | 5.92 | A | A | A | 1.42 |
| 7 | 70.40 | 2.53 | 1.73 | 5.74 | A | A | A | 0.37 |
| 8 | 65.81 | 0.97 | -4.90 | 4.72 | A | A | A | 1.05 |
| 9 | 71 | 8 | 2.60 | 12.13 | A | A | A | 0.56 |
| 10 | 65.25 | 0.94 | -5.71 | 4.71 | A | A | A | 1.22 |
| 11 | 70.6 | 2.4 | 2.02 | 5.62 | A | A | A | 0.43 |
| 12 | 81.1 | 1.4 | 17.20 | 4.80 | A | N | W | 3.68 |
| 13 | 68.28 | 2.55 | -1.33 | 5.83 | A | A | A | 0.28 |
| 14 | 69.4 | 3.8 | 0.29 | 7.07 | A | A | A | 0.06 |
| 15 | 72.1 | 3.6 | 4.19 | 6.71 | A | A | A | 0.90 |
| 16 | 70.7 | 1.6 | 2.17 | 5.02 | A | A | A | 0.46 |
| 17 | 69.20 | 2.15 | 0.00 | 5.45 | A | A | A | 0.00 |
| 18 | 70.9 | 4.2 | 2.46 | 7.43 | A | A | A | 0.53 |
| 19 | 66.2 | 2.8 | -4.34 | 6.16 | A | A | A | 0.93 |
| 20 | 72.6 | 1.5 | 4.91 | 4.93 | A | A | A | 1.05 |
| 21 | 70.3 | 5.0 | 1.59 | 8.41 | A | A | A | 0.34 |
| 22 | 73.2 | 1.5 | 5.78 | 4.93 | A | A | A | 1.24 |
| 25 | 70 | 3 | 1.16 | 6.20 | A | A | A | 0.25 |
| 26 | 75.45 | 4.07 | 9.03 | 7.01 | A | A | A | 1.93 |
| 27 | 71.4 | 1.9 | 3.18 | 5.21 | A | A | A | 0.68 |
| 29 | 77.32 | 1.86 | 11.73 | 5.08 | A | A | A | 2.51 |
| 31 | 71.40 | 2.60 | 3.18 | 5.77 | A | A | A | 0.68 |
| 32 | 71.1 | 7.8 | 2.75 | 11.85 | A | A | A | 0.59 |
| 33 | 68.1 | 2.9 | -1.59 | 6.18 | A | A | A | 0.34 |
| 34 | 69.1 | 3.5 | -0.14 | 6.76 | A | A | A | 0.03 |

TABLE 6. Evaluation Results for Cs-137 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 35 | 66.3 | 4.4 | -4.19 | 8.01 | A | A | A | 0.90 |
| 36 | 69.7 | 3.2 | 0.72 | 6.41 | A | A | A | 0.15 |
| 38 | 67.1 | 4.0 | -3.03 | 7.46 | A | A | A | 0.65 |
| 39 | 68.7 | 1.7 | -0.72 | 5.12 | A | A | A | 0.15 |
| 41 | 68.6 | 3.4 | -0.87 | 6.68 | A | A | A | 0.19 |
| 42 | 81.60 | 1.38 | 17.92 | 4.79 | A | N | W | 3.84 |
| 43 | 69.82 | 2.13 | 0.90 | 5.42 | A | A | A | 0.19 |
| 45 | 140.6 | 7.5 | 103.18 | 6.97 | N | N | N | 22.11 |
| 46 | 139.81 | 3.76 | 102.04 | 5.23 | N | N | N | 21.86 |
| 47 | 69.5 | 1.3 | 0.43 | 4.85 | A | A | A | 0.09 |
| 48 | 69.6 | 2.5 | 0.58 | 5.74 | A | A | A | 0.12 |
| 50 | 71.95 | 2.09 | 3.97 | 5.34 | A | A | A | 0.85 |
| 51 | 69 | 5 | -0.29 | 8.52 | A | A | A | 0.06 |
| 52 | 69.7 | 2.7 | 0.72 | 5.92 | A | A | A | 0.15 |
| 53 | 68.3 | 4.2 | -1.30 | 7.61 | A | A | A | 0.28 |
| 55 | 69.4 | 1.7 | 0.29 | 5.11 | A | A | A | 0.06 |
| 56 | 72.03 | 2.08 | 4.09 | 5.33 | A | A | A | 0.88 |
| 59 | 68.9 | 4.8 | -0.43 | 8.28 | A | A | A | 0.09 |
| 60 | 76.4 | 13.4 | 10.40 | 18.10 | A | A | A | 2.23 |
| 61 | 69.42 | 2.10 | 0.32 | 5.41 | A | A | A | 0.07 |
| 63 | 68.02 | 2.09 | -1.71 | 5.43 | A | A | A | 0.37 |
| 64 | 70.8 | 3.4 | 2.31 | 6.57 | A | A | A | 0.50 |
| 66 | 75.2 | 4.5 | 8.67 | 7.48 | A | A | A | 1.86 |
| 68 | 68.5 | 3.8 | -1.01 | 7.13 | A | A | A | 0.22 |
| 69 | 70.92 | 1.72 | 2.49 | 5.09 | A | A | A | 0.53 |
| 70 | 69.83 | 0.89 | 0.91 | 4.66 | A | A | A | 0.20 |
| 72 | 68.613 | 4.531 | -0.85 | 7.98 | A | A | A | 0.18 |
| 74 | 75.0 | 3.5 | 8.38 | 6.47 | A | A | A | 1.80 |
| 75 | 69.1 | 2.3 | -0.14 | 5.58 | A | A | A | 0.03 |
| 76 | 69.00 | 0.39 | -0.29 | 4.52 | A | A | A | 0.06 |
| 77 | 88.5 | 0.9 | 27.89 | 4.59 | N | N | N | 5.98 |
| 78 | 68.7 | 0.75 | -0.72 | 4.61 | A | A | A | 0.15 |
| 80 | 73.7 | 5.2 | 6.50 | 8.36 | A | A | A | 1.39 |
| 81 | 74.8 | 5.5 | 8.09 | 8.61 | A | A | A | 1.73 |
| 84 | 72.47 | 1.08 | 4.73 | 4.72 | A | A | A | 1.01 |
| 85 | 73.08 | 6.16 | 5.61 | 9.55 | A | A | A | 1.20 |
| 86 | 68.4 | 2.33 | -1.16 | 5.63 | A | A | A | 0.25 |
| 87 | 76 | 3 | 9.83 | 5.97 | A | A | A | 2.11 |
| 88 | 68.6 | 4.8 | -0.87 | 8.31 | A | A | A | 0.19 |
| 89 | 71.98 | 2.6 | 4.02 | 5.75 | A | A | A | 0.86 |
| 92 | 70.3 | 7.4 | 1.59 | 11.44 | A | A | A | 0.34 |
| 93 | 69.85 | 0.95 | 0.94 | 4.68 | A | A | A | 0.20 |
| 94 | 65.88 | 0.39 | -4.80 | 4.52 | A | A | A | 1.03 |
| 95 | 66.39 | 0.38 | -4.06 | 4.52 | A | A | A | 0.87 |
| 96 | 69.4 | 4.2 | 0.29 | 7.53 | A | A | A | 0.06 |
| 98 | 71.2 | 5.0 | 2.89 | 8.33 | A | A | A | 0.62 |
| 99 | 68 | 2.1 | -1.73 | 5.44 | A | A | A | 0.37 |
| 101 | 72.8 | 2.0 | 5.20 | 5.26 | A | A | A | 1.11 |
| 102 | 73.9 | 2.3 | 6.79 | 5.45 | A | A | A | 1.46 |
| 103 | 68.93 | 1.46 | -0.39 | 4.96 | A | A | A | 0.08 |
| 105 | 71.83 | 0.92 | 3.80 | 4.66 | A | A | A | 0.81 |
| 106 | 69.5 | 0.4 | 0.43 | 4.52 | A | A | A | 0.09 |
| 107 | 81.9 | 1.4 | 18.35 | 4.79 | A | N | W | 3.93 |
| 108 | 70.3702 | 0.5407 | 1.69 | 4.55 | A | A | A | 0.36 |

TABLE 6. Evaluation Results for Cs-137 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 109 | 74.1 | 3.3 | 7.08 | 6.32 | A | A | A | 1.52 |
| 110 | 68.878 | 1.540 | -0.47 | 5.01 | A | A | A | 0.10 |
| 111 | 70.7 | 4.8 | 2.17 | 8.13 | A | A | A | 0.46 |
| 112 | 79.9779 | 3.185 | 15.58 | 5.99 | A | N | W | 3.34 |
| 113 | 70.3 | 3.5 | 1.59 | 6.70 | A | A | A | 0.34 |
| 116 | 69.8 | 1.3 | 0.87 | 4.85 | A | A | A | 0.19 |
| 118 | 65.9 | 5.6 | -4.77 | 9.61 | A | A | A | 1.02 |
| 120 | 67.9 | 2.8 | -1.88 | 6.09 | A | A | A | 0.40 |
| 121 | 68.1 | 3.0 | -1.59 | 6.28 | A | A | A | 0.34 |
| 123 | 79 | 5 | 14.16 | 7.75 | A | A | A | 3.03 |
| 124 | 70.6 | 1.2 | 2.02 | 4.79 | A | A | A | 0.43 |
| 125 | 68.9 | 2.4 | -0.43 | 5.67 | A | A | A | 0.09 |
| 126 | 62.3 | 3.8 | -9.97 | 7.57 | A | A | A | 2.14 |
| 129 | 69.6 | 2.0 | 0.58 | 5.32 | A | A | A | 0.12 |
| 130 | 70 | 3 | 1.16 | 6.20 | A | A | A | 0.25 |
| 132 | 65.6 | 3.9 | -5.20 | 7.44 | A | A | A | 1.11 |
| 133 | 68.2 | 3.4 | -1.45 | 6.70 | A | A | A | 0.31 |
| 134 | 70.66 | 4.31 | 2.11 | 7.57 | A | A | A | 0.45 |
| 135 | 68.7 | 2.7 | -0.72 | 5.96 | A | A | A | 0.15 |
| 137 | 63.958 | 1.692 | -7.58 | 5.20 | A | A | A | 1.62 |
| 139 | 72.559 | 1.0625 | 4.85 | 4.71 | A | A | A | 1.04 |
| 140 | 69.1 | 2.3 | -0.14 | 5.58 | A | A | A | 0.03 |
| 141 | 73.1 | 4.2 | 5.64 | 7.29 | A | A | A | 1.21 |
| 143 | 67.9 | 4.0 | -1.88 | 7.40 | A | A | A | 0.40 |
| 144 | 70.5 | 4.21 | 1.88 | 7.47 | A | A | A | 0.40 |
| 146 | 68.2 | 4.2 | -1.45 | 7.62 | A | A | A | 0.31 |
| 147 | 68.2 | 6.1 | -1.45 | 10.00 | A | A | A | 0.31 |
| 148 | 72.94 | 6.30 | 5.40 | 9.73 | A | A | A | 1.16 |
| 149 | 69.2 | 1.6 | 0.00 | 5.04 | A | A | A | 0.00 |
| 150 | 73.1 | 2.3 | 5.64 | 5.47 | A | A | A | 1.21 |
| 151 | 69.2 | 2.0 | 0.00 | 5.33 | A | A | A | 0.00 |
| 152 | 69.1 | 2.4 | -0.14 | 5.67 | A | A | A | 0.03 |
| 153 | 71.4 | 5.0 | 3.18 | 8.31 | A | A | A | 0.68 |
| 154 | 69.3 | 1.5 | 0.14 | 4.98 | A | A | A | 0.03 |
| 155 | 69.7 | 3.4 | 0.72 | 6.62 | A | A | A | 0.15 |
| 156 | 68.9 | 4.4 | -0.43 | 7.80 | A | A | A | 0.09 |
| 157 | 68.26 | 3.76 | -1.36 | 7.10 | A | A | A | 0.29 |
| 158 | 77.5 | 7.1 | 11.99 | 10.20 | A | A | A | 2.57 |
| 159 | 62.7 | 0.8 | -9.39 | 4.66 | A | A | A | 2.01 |
| 160 | 71.4 | 2.5 | 3.18 | 5.69 | A | A | A | 0.68 |
| 161 | 67.7 | 2.0 | -2.17 | 5.37 | A | A | A | 0.46 |
| 162 | 68.6 | 1.9 | -0.87 | 5.27 | A | A | A | 0.19 |
| 163 | 70.3 | 4.2 | 1.59 | 7.47 | A | A | A | 0.34 |
| 164 | 70 | 4 | 1.16 | 7.26 | A | A | A | 0.25 |
| 165 | 79.8 | 4.7 | 15.32 | 7.40 | A | A | A | 3.28 |
| 166 | 70.3 | 3.8 | 1.59 | 7.02 | A | A | A | 0.34 |
| 168 | 68.02 | 4.11 | -1.71 | 7.52 | A | A | A | 0.37 |
| 169 | 68.42 | 0.91 | -1.13 | 4.67 | A | A | A | 0.24 |
| 171 | 71.25 | 6.50 | 2.96 | 10.16 | A | A | A | 0.63 |
| 172 | 80.5 | 10.9 | 16.33 | 14.26 | A | A | A | 3.50 |
| 173 | 79.4 | 3.3 | 14.74 | 6.11 | A | A | A | 3.16 |
| 176 | 69.0 | 3.97 | -0.29 | 7.29 | A | A | A | 0.06 |
| 177 | 78 | 5 | 12.72 | 7.82 | A | A | A | 2.72 |
| 179 | 70.8 | 1.7 | 2.31 | 5.08 | A | A | A | 0.50 |

TABLE 6. Evaluation Results for Cs-137 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 180 | 69.80 | 2.3 | 0.87 | 5.56 | A | A | A | 0.19 |
| 181 | 68.035 | 1.519 | -1.68 | 5.01 | A | A | A | 0.36 |
| 182 | 73.84 | 4.56 | 6.71 | 7.63 | A | A | A | 1.44 |
| 183 | 70.19 | 2.22 | 1.43 | 5.48 | A | A | A | 0.31 |
| 184 | 68.13 | 0.66 | -1.55 | 4.58 | A | A | A | 0.33 |
| 185 | 69.5 | 5.0 | 0.43 | 8.47 | A | A | A | 0.09 |
| 186 | 70.7 | 4.4 | 2.17 | 7.67 | A | A | A | 0.46 |
| 187 | 488 | 12 | 605.20 | 5.11 | N | N | N | 129.66 |
| 189 | 74.0 | 4.6 | 6.94 | 7.66 | A | A | A | 1.49 |
| 190 | 71.7 | 7.2 | 3.61 | 11.00 | A | A | A | 0.77 |
| 191 | 69.9 | 2.4 | 1.01 | 5.64 | A | A | A | 0.22 |
| 192 | 0.041 | 0.003 | -99.94 | 8.58 | N | N | N | 21.41 |
| 193 | 68.07 | 6.81 | -1.63 | 10.96 | A | A | A | 0.35 |
| 195 | 72.0 | 2.70 | 4.05 | 5.84 | A | A | A | 0.87 |
| 196 | 65.4 | 2.7 | -5.49 | 6.09 | A | A | A | 1.18 |
| 197 | 68.8 | 2.3 | -0.58 | 5.59 | A | A | A | 0.12 |
| 198 | 66.46 | 2.95 | -3.96 | 6.31 | A | A | A | 0.85 |
| 199 | 64.94 | 1.56 | -6.16 | 5.08 | A | A | A | 1.32 |
| 200 | 71.47 | 1.89 | 3.28 | 5.20 | A | A | A | 0.70 |
| 201 | 69.0 | 6.9 | -0.29 | 10.96 | A | A | A | 0.06 |
| 202 | 68.51 | 4.03 | -1.00 | 7.39 | A | A | A | 0.21 |
| 205 | 69.8 | 2.7 | 0.87 | 5.92 | A | A | A | 0.19 |
| 206 | 70.2 | 2.8 | 1.45 | 6.00 | A | A | A | 0.31 |
| 207 | 68.9 | 4.0 | -0.43 | 7.33 | A | A | A | 0.09 |
| 208 | 59.2 | 2.71 | -14.45 | 6.40 | A | A | A | 3.10 |
| 209 | 71.7 | 4.6 | 3.61 | 7.82 | A | A | A | 0.77 |
| 210 | 71.4 | 5.2 | 3.18 | 8.55 | A | A | A | 0.68 |
| 211 | 69.9 | 5.3 | 1.01 | 8.81 | A | A | A | 0.22 |
| 212 | 69.1 | 1.4 | -0.14 | 4.92 | A | A | A | 0.03 |
| 213 | 70.5 | 4.3 | 1.88 | 7.57 | A | A | A | 0.40 |
| 214 | 70.1 | 4.9 | 1.30 | 8.30 | A | A | A | 0.28 |
| 215 | 63.7 | 3.97 | -7.95 | 7.68 | A | A | A | 1.70 |
| 216 | 68.16 | 4.16 | -1.50 | 7.57 | A | A | A | 0.32 |
| 219 | 70.9 | 5.7 | 2.46 | 9.20 | A | A | A | 0.53 |
| 221 | 75.53 | 3.16 | 9.15 | 6.13 | A | A | A | 1.96 |
| 222 | 83.4 | 2.25 | 20.52 | 5.23 | N | N | N | 4.40 |
| 223 | 66.4 | 2.4 | -4.05 | 5.76 | A | A | A | 0.87 |
| 224 | 75.3 | 3.3 | 8.82 | 6.27 | A | A | A | 1.89 |
| 225 | 68 | 3 | -1.73 | 6.29 | A | A | A | 0.37 |
| 227 | 68.99 | 1.03 | -0.30 | 4.72 | A | A | A | 0.07 |
| 228 | 68.4 | 2.7 | -1.16 | 5.97 | A | A | A | 0.25 |
| 229 | 70.5 | 2.23 | 1.88 | 5.48 | A | A | A | 0.40 |
| 230 | 68.1 | 2.6 | -1.59 | 5.89 | A | A | A | 0.34 |
| 231 | 72.81 | 13.03 | 5.22 | 18.45 | A | A | A | 1.12 |
| 232 | 69.4304 | 3.0280 | 0.33 | 6.25 | A | A | A | 0.07 |
| 233 | 66.44 | 0.76 | -3.99 | 4.62 | A | A | A | 0.85 |
| 234 | 75.12 | 1.95 | 8.55 | 5.18 | A | A | A | 1.83 |
| 235 | 80.78 | 2.38 | 16.73 | 5.36 | A | N | W | 3.59 |
| 237 | 67.4 | 4.1 | -2.60 | 7.55 | A | A | A | 0.56 |
| 238 | 71.19 | 9.25 | 2.88 | 13.74 | A | A | A | 0.62 |
| 239 | 66.8 | 2.3 | -3.47 | 5.65 | A | A | A | 0.74 |
| 240 | 75.0 | 2.2 | 8.38 | 5.35 | A | A | A | 1.80 |
| 241 | 73.5 | 4.4 | 6.21 | 7.48 | A | A | A | 1.33 |
| 242 | 67.6 | 3.5 | -2.31 | 6.85 | A | A | A | 0.50 |

TABLE 6. Evaluation Results for Cs-137 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 243 | 64.5 | 4.8 | -6.79 | 8.69 | A | A | A | 1.46 |
| 245 | 67.2 | 1.07 | -2.89 | 4.75 | A | A | A | 0.62 |
| 246 | 62.32 | 1.75 | -9.94 | 5.29 | A | A | A | 2.13 |
| 248 | 68.17 | 2.05 | -1.49 | 5.40 | A | A | A | 0.32 |
| 249 | 70.5 | 12.8 | 1.88 | 18.70 | A | A | A | 0.40 |
| 250 | 58.9 | 2.3 | -14.88 | 5.94 | A | A | A | 3.19 |
| 251 | 74.6 | 1.0 | 7.80 | 4.68 | A | A | A | 1.67 |
| 252 | 70.4041 | 1.59 | 1.74 | 5.02 | A | A | A | 0.37 |
| 254 | 68.7 | 3.1 | -0.72 | 6.36 | A | A | A | 0.15 |
| 255 | 72.0 | 2.1 | 4.05 | 5.35 | A | A | A | 0.87 |
| 256 | 65.01 | 0.46 | -6.05 | 4.54 | A | A | A | 1.30 |
| 257 | 76 | 3 | 9.83 | 5.97 | A | A | A | 2.11 |
| 258 | 68.7 | 3.06 | -0.72 | 6.32 | A | A | A | 0.15 |
| 259 | 62.14 | 0.14 | -10.20 | 4.49 | A | A | A | 2.19 |
| 260 | 71.4 | 3.7 | 3.18 | 6.85 | A | A | A | 0.68 |
| 261 | 68.7 | 3.5 | -0.72 | 6.78 | A | A | A | 0.15 |
| 262 | 68.5 | 4.0 | -1.01 | 7.36 | A | A | A | 0.22 |
| 263 | 68.06 | 2 | -1.65 | 5.36 | A | A | A | 0.35 |
| 264 | 80.8 | 1.7 | 16.76 | 4.95 | A | N | W | 3.59 |
| 266 | 67.6 | 4.1 | -2.31 | 7.54 | A | A | A | 0.50 |
| 267 | 126.73 | 8.17 | 83.14 | 7.85 | N | N | N | 17.81 |
| 268 | 63.59 | 1.27 | -8.11 | 4.90 | A | A | A | 1.74 |
| 270 | 69.9 | 7.57 | 1.01 | 11.72 | A | A | A | 0.22 |
| 273 | 71.46 | 1.60 | 3.27 | 5.01 | A | A | A | 0.70 |
| 274 | 66.1 | 3.7 | -4.48 | 7.17 | A | A | A | 0.96 |
| 275 | 68.9 | 4.26 | -0.43 | 7.64 | A | A | A | 0.09 |
| 279 | 69.4 | 4.2 | 0.29 | 7.53 | A | A | A | 0.06 |
| 281 | 68 | 3 | -1.73 | 6.29 | A | A | A | 0.37 |
| 283 | 69.9 | 3.2 | 1.01 | 6.41 | A | A | A | 0.22 |
| 286 | 71.0 | 2.8 | 2.60 | 5.97 | A | A | A | 0.56 |
| 287 | 67.9 | 1.2 | -1.88 | 4.82 | A | A | A | 0.40 |
| 289 | 69.4 | 3.5 | 0.29 | 6.75 | A | A | A | 0.06 |
| 290 | 80.34 | 0.99 | 16.10 | 4.65 | A | N | W | 3.45 |
| 291 | 66.6 | 2.5 | -3.76 | 5.84 | A | A | A | 0.80 |
| 295 | 65.0 | 4.7 | -6.07 | 8.51 | A | A | A | 1.30 |
| 297 | 61.8 | 4.0 | -10.69 | 7.87 | A | A | A | 2.29 |
| 298 | 61.8 | 1.3 | -10.69 | 4.95 | A | A | A | 2.29 |
| 299 | 84.9 | 5.2 | 22.69 | 7.59 | N | N | N | 4.86 |
| 304 | 50.37 | 3.78 | -27.21 | 8.74 | N | N | N | 5.83 |
| 307 | 69.8 | 0.6 | 0.87 | 4.56 | A | A | A | 0.19 |
| 308 | 63.73 | 3.88 | -7.90 | 7.56 | A | A | A | 1.69 |
| 310 | 68.9 | 1.5 | -0.43 | 4.98 | A | A | A | 0.09 |
| 312 | 66.69 | 1.06 | -3.63 | 4.75 | A | A | A | 0.78 |
| 313 | 66.7 | 1.7 | -3.61 | 5.15 | A | A | A | 0.77 |
| 314 | 68.38 | 3.16 | -1.18 | 6.44 | A | A | A | 0.25 |
| 316 | 70.546 | 5.644 | 1.95 | 9.17 | A | A | A | 0.42 |
| 317 | 71 |  | 2.60 |  | A |  |  | 0.56 |
| 320 | 64.11 | 2.42 | -7.36 | 5.86 | A | A | A | 1.58 |
| 321 | 67.96 | 0.41 | -1.79 | 4.52 | A | A | A | 0.38 |
| 322 | 75.7 | 8.3 | 9.39 | 11.84 | A | A | A | 2.01 |
| 323 | 68.3 | 1.3 | -1.30 | 4.87 | A | A | A | 0.28 |
| 324 | 63.62 | 1.24 | -8.06 | 4.89 | A | A | A | 1.73 |
| 325 | 67.8 | 4.0 | -2.02 | 7.41 | A | A | A | 0.43 |

## Reported Results for Am-241 Determination in Sample 1

TABLE 7. Target values

| Parameter | Value |
| --- | --- |
| Target Value, Bq/kg | 60.9 |
| UNC, Bq/kg | 2.7 |
| MARB, % | 20 |

FIGURE 7. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

FIGURE 8. Reported results and evaluation parameters for Am-241. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

TABLE 8. Evaluation Results for Am-241 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 70.07 | 2.09 | 15.06 | 5.34 | A | N | W | 2.30 |
| 2 | 65.3 | 3.5 | 7.22 | 6.96 | A | A | A | 1.11 |
| 3 | 61.0 | 3.6 | 0.16 | 7.38 | A | A | A | 0.03 |
| 4 | 27.2 | 1.4 | -55.34 | 6.79 | N | N | N | 8.47 |
| 5 | 60.4 | 8 | -0.82 | 13.97 | A | A | A | 0.13 |
| 6 | 45.9 | 2.5 | -24.63 | 7.02 | N | N | N | 3.77 |
| 7 | 53.09 | 4.49 | -12.82 | 9.55 | A | A | A | 1.96 |
| 8 | 66.11 | 1.01 | 8.56 | 4.69 | A | A | A | 1.31 |
| 9 | 65 | 11 | 6.73 | 17.49 | A | A | A | 1.03 |
| 10 | 58.7 | 1.2 | -3.61 | 4.88 | A | A | A | 0.55 |
| 11 | 58.7 | 2.1 | -3.61 | 5.70 | A | A | A | 0.55 |
| 12 | 84.9 | 1.5 | 39.41 | 4.77 | N | N | N | 6.03 |
| 13 | 61.76 | 2.29 | 1.41 | 5.78 | A | A | A | 0.22 |
| 14 | 61.7 | 4.5 | 1.31 | 8.54 | A | A | A | 0.20 |
| 15 | 68.4 | 6.3 | 12.32 | 10.22 | A | A | A | 1.88 |
| 16 | 60.2 | 2.9 | -1.15 | 6.55 | A | A | A | 0.18 |
| 17 | 63.24 | 3.40 | 3.84 | 6.97 | A | A | A | 0.59 |
| 18 | 63.7 | 4.6 | 4.60 | 8.47 | A | A | A | 0.70 |
| 19 | 61.7 | 3.0 | 1.31 | 6.58 | A | A | A | 0.20 |
| 20 | 63.2 | 2.6 | 3.78 | 6.05 | A | A | A | 0.58 |
| 21 | 62.9 | 4.7 | 3.28 | 8.69 | A | A | A | 0.50 |
| 22 | 63.5 | 2.0 | 4.27 | 5.44 | A | A | A | 0.65 |
| 25 | 61 | 4 | 0.16 | 7.92 | A | A | A | 0.03 |
| 26 | 58.69 | 3.31 | -3.63 | 7.17 | A | A | A | 0.56 |
| 27 | 66.3 | 3.1 | 8.87 | 6.44 | A | A | A | 1.36 |
| 29 | 71.56 | 2.41 | 17.50 | 5.57 | A | N | W | 2.68 |
| 31 | 63.90 | 4.02 | 4.93 | 7.70 | A | A | A | 0.75 |
| 32 | 64.3 | 10.3 | 5.58 | 16.62 | A | A | A | 0.85 |
| 33 | 65.2 | 2.5 | 7.06 | 5.86 | A | A | A | 1.08 |
| 34 | 61.5 | 3.8 | 0.99 | 7.60 | A | A | A | 0.15 |

TABLE 8. Evaluation Results for Am-241 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 35 | 56.8 | 6.5 | -6.73 | 12.27 | A | A | A | 1.03 |
| 36 | 61 | 3.4 | 0.16 | 7.12 | A | A | A | 0.03 |
| 38 | 60.6 | 4.5 | -0.49 | 8.65 | A | A | A | 0.08 |
| 39 | 62.6 | 2.4 | 2.79 | 5.86 | A | A | A | 0.43 |
| 41 | 63.5 | 3.1 | 4.27 | 6.59 | A | A | A | 0.65 |
| 43 | 63.38 | 2.59 | 4.07 | 6.03 | A | A | A | 0.62 |
| 45 | 100.5 | 5.4 | 65.02 | 6.97 | N | N | N | 9.95 |
| 46 | 75.36 | 2.22 | 23.74 | 5.32 | N | N | N | 3.63 |
| 47 | 139 | 17 | 128.24 | 13.01 | N | N | N | 19.62 |
| 48 | 59.5 | 2.1 | -2.30 | 5.67 | A | A | A | 0.35 |
| 50 | 68.42 | 9.90 | 12.35 | 15.13 | A | A | A | 1.89 |
| 51 | 64 | 7 | 5.09 | 11.80 | A | A | A | 0.78 |
| 52 | 61.5 | 2.7 | 0.99 | 6.24 | A | A | A | 0.15 |
| 53 | 59.7 | 2.1 | -1.97 | 5.66 | A | A | A | 0.30 |
| 56 | 62.27 | 1.63 | 2.25 | 5.15 | A | A | A | 0.34 |
| 59 | 61.7 | 4.7 | 1.31 | 8.81 | A | A | A | 0.20 |
| 63 | 78.51 | 3.50 | 28.92 | 6.29 | N | N | N | 4.42 |
| 64 | 62.4 | 3.5 | 2.46 | 7.15 | A | A | A | 0.38 |
| 66 | 66.2 | 6.8 | 8.70 | 11.19 | A | A | A | 1.33 |
| 68 | 68.1 | 4.6 | 11.82 | 8.08 | A | A | A | 1.81 |
| 69 | 68.29 | 4.80 | 12.13 | 8.31 | A | A | A | 1.86 |
| 70 | 51.51 | 1.97 | -15.42 | 5.86 | A | N | W | 2.36 |
| 72 | 61.803 | 4.261 | 1.48 | 8.20 | A | A | A | 0.23 |
| 74 | 62.8 | 3.3 | 3.12 | 6.88 | A | A | A | 0.48 |
| 75 | 83.6 | 6.1 | 37.27 | 8.54 | N | N | N | 5.70 |
| 76 | 61.30 | 1.54 | 0.66 | 5.10 | A | A | A | 0.10 |
| 78 | 62.3 | 1.0 | 2.30 | 4.72 | A | A | A | 0.35 |
| 80 | 81.6 | 5.9 | 33.99 | 8.48 | N | N | N | 5.20 |
| 81 | 61.9 | 4.9 | 1.64 | 9.07 | A | A | A | 0.25 |
| 84 | 61.32 | 1.11 | 0.69 | 4.79 | A | A | A | 0.11 |
| 85 | 64.30 | 6.03 | 5.58 | 10.37 | A | A | A | 0.85 |
| 86 | 66.5 | 4.23 | 9.20 | 7.75 | A | A | A | 1.41 |
| 87 | 65 | 4 | 6.73 | 7.58 | A | A | A | 1.03 |
| 88 | 61.0 | 4.5 | 0.16 | 8.61 | A | A | A | 0.03 |
| 91 | 61 | 7 | 0.16 | 12.30 | A | A | A | 0.03 |
| 92 | 64.7 | 7.0 | 6.24 | 11.69 | A | A | A | 0.95 |
| 93 | 62.66 | 0.95 | 2.89 | 4.69 | A | A | A | 0.44 |
| 94 | 58.48 | 0.78 | -3.97 | 4.63 | A | A | A | 0.61 |
| 95 | 57.50 | 0.73 | -5.58 | 4.61 | A | A | A | 0.85 |
| 96 | 66.2 | 4.4 | 8.70 | 7.99 | A | A | A | 1.33 |
| 98 | 60.8 | 5.0 | -0.16 | 9.34 | A | A | A | 0.03 |
| 99 | 59.1 | 4.8 | -2.96 | 9.25 | A | A | A | 0.45 |
| 101 | 62.1 | 1.9 | 1.97 | 5.39 | A | A | A | 0.30 |
| 102 | 68.7 | 2.1 | 12.81 | 5.39 | A | A | A | 1.96 |
| 103 | 61.41 | 1.83 | 0.84 | 5.34 | A | A | A | 0.13 |
| 105 | 57.82 | 0.88 | -5.06 | 4.69 | A | A | A | 0.77 |
| 106 | 62.6 | 1.2 | 2.79 | 4.83 | A | A | A | 0.43 |
| 107 | 70.9 | 1.6 | 16.42 | 4.97 | A | N | W | 2.51 |
| 109 | 63.9 | 3.2 | 4.93 | 6.69 | A | A | A | 0.75 |
| 111 | 67.7 | 4.3 | 11.17 | 7.75 | A | A | A | 1.71 |
| 113 | 63.1 | 3.4 | 3.61 | 6.98 | A | A | A | 0.55 |
| 116 | 63.5 | 1.9 | 4.27 | 5.35 | A | A | A | 0.65 |
| 118 | 63.2 | 5.6 | 3.78 | 9.91 | A | A | A | 0.58 |
| 120 | 65.4 | 2.9 | 7.39 | 6.27 | A | A | A | 1.13 |

TABLE 8. Evaluation Results for Am-241 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 121 | 62.1 | 4.4 | 1.97 | 8.36 | A | A | A | 0.30 |
| 123 | 76 | 5 | 24.79 | 7.93 | N | N | N | 3.79 |
| 124 | 62.4 | 1.4 | 2.46 | 4.97 | A | A | A | 0.38 |
| 125 | 59.4 | 2.6 | -2.46 | 6.23 | A | A | A | 0.38 |
| 126 | 56.8 | 5.8 | -6.73 | 11.13 | A | A | A | 1.03 |
| 129 | 64.1 | 2.0 | 5.25 | 5.42 | A | A | A | 0.80 |
| 130 | 56 | 3 | -8.05 | 6.95 | A | A | A | 1.23 |
| 132 | 58.7 | 3.8 | -3.61 | 7.85 | A | A | A | 0.55 |
| 133 | 61.9 | 3.2 | 1.64 | 6.81 | A | A | A | 0.25 |
| 134 | 63.62 | 6.62 | 4.47 | 11.31 | A | A | A | 0.68 |
| 135 | 64.1 | 3.9 | 5.25 | 7.53 | A | A | A | 0.80 |
| 137 | 61.6 | 2.39 | 1.15 | 5.89 | A | A | A | 0.18 |
| 139 | 58.066 | 1.4028 | -4.65 | 5.05 | A | A | A | 0.71 |
| 141 | 62.1 | 3.1 | 1.97 | 6.68 | A | A | A | 0.30 |
| 144 | 62.3 | 8.6 | 2.30 | 14.50 | A | A | A | 0.35 |
| 146 | 59.5 | 6.0 | -2.30 | 11.02 | A | A | A | 0.35 |
| 147 | 59.2 | 5.4 | -2.79 | 10.14 | A | A | A | 0.43 |
| 148 | 60.66 | 7.65 | -0.39 | 13.37 | A | A | A | 0.06 |
| 149 | 62.6 | 1.5 | 2.79 | 5.04 | A | A | A | 0.43 |
| 150 | 66.9 | 2.1 | 9.85 | 5.43 | A | A | A | 1.51 |
| 151 | 61.0 | 3.0 | 0.16 | 6.62 | A | A | A | 0.03 |
| 152 | 60.3 | 2.5 | -0.99 | 6.07 | A | A | A | 0.15 |
| 153 | 63.0 | 4.5 | 3.45 | 8.41 | A | A | A | 0.53 |
| 154 | 48.0 | 2.6 | -21.18 | 7.00 | N | N | N | 3.24 |
| 155 | 68.3 | 3.6 | 12.15 | 6.89 | A | A | A | 1.86 |
| 156 | 62.5 | 4.5 | 2.63 | 8.46 | A | A | A | 0.40 |
| 157 | 59.79 | 3.72 | -1.82 | 7.64 | A | A | A | 0.28 |
| 159 | 56.8 | 1.5 | -6.73 | 5.16 | A | A | A | 1.03 |
| 160 | 62.8 | 2.5 | 3.12 | 5.96 | A | A | A | 0.48 |
| 161 | 53.8 | 2.3 | -11.66 | 6.16 | A | A | A | 1.78 |
| 162 | 62.7 | 3.1 | 2.96 | 6.64 | A | A | A | 0.45 |
| 163 | 57.4 | 5.7 | -5.75 | 10.88 | A | A | A | 0.88 |
| 164 | 64 | 3 | 5.09 | 6.45 | A | A | A | 0.78 |
| 165 | 65.5 | 4.7 | 7.55 | 8.43 | A | A | A | 1.16 |
| 166 | 64.5 | 3.5 | 5.91 | 7.01 | A | A | A | 0.90 |
| 168 | 52.69 | 5.41 | -13.48 | 11.18 | A | A | A | 2.06 |
| 169 | 62.09 | 0.83 | 1.95 | 4.63 | A | A | A | 0.30 |
| 171 | 57.28 | 5.30 | -5.94 | 10.26 | A | A | A | 0.91 |
| 172 | 79.7 | 11.1 | 30.87 | 14.62 | N | A | N | 4.72 |
| 177 | 70 | 5 | 14.94 | 8.41 | A | A | A | 2.29 |
| 179 | 61.5 | 1.5 | 0.99 | 5.06 | A | A | A | 0.15 |
| 180 | 62.78 | 3.1 | 3.09 | 6.64 | A | A | A | 0.47 |
| 181 | 61.179 | 3.611 | 0.46 | 7.38 | A | A | A | 0.07 |
| 182 | 85.51 | 5.19 | 40.41 | 7.52 | N | N | N | 6.18 |
| 183 | 61.84 | 2.1 | 1.54 | 5.58 | A | A | A | 0.24 |
| 184 | 61.68 | 0.50 | 1.28 | 4.51 | A | A | A | 0.20 |
| 185 | 61.6 | 4.6 | 1.15 | 8.68 | A | A | A | 0.18 |
| 186 | 60.1 | 3.7 | -1.31 | 7.59 | A | A | A | 0.20 |
| 187 | 88.1 | 2.2 | 44.66 | 5.09 | N | N | N | 6.83 |
| 189 | 68.0 | 6.9 | 11.66 | 11.07 | A | A | A | 1.78 |
| 190 | 63.9 | 7.0 | 4.93 | 11.82 | A | A | A | 0.75 |
| 191 | 63.3 | 2.5 | 3.94 | 5.94 | A | A | A | 0.60 |
| 192 | 0.023 | 0.003 | -99.96 | 13.78 | N | N | N | 15.30 |
| 193 | 57.51 | 8.05 | -5.57 | 14.68 | A | A | A | 0.85 |

TABLE 8. Evaluation Results for Am-241 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 195 | 63.6 | 3.40 | 4.43 | 6.95 | A | A | A | 0.68 |
| 196 | 60.7 | 4.3 | -0.33 | 8.36 | A | A | A | 0.05 |
| 197 | 62.9 | 2.5 | 3.28 | 5.95 | A | A | A | 0.50 |
| 198 | 61.04 | 3.67 | 0.23 | 7.47 | A | A | A | 0.04 |
| 199 | 79.52 | 1.75 | 30.57 | 4.95 | N | N | N | 4.68 |
| 205 | 63.7 | 2.3 | 4.60 | 5.72 | A | A | A | 0.70 |
| 206 | 63.5 | 6.3 | 4.27 | 10.87 | A | A | A | 0.65 |
| 207 | 59.8 | 3.5 | -1.81 | 7.34 | A | A | A | 0.28 |
| 208 | 58.4 | 2.6 | -4.11 | 6.28 | A | A | A | 0.63 |
| 209 | 61.9 | 4.6 | 1.64 | 8.65 | A | A | A | 0.25 |
| 211 | 62.5 | 5.4 | 2.63 | 9.71 | A | A | A | 0.40 |
| 212 | 60.8 | 2.4 | -0.16 | 5.94 | A | A | A | 0.03 |
| 213 | 62.3 | 3.8 | 2.30 | 7.54 | A | A | A | 0.35 |
| 216 | 61.45 | 2.98 | 0.90 | 6.57 | A | A | A | 0.14 |
| 219 | 62.3 | 14.7 | 2.30 | 24.01 | A | N | W | 0.35 |
| 222 | 30.9 | 2.16 | -49.26 | 8.28 | N | N | N | 7.54 |
| 223 | 61.1 | 2.8 | 0.33 | 6.38 | A | A | A | 0.05 |
| 224 | 62.2 | 3.4 | 2.13 | 7.04 | A | A | A | 0.33 |
| 225 | 60 | 3 | -1.48 | 6.68 | A | A | A | 0.23 |
| 227 | 51.83 | 2.3 | -14.89 | 6.27 | A | A | A | 2.28 |
| 228 | 60.0 | 2.4 | -1.48 | 5.97 | A | A | A | 0.23 |
| 230 | 62.35 | 2.4 | 2.38 | 5.87 | A | A | A | 0.36 |
| 231 | 56.24 | 12.14 | -7.65 | 22.04 | A | N | W | 1.17 |
| 232 | 62.6082 | 3.5978 | 2.80 | 7.26 | A | A | A | 0.43 |
| 233 | 60.36 | 0.99 | -0.89 | 4.73 | A | A | A | 0.14 |
| 234 | 77.16 | 1.93 | 26.70 | 5.09 | N | N | N | 4.09 |
| 235 | 80.58 | 6.14 | 32.32 | 8.82 | N | N | N | 4.94 |
| 237 | 63.4 | 6.4 | 4.11 | 11.03 | A | A | A | 0.63 |
| 238 | 60.96 | 7.924 | 0.10 | 13.73 | A | A | A | 0.02 |
| 239 | 62.5 | 2.3 | 2.63 | 5.76 | A | A | A | 0.40 |
| 240 | 66.6 | 2.3 | 9.36 | 5.62 | A | A | A | 1.43 |
| 241 | 67.7 | 5.2 | 11.17 | 8.87 | A | A | A | 1.71 |
| 242 | 62.3 | 3.2 | 2.30 | 6.79 | A | A | A | 0.35 |
| 246 | 45.26 | 3.90 | -25.68 | 9.69 | N | N | N | 3.93 |
| 248 | 57.80 | 1.45 | -5.09 | 5.09 | A | A | A | 0.78 |
| 249 | 72.1 | 7.6 | 18.39 | 11.44 | A | A | A | 2.81 |
| 250 | 44.5 | 2.1 | -26.93 | 6.48 | N | N | N | 4.12 |
| 251 | 64.9 | 1.4 | 6.57 | 4.93 | A | A | A | 1.01 |
| 254 | 62.3 | 4.7 | 2.30 | 8.75 | A | A | A | 0.35 |
| 255 | 60.6 | 2.0 | -0.49 | 5.53 | A | A | A | 0.08 |
| 256 | 59.79 | 0.93 | -1.82 | 4.70 | A | A | A | 0.28 |
| 257 | 78 | 3 | 28.08 | 5.87 | N | N | N | 4.30 |
| 258 | 61.9 | 3.01 | 1.64 | 6.58 | A | A | A | 0.25 |
| 259 | 52.48 | 0.14 | -13.83 | 4.44 | A | N | W | 2.12 |
| 261 | 61.6 | 4.7 | 1.15 | 8.82 | A | A | A | 0.18 |
| 262 | 65.0 | 5.5 | 6.73 | 9.55 | A | A | A | 1.03 |
| 263 | 59.32 | 1.77 | -2.59 | 5.34 | A | A | A | 0.40 |
| 266 | 58.1 | 5.9 | -4.60 | 11.08 | A | A | A | 0.70 |
| 267 | 58.51 | 1.02 | -3.92 | 4.76 | A | A | A | 0.60 |
| 268 | 61.9 | 2.5 | 1.64 | 6.00 | A | A | A | 0.25 |
| 270 | 62.4 | 4.46 | 2.46 | 8.41 | A | A | A | 0.38 |
| 274 | 60.3 | 5.1 | -0.99 | 9.55 | A | A | A | 0.15 |
| 275 | 63 | 3.94 | 3.45 | 7.67 | A | A | A | 0.53 |
| 279 | 65.4 | 6.6 | 7.39 | 11.02 | A | A | A | 1.13 |

TABLE 8. Evaluation Results for Am-241 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 281 | 51 | 2 | -16.26 | 5.92 | A | N | W | 2.49 |
| 283 | 64.9 | 3.4 | 6.57 | 6.86 | A | A | A | 1.01 |
| 287 | 63 | 0.6 | 3.45 | 4.53 | A | A | A | 0.53 |
| 289 | 66.5 | 3.5 | 9.20 | 6.88 | A | A | A | 1.41 |
| 290 | 67.90 | 2.13 | 11.49 | 5.43 | A | A | A | 1.76 |
| 291 | 63.3 | 2.3 | 3.94 | 5.73 | A | A | A | 0.60 |
| 295 | 467.0 | 41.5 | 666.83 | 9.93 | N | N | N | 102.04 |
| 298 | 59.4 | 1.2 | -2.46 | 4.87 | A | A | A | 0.38 |
| 299 | 46.6 | 3.0 | -23.48 | 7.82 | N | N | N | 3.59 |
| 304 | 66.48 | 5.32 | 9.16 | 9.15 | A | A | A | 1.40 |
| 310 | 65.4 | 1.5 | 7.39 | 4.99 | A | A | A | 1.13 |
| 312 | 56.64 | 1.04 | -7.00 | 4.80 | A | A | A | 1.07 |
| 313 | 69.4 | 3.6 | 13.96 | 6.82 | A | A | A | 2.14 |
| 314 | 67.72 | 3.24 | 11.20 | 6.52 | A | A | A | 1.71 |
| 316 | 64.826 | 12.317 | 6.45 | 19.51 | A | A | A | 0.99 |
| 317 | 64.6 |  | 6.08 |  | A |  |  | 0.93 |
| 320 | 67.52 | 3.71 | 10.87 | 7.06 | A | A | A | 1.66 |
| 321 | 62.89 | 0.81 | 3.27 | 4.62 | A | A | A | 0.50 |
| 322 | 55.0 | 8.4 | -9.69 | 15.90 | A | A | A | 1.48 |
| 323 | 59.7 | 1.6 | -1.97 | 5.18 | A | A | A | 0.30 |
| 324 | 62.97 | 2.41 | 3.40 | 5.86 | A | A | A | 0.52 |
| 325 | 62.6 | 5.5 | 2.79 | 9.84 | A | A | A | 0.43 |

## Reported Results for Sr-90 Determination in Sample 2

TABLE 9. Target values

| Parameter | Value |
| --- | --- |
| Target Value, Bq/kg | 146.8 |
| UNC, Bq/kg | 8.4 |
| MARB, % | 30 |

FIGURE 9. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

FIGURE 10. Reported results and evaluation parameters for Sr-90. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

TABLE 10. Evaluation Results for Sr-90 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 5 | 151 | 15 | 2.86 | 11.46 | A | A | A | 0.39 |
| 6 | 145.1 | 7.3 | -1.16 | 7.62 | A | A | A | 0.16 |
| 9 | 156 | 19 | 6.27 | 13.46 | A | A | A | 0.85 |
| 10 | 147.0 | 5.9 | 0.14 | 6.99 | A | A | A | 0.02 |
| 11 | 147.6 | 7.4 | 0.54 | 7.61 | A | A | A | 0.07 |
| 14 | 138.5 | 4.5 | -5.65 | 6.58 | A | A | A | 0.77 |
| 18 | 147 | 13 | 0.14 | 10.53 | A | A | A | 0.02 |
| 19 | 141 | 10.6 | -3.95 | 9.45 | A | A | A | 0.54 |
| 25 | 143.3 | 7.3 | -2.38 | 7.66 | A | A | A | 0.32 |
| 26 | 143.9 | 23.9 | -1.98 | 17.57 | A | A | A | 0.27 |
| 32 | 153.4 | 19.5 | 4.50 | 13.94 | A | A | A | 0.61 |
| 34 | 141.5 | 12.0 | -3.61 | 10.23 | A | A | A | 0.49 |
| 35 | 160 | 12 | 8.99 | 9.43 | A | A | A | 1.22 |
| 39 | 123 | 4 | -16.21 | 6.58 | A | A | A | 2.21 |
| 41 | 117.0 | 6.5 | -20.30 | 7.98 | A | A | A | 2.76 |
| 50 | 10.95 | 2.92 | -92.54 | 27.27 | N | N | N | 12.60 |
| 58 | 148.00 | 7.45 | 0.82 | 7.62 | A | A | A | 0.11 |
| 59 | 140 | 7.0 | -4.63 | 7.60 | A | A | A | 0.63 |
| 60 | 146.3 | 18.1 | -0.34 | 13.63 | A | A | A | 0.05 |
| 61 | 82.35 | 0.73 | -43.90 | 5.79 | N | N | N | 5.98 |
| 76 | 147.0 | 6.2 | 0.14 | 7.11 | A | A | A | 0.02 |
| 79 | 140 | 6 | -4.63 | 7.15 | A | A | A | 0.63 |
| 81 | 152.5 | 16.7 | 3.88 | 12.36 | A | A | A | 0.53 |
| 84 | 172 | 22 | 17.17 | 14.01 | A | A | A | 2.34 |
| 86 | 142.7 | 16.6 | -2.79 | 12.96 | A | A | A | 0.38 |
| 88 | 145 | 7 | -1.23 | 7.49 | A | A | A | 0.17 |
| 91 | 113 | 29 | -23.02 | 26.29 | A | A | A | 3.14 |
| 92 | 149 | 18 | 1.50 | 13.37 | A | A | A | 0.20 |
| 93 | 130.5 | 2.4 | -11.10 | 6.01 | A | A | A | 1.51 |
| 99 | 163.0 | 6.7 | 11.04 | 7.05 | A | A | A | 1.50 |

TABLE 10. Evaluation Results for Sr-90 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 102 | 154.3 | 2.2 | 5.11 | 5.90 | A | A | A | 0.70 |
| 105 | 147.2 | 1.1 | 0.27 | 5.77 | A | A | A | 0.04 |
| 107 | 145.3 | 0.3 | -1.02 | 5.73 | A | A | A | 0.14 |
| 111 | 128 | 19 | -12.81 | 15.91 | A | A | A | 1.74 |
| 112 | 66.26 | 2.9109 | -54.86 | 7.21 | N | N | N | 7.47 |
| 116 | 136.2 | 10.3 | -7.22 | 9.48 | A | A | A | 0.98 |
| 123 | 138 | 9 | -5.99 | 8.68 | A | A | A | 0.82 |
| 132 | 138 | 16 | -5.99 | 12.93 | A | A | A | 0.82 |
| 137 | 158.7 | 6.44 | 8.11 | 7.01 | A | A | A | 1.10 |
| 141 | 154 | 5.6 | 4.90 | 6.78 | A | A | A | 0.67 |
| 143 | 127 | 5.4 | -13.49 | 7.13 | A | A | A | 1.84 |
| 146 | 161 | 28 | 9.67 | 18.31 | A | A | A | 1.32 |
| 148 | 147.5 | 1.93 | 0.48 | 5.87 | A | A | A | 0.06 |
| 150 | 144.6 | 2.1 | -1.50 | 5.90 | A | A | A | 0.20 |
| 151 | 151.0 | 12.0 | 2.86 | 9.79 | A | A | A | 0.39 |
| 152 | 151 | 8 | 2.86 | 7.80 | A | A | A | 0.39 |
| 154 | 144.0 | 8.9 | -1.91 | 8.42 | A | A | A | 0.26 |
| 155 | 149.5 | 6.9 | 1.84 | 7.35 | A | A | A | 0.25 |
| 156 | 148 | 7 | 0.82 | 7.42 | A | A | A | 0.11 |
| 162 | 154.37 | 15.44 | 5.16 | 11.52 | A | A | A | 0.70 |
| 166 | 144.1 | 6.3 | -1.84 | 7.20 | A | A | A | 0.25 |
| 171 | 121.43 | 12.00 | -17.28 | 11.42 | A | A | A | 2.35 |
| 172 | 103.1 | 20.0 | -29.77 | 20.22 | A | A | A | 4.05 |
| 176 | 140 | 5.10 | -4.63 | 6.78 | A | A | A | 0.63 |
| 180 | 142.5 | 2.5 | -2.93 | 5.98 | A | A | A | 0.40 |
| 181 | 141.41 | 0.01 | -3.67 | 5.72 | A | A | A | 0.50 |
| 182 | 143.14 | 0.42 | -2.49 | 5.73 | A | A | A | 0.34 |
| 185 | 145 | 9 | -1.23 | 8.44 | A | A | A | 0.17 |
| 186 | 145.3 | 5.9 | -1.02 | 7.02 | A | A | A | 0.14 |
| 189 | 193.9 | 19.5 | 32.08 | 11.57 | N | N | N | 4.37 |
| 190 | 140 | 15 | -4.63 | 12.15 | A | A | A | 0.63 |
| 195 | 131.7 | 10.0 | -10.29 | 9.51 | A | A | A | 1.40 |
| 197 | 145 | 6 | -1.23 | 7.06 | A | A | A | 0.17 |
| 200 | 142.89 | 5.27 | -2.66 | 6.81 | A | A | A | 0.36 |
| 201 | 146 | 15 | -0.54 | 11.76 | A | A | A | 0.07 |
| 210 | 143 | 11 | -2.59 | 9.59 | A | A | A | 0.35 |
| 211 | 142.4 | 6.7 | -3.00 | 7.41 | A | A | A | 0.41 |
| 212 | 146 | 6 | -0.54 | 7.04 | A | A | A | 0.07 |
| 213 | 146 | 10 | -0.54 | 8.92 | A | A | A | 0.07 |
| 214 | 47.9 | 5.03 | -67.37 | 11.96 | N | N | N | 9.17 |
| 216 | 139.4 | 1.83 | -5.04 | 5.87 | A | A | A | 0.69 |
| 219 | 149.9 | 5.40 | 2.11 | 6.76 | A | A | A | 0.29 |
| 220 | 150 | 8 | 2.18 | 7.82 | A | A | A | 0.30 |
| 227 | 135 | 4 | -8.04 | 6.44 | A | A | A | 1.09 |
| 233 | 116 | 4.64 | -20.98 | 6.98 | A | N | W | 2.86 |
| 238 | 167.21 | 21.74 | 13.90 | 14.21 | A | A | A | 1.89 |
| 240 | 143 | 5 | -2.59 | 6.71 | A | A | A | 0.35 |
| 242 | 142 | 5 | -3.27 | 6.72 | A | A | A | 0.45 |
| 243 | 153.3 | 11.1 | 4.43 | 9.23 | A | A | A | 0.60 |
| 245 | 103.4 | 6.9 | -29.56 | 8.79 | A | N | W | 4.03 |
| 249 | 131 | 1.5 | -10.76 | 5.84 | A | A | A | 1.47 |
| 251 | 141 | 13 | -3.95 | 10.85 | A | A | A | 0.54 |
| 254 | 132 | 28 | -10.08 | 21.97 | A | A | A | 1.37 |
| 255 | 147.7 | 0.3 | 0.61 | 5.73 | A | A | A | 0.08 |

TABLE 10. Evaluation Results for Sr-90 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 256 | 141.100 | 0.642 | -3.88 | 5.74 | A | A | A | 0.53 |
| 260 | 151.4 | 5.2 | 3.13 | 6.67 | A | A | A | 0.43 |
| 261 | 156 | 7.9 | 6.27 | 7.64 | A | A | A | 0.85 |
| 268 | 144.4 | 17.4 | -1.63 | 13.34 | A | A | A | 0.22 |
| 270 | 147 | 2.9 | 0.14 | 6.05 | A | A | A | 0.02 |
| 283 | 139 | 6.3 | -5.31 | 7.30 | A | A | A | 0.72 |
| 286 | 141 | 10 | -3.95 | 9.11 | A | A | A | 0.54 |
| 289 | 140 | 5 | -4.63 | 6.75 | A | A | A | 0.63 |
| 291 | 141.9 | 7.5 | -3.34 | 7.79 | A | A | A | 0.45 |
| 305 | 22.2 | 0.53 | -84.88 | 6.20 | N | N | N | 11.56 |
| 306 | 1441.35 | 56.35 | 881.85 | 6.93 | N | N | N | 120.09 |
| 310 | 147.3 | 4.6 | 0.34 | 6.52 | A | A | A | 0.05 |
| 313 | 121.4 | 0.2 | -17.30 | 5.72 | A | N | W | 2.36 |
| 320 | 143.59 | 12.92 | -2.19 | 10.66 | A | A | A | 0.30 |
| 322 | 139.9 | 11.1 | -4.70 | 9.78 | A | A | A | 0.64 |
| 324 | 136.82 | 17.06 | -6.80 | 13.72 | A | A | A | 0.93 |

## Reported Results for Ba-133 Determination in Sample 2

TABLE 11. Target values

| Parameter | Value |
| --- | --- |
| Target Value, Bq/kg | 147.3 |
| UNC, Bq/kg | 6.6 |
| MARB, % | 20 |

FIGURE 11. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

FIGURE 12. Reported results and evaluation parameters for Ba-133. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

TABLE 12. Evaluation Results for Ba-133 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 129.97 | 3.51 | -11.77 | 5.23 | A | A | A | 2.42 |
| 2 | 143.3 | 7.2 | -2.72 | 6.73 | A | A | A | 0.56 |
| 3 | 139.3 | 7.9 | -5.43 | 7.23 | A | A | A | 1.12 |
| 4 | 142.2 | 8.2 | -3.46 | 7.30 | A | A | A | 0.71 |
| 5 | 159.3 | 12.02 | 8.15 | 8.78 | A | A | A | 1.67 |
| 6 | 127.7 | 11.6 | -13.31 | 10.13 | A | A | A | 2.73 |
| 7 | 139.28 | 6.21 | -5.44 | 6.32 | A | A | A | 1.12 |
| 9 | 145 | 16 | -1.56 | 11.91 | A | A | A | 0.32 |
| 10 | 121.9 | 0.90 | -17.24 | 4.54 | A | N | W | 3.54 |
| 11 | 147.5 | 5.0 | 0.14 | 5.62 | A | A | A | 0.03 |
| 12 | 145.6 | 6.4 | -1.15 | 6.28 | A | A | A | 0.24 |
| 13 | 143.51 | 5.64 | -2.57 | 5.96 | A | A | A | 0.53 |
| 14 | 148.1 | 10.5 | 0.54 | 8.39 | A | A | A | 0.11 |
| 15 | 152.8 | 7.7 | 3.73 | 6.74 | A | A | A | 0.77 |
| 16 | 146.1 | 4.3 | -0.81 | 5.36 | A | A | A | 0.17 |
| 17 | 141.53 | 5.15 | -3.92 | 5.77 | A | A | A | 0.80 |
| 18 | 152 | 13 | 3.19 | 9.66 | A | A | A | 0.66 |
| 19 | 145.9 | 13.7 | -0.95 | 10.40 | A | A | A | 0.20 |
| 20 | 168 | 7 | 14.05 | 6.12 | A | A | A | 2.89 |
| 21 | 140 | 10 | -4.96 | 8.43 | A | A | A | 1.02 |
| 22 | 151.8 | 2.7 | 3.05 | 4.82 | A | A | A | 0.63 |
| 25 | 145 | 7 | -1.56 | 6.59 | A | A | A | 0.32 |
| 26 | 136.63 | 7.36 | -7.24 | 7.01 | A | A | A | 1.49 |
| 27 | 150 | 16 | 1.83 | 11.57 | A | A | A | 0.38 |
| 29 | 132.46 | 2.80 | -10.07 | 4.95 | A | A | A | 2.07 |
| 31 | 130.89 | 7.53 | -11.14 | 7.29 | A | A | A | 2.29 |
| 32 | 145.0 | 16.0 | -1.56 | 11.91 | A | A | A | 0.32 |
| 33 | 145 | 7 | -1.56 | 6.59 | A | A | A | 0.32 |
| 34 | 146.0 | 6.2 | -0.88 | 6.17 | A | A | A | 0.18 |
| 35 | 128 | 6.6 | -13.10 | 6.83 | A | A | A | 2.69 |

TABLE 12. Evaluation Results for Ba-133 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 36 | 145.8 | 6.8 | -1.02 | 6.47 | A | A | A | 0.21 |
| 38 | 155 | 8 | 5.23 | 6.83 | A | A | A | 1.07 |
| 39 | 164.2 | 3.6 | 11.47 | 4.99 | A | A | A | 2.36 |
| 41 | 144.0 | 6.0 | -2.24 | 6.12 | A | A | A | 0.46 |
| 42 | 110.43 | 1.83 | -25.03 | 4.78 | N | N | N | 5.14 |
| 43 | 144.65 | 3.33 | -1.80 | 5.04 | A | A | A | 0.37 |
| 46 | 243.86 | 4.78 | 65.55 | 4.89 | N | N | N | 13.47 |
| 47 | 138 | 3.9 | -6.31 | 5.30 | A | A | A | 1.30 |
| 48 | 148.1 | 5.3 | 0.54 | 5.73 | A | A | A | 0.11 |
| 50 | 137.59 | 5.55 | -6.59 | 6.03 | A | A | A | 1.35 |
| 51 | 132 | 8 | -10.39 | 7.54 | A | A | A | 2.13 |
| 52 | 120.3 | 5.1 | -18.33 | 6.17 | A | N | W | 3.77 |
| 53 | 132.6 | 5.4 | -9.98 | 6.05 | A | A | A | 2.05 |
| 55 | 137.1 | 2.6 | -6.92 | 4.87 | A | A | A | 1.42 |
| 56 | 148.8 | 3.87 | 1.02 | 5.18 | A | A | A | 0.21 |
| 59 | 145 | 10.2 | -1.56 | 8.34 | A | A | A | 0.32 |
| 61 | 137.53 | 4.35 | -6.63 | 5.48 | A | A | A | 1.36 |
| 63 | 152.0 | 5.1 | 3.19 | 5.60 | A | A | A | 0.66 |
| 64 | 140.2 | 8 | -4.82 | 7.26 | A | A | A | 0.99 |
| 66 | 166 | 5.8 | 12.70 | 5.68 | A | A | A | 2.61 |
| 68 | 156.2 | 8.8 | 6.04 | 7.20 | A | A | A | 1.24 |
| 69 | 138.41 | 2.04 | -6.04 | 4.72 | A | A | A | 1.24 |
| 70 | 145.80 | 2.07 | -1.02 | 4.70 | A | A | A | 0.21 |
| 72 | 149.674 | 10.033 | 1.61 | 8.06 | A | A | A | 0.33 |
| 74 | 142.8 | 5.5 | -3.05 | 5.91 | A | A | A | 0.63 |
| 75 | 157.9 | 4.8 | 7.20 | 5.41 | A | A | A | 1.48 |
| 76 | 139.00 | 1.27 | -5.63 | 4.57 | A | A | A | 1.16 |
| 77 | 192.1 | 2.4 | 30.41 | 4.65 | N | N | N | 6.25 |
| 80 | 161.5 | 10.3 | 9.64 | 7.79 | A | A | A | 1.98 |
| 81 | 147.0 | 10.3 | -0.20 | 8.32 | A | A | A | 0.04 |
| 84 | 141.15 | 2.88 | -4.18 | 4.92 | A | A | A | 0.86 |
| 85 | 150.28 | 12.33 | 2.02 | 9.35 | A | A | A | 0.42 |
| 86 | 142 | 2.7 | -3.60 | 4.87 | A | A | A | 0.74 |
| 87 | 145 | 5 | -1.56 | 5.65 | A | A | A | 0.32 |
| 88 | 143.7 | 9.5 | -2.44 | 7.99 | A | A | A | 0.50 |
| 92 | 150 | 10 | 1.83 | 8.03 | A | A | A | 0.38 |
| 93 | 129.34 | 2.30 | -12.19 | 4.82 | A | A | A | 2.50 |
| 94 | 139.31 | 1.74 | -5.42 | 4.65 | A | A | A | 1.11 |
| 95 | 141.35 | 1.77 | -4.04 | 4.65 | A | A | A | 0.83 |
| 96 | 139.2 | 8.4 | -5.50 | 7.52 | A | A | A | 1.13 |
| 98 | 146.0 | 4.3 | -0.88 | 5.36 | A | A | A | 0.18 |
| 99 | 152 | 5.1 | 3.19 | 5.60 | A | A | A | 0.66 |
| 101 | 146.9 | 4.6 | -0.27 | 5.47 | A | A | A | 0.06 |
| 102 | 130.7 | 4.1 | -11.27 | 5.47 | A | A | A | 2.32 |
| 103 | 146.12 | 3.6 | -0.80 | 5.11 | A | A | A | 0.16 |
| 105 | 151.4 | 7.9 | 2.78 | 6.88 | A | A | A | 0.57 |
| 106 | 144.2 | 0.6 | -2.10 | 4.50 | A | A | A | 0.43 |
| 107 | 152 | 5 | 3.19 | 5.56 | A | A | A | 0.66 |
| 108 | 143.0811 | 0.7052 | -2.86 | 4.51 | A | A | A | 0.59 |
| 109 | 143.4 | 5.6 | -2.65 | 5.94 | A | A | A | 0.54 |
| 111 | 145.2 | 9.6 | -1.43 | 7.99 | A | A | A | 0.29 |
| 113 | 146.0 | 7.3 | -0.88 | 6.71 | A | A | A | 0.18 |
| 116 | 143.9 | 1.1 | -2.31 | 4.55 | A | A | A | 0.47 |
| 118 | 135.7 | 13.1 | -7.88 | 10.64 | A | A | A | 1.62 |

TABLE 12. Evaluation Results for Ba-133 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 120 | 152 | 6 | 3.19 | 5.97 | A | A | A | 0.66 |
| 121 | 144 | 12 | -2.24 | 9.46 | A | A | A | 0.46 |
| 123 | 127 | 8 | -13.78 | 7.73 | A | A | A | 2.83 |
| 124 | 143.6 | 4.5 | -2.51 | 5.47 | A | A | A | 0.52 |
| 125 | 132 | 2.3 | -10.39 | 4.81 | A | A | A | 2.13 |
| 126 | 129 | 11.8 | -12.42 | 10.19 | A | A | A | 2.55 |
| 129 | 144 | 6 | -2.24 | 6.12 | A | A | A | 0.46 |
| 130 | 144 | 5 | -2.24 | 5.67 | A | A | A | 0.46 |
| 132 | 126 | 7.6 | -14.46 | 7.51 | A | A | A | 2.97 |
| 133 | 140.8 | 7.1 | -4.41 | 6.75 | A | A | A | 0.91 |
| 134 | 102.31 | 5.09 | -30.54 | 6.70 | N | N | N | 6.27 |
| 135 | 144.2 | 5.5 | -2.10 | 5.88 | A | A | A | 0.43 |
| 137 | 116.6 | 2.14 | -20.84 | 4.84 | N | N | N | 4.28 |
| 144 | 140.8 | 12.8 | -4.41 | 10.14 | A | A | A | 0.91 |
| 146 | 154.3 | 6.3 | 4.75 | 6.06 | A | A | A | 0.98 |
| 147 | 146 | 13 | -0.88 | 9.97 | A | A | A | 0.18 |
| 148 | 150.83 | 14.26 | 2.40 | 10.46 | A | A | A | 0.49 |
| 149 | 158.3 | 3.3 | 7.47 | 4.94 | A | A | A | 1.53 |
| 150 | 137.2 | 4.3 | -6.86 | 5.47 | A | A | A | 1.41 |
| 151 | 141.0 | 6.0 | -4.28 | 6.18 | A | A | A | 0.88 |
| 152 | 134 | 5 | -9.03 | 5.83 | A | A | A | 1.85 |
| 153 | 146 | 11 | -0.88 | 8.77 | A | A | A | 0.18 |
| 154 | 143.4 | 2.5 | -2.65 | 4.81 | A | A | A | 0.54 |
| 155 | 148.3 | 6.9 | 0.68 | 6.46 | A | A | A | 0.14 |
| 156 | 141 | 9 | -4.28 | 7.80 | A | A | A | 0.88 |
| 157 | 139.82 | 7.53 | -5.08 | 7.01 | A | A | A | 1.04 |
| 158 | 161.1 | 14.2 | 9.37 | 9.89 | A | A | A | 1.92 |
| 159 | 129.3 | 2.6 | -12.22 | 4.91 | A | A | A | 2.51 |
| 160 | 126 | 4 | -14.46 | 5.49 | A | N | W | 2.97 |
| 161 | 145.3 | 4.1 | -1.36 | 5.30 | A | A | A | 0.28 |
| 162 | 151.6 | 2.4 | 2.92 | 4.75 | A | A | A | 0.60 |
| 163 | 146.0 | 6.4 | -0.88 | 6.27 | A | A | A | 0.18 |
| 164 | 145 | 6 | -1.56 | 6.10 | A | A | A | 0.32 |
| 165 | 145.4 | 8.7 | -1.29 | 7.48 | A | A | A | 0.26 |
| 166 | 142.2 | 7.9 | -3.46 | 7.14 | A | A | A | 0.71 |
| 168 | 148.66 | 9.03 | 0.92 | 7.55 | A | A | A | 0.19 |
| 169 | 145.44 | 1.52 | -1.26 | 4.60 | A | A | A | 0.26 |
| 171 | 145.89 | 13.20 | -0.96 | 10.10 | A | A | A | 0.20 |
| 172 | 145.9 | 20.2 | -0.95 | 14.55 | A | A | A | 0.20 |
| 173 | 139.0 | 5.0 | -5.63 | 5.75 | A | A | A | 1.16 |
| 176 | 142 | 8.96 | -3.60 | 7.74 | A | A | A | 0.74 |
| 177 | 139 | 14 | -5.63 | 11.02 | A | A | A | 1.16 |
| 179 | 147.7 | 1.5 | 0.27 | 4.59 | A | A | A | 0.06 |
| 180 | 147.89 | 5.01 | 0.40 | 5.62 | A | A | A | 0.08 |
| 181 | 127.720 | 3.515 | -13.29 | 5.26 | A | A | A | 2.73 |
| 182 | 142.98 | 7.94 | -2.93 | 7.14 | A | A | A | 0.60 |
| 184 | 149.52 | 0.70 | 1.51 | 4.51 | A | A | A | 0.31 |
| 185 | 146.4 | 8.8 | -0.61 | 7.50 | A | A | A | 0.13 |
| 186 | 131.0 | 7.9 | -11.07 | 7.51 | A | A | A | 2.27 |
| 189 | 144.0 | 5.9 | -2.24 | 6.07 | A | A | A | 0.46 |
| 190 | 147 | 17 | -0.20 | 12.40 | A | A | A | 0.04 |
| 191 | 144.9 | 6.7 | -1.63 | 6.44 | A | A | A | 0.33 |
| 193 | 143.4 | 14.34 | -2.65 | 10.96 | A | A | A | 0.54 |
| 195 | 148.1 | 5.50 | 0.54 | 5.82 | A | A | A | 0.11 |

TABLE 12. Evaluation Results for Ba-133 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 196 | 149 | 12 | 1.15 | 9.22 | A | A | A | 0.24 |
| 197 | 145 | 6 | -1.56 | 6.10 | A | A | A | 0.32 |
| 198 | 137.47 | 7.38 | -6.67 | 6.99 | A | A | A | 1.37 |
| 199 | 128.20 | 2.70 | -12.97 | 4.95 | A | N | W | 2.66 |
| 200 | 147.85 | 2.99 | 0.37 | 4.92 | A | A | A | 0.08 |
| 201 | 147 | 15 | -0.20 | 11.14 | A | A | A | 0.04 |
| 202 | 140.77 | 8.16 | -4.43 | 7.33 | A | A | A | 0.91 |
| 205 | 144 | 7 | -2.24 | 6.61 | A | A | A | 0.46 |
| 206 | 148.4 | 8.1 | 0.75 | 7.06 | A | A | A | 0.15 |
| 207 | 134 | 7.0 | -9.03 | 6.88 | A | A | A | 1.85 |
| 208 | 148.28 | 9.61 | 0.67 | 7.88 | A | A | A | 0.14 |
| 209 | 146 | 7.6 | -0.88 | 6.87 | A | A | A | 0.18 |
| 211 | 147 | 13 | -0.20 | 9.91 | A | A | A | 0.04 |
| 212 | 145 | 4 | -1.56 | 5.26 | A | A | A | 0.32 |
| 213 | 144 | 9 | -2.24 | 7.69 | A | A | A | 0.46 |
| 214 | 157.1 | 10.9 | 6.65 | 8.26 | A | A | A | 1.37 |
| 215 | 137 | 12.5 | -6.99 | 10.16 | A | A | A | 1.44 |
| 216 | 145.6 | 8.79 | -1.15 | 7.52 | A | A | A | 0.24 |
| 219 | 146 | 7.8 | -0.88 | 6.97 | A | A | A | 0.18 |
| 221 | 146.98 | 4.35 | -0.22 | 5.37 | A | A | A | 0.04 |
| 222 | 122 | 3.78 | -17.18 | 5.45 | A | N | W | 3.53 |
| 223 | 142.8 | 7.5 | -3.05 | 6.90 | A | A | A | 0.63 |
| 224 | 143.4 | 3.1 | -2.65 | 4.97 | A | A | A | 0.54 |
| 225 | 146 | 7 | -0.88 | 6.56 | A | A | A | 0.18 |
| 227 | 147.06 | 4.15 | -0.16 | 5.30 | A | A | A | 0.03 |
| 228 | 147.8 | 5.2 | 0.34 | 5.70 | A | A | A | 0.07 |
| 230 | 149 | 6 | 1.15 | 6.02 | A | A | A | 0.24 |
| 231 | 161.3 | 32.74 | 9.50 | 20.79 | A | N | W | 1.95 |
| 232 | 143.0344 | 8.4405 | -2.90 | 7.41 | A | A | A | 0.59 |
| 233 | 126.61 | 1.02 | -14.05 | 4.55 | A | N | W | 2.89 |
| 234 | 138.22 | 4.7 | -6.16 | 5.62 | A | A | A | 1.27 |
| 235 | 140.32 | 2.75 | -4.74 | 4.89 | A | A | A | 0.97 |
| 237 | 151.5 | 9.1 | 2.85 | 7.49 | A | A | A | 0.59 |
| 238 | 144.43 | 18.776 | -1.95 | 13.75 | A | A | A | 0.40 |
| 239 | 144 | 5 | -2.24 | 5.67 | A | A | A | 0.46 |
| 240 | 153.5 | 4.4 | 4.21 | 5.32 | A | A | A | 0.86 |
| 241 | 150 | 8 | 1.83 | 6.97 | A | A | A | 0.38 |
| 242 | 145.8 | 7.7 | -1.02 | 6.93 | A | A | A | 0.21 |
| 245 | 142.2 | 5.8 | -3.46 | 6.06 | A | A | A | 0.71 |
| 246 | 129.34 | 2.32 | -12.19 | 4.83 | A | A | A | 2.50 |
| 248 | 148.78 | 4.67 | 1.00 | 5.47 | A | A | A | 0.21 |
| 249 | 137 | 4 | -6.99 | 5.35 | A | A | A | 1.44 |
| 250 | 132 | 8 | -10.39 | 7.54 | A | A | A | 2.13 |
| 251 | 145 | 2.9 | -1.56 | 4.91 | A | A | A | 0.32 |
| 252 | 139.0287 | 3.5 | -5.62 | 5.14 | A | A | A | 1.15 |
| 254 | 139.0 | 8.5 | -5.63 | 7.58 | A | A | A | 1.16 |
| 255 | 148.6 | 4.6 | 0.88 | 5.45 | A | A | A | 0.18 |
| 256 | 135.75 | 3.34 | -7.84 | 5.11 | A | A | A | 1.61 |
| 257 | 141 | 5 | -4.28 | 5.71 | A | A | A | 0.88 |
| 258 | 131 | 5.62 | -11.07 | 6.20 | A | A | A | 2.27 |
| 259 | 139 | 0.36 | -5.63 | 4.49 | A | A | A | 1.16 |
| 260 | 144.3 | 14.0 | -2.04 | 10.69 | A | A | A | 0.42 |
| 261 | 141 | 7.1 | -4.28 | 6.74 | A | A | A | 0.88 |
| 262 | 137 | 8.0 | -6.99 | 7.36 | A | A | A | 1.44 |

TABLE 12. Evaluation Results for Ba-133 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 263 | 149.02 | 4.34 | 1.17 | 5.34 | A | A | A | 0.24 |
| 266 | 138.7 | 6.0 | -5.84 | 6.23 | A | A | A | 1.20 |
| 267 | 253.93 | 8.66 | 72.39 | 5.63 | N | N | N | 14.87 |
| 268 | 141.5 | 5.7 | -3.94 | 6.03 | A | A | A | 0.81 |
| 270 | 147 | 15.8 | -0.20 | 11.64 | A | A | A | 0.04 |
| 274 | 138.2 | 8.0 | -6.18 | 7.32 | A | A | A | 1.27 |
| 275 | 147 | 9.21 | -0.20 | 7.70 | A | A | A | 0.04 |
| 279 | 152.2 | 7.5 | 3.33 | 6.66 | A | A | A | 0.68 |
| 281 | 138 | 9 | -6.31 | 7.91 | A | A | A | 1.30 |
| 283 | 144 | 2.7 | -2.24 | 4.86 | A | A | A | 0.46 |
| 286 | 141.0 | 10.7 | -4.28 | 8.81 | A | A | A | 0.88 |
| 287 | 138.4 | 1.58 | -6.04 | 4.62 | A | A | A | 1.24 |
| 288 | 147 | 11 | -0.20 | 8.72 | A | A | A | 0.04 |
| 289 | 148 | 6 | 0.48 | 6.04 | A | A | A | 0.10 |
| 290 | 107.07 | 8.4 | -27.31 | 9.03 | N | N | N | 5.61 |
| 291 | 140 | 5 | -4.96 | 5.73 | A | A | A | 1.02 |
| 295 | 139.0 | 12.4 | -5.63 | 9.98 | A | A | A | 1.16 |
| 297 | 120.0 | 11.0 | -18.53 | 10.20 | A | A | A | 3.81 |
| 298 | 146.2 | 2.9 | -0.75 | 4.90 | A | A | A | 0.15 |
| 299 | 144.7 | 8.9 | -1.77 | 7.61 | A | A | A | 0.36 |
| 304 | 142.84 | 10.00 | -3.03 | 8.31 | A | A | A | 0.62 |
| 305 | 136 | 6.31 | -7.67 | 6.45 | A | A | A | 1.58 |
| 307 | 129.4 | 0.1 | -12.15 | 4.48 | A | N | W | 2.50 |
| 310 | 144.2 | 4.1 | -2.10 | 5.31 | A | A | A | 0.43 |
| 312 | 150.58 | 2.34 | 2.23 | 4.74 | A | A | A | 0.46 |
| 313 | 141.7 | 3.4 | -3.80 | 5.08 | A | A | A | 0.78 |
| 314 | 154.61 | 7.14 | 4.96 | 6.43 | A | A | A | 1.02 |
| 317 | 133 |  | -9.71 |  | A |  |  | 1.99 |
| 320 | 146.40 | 4.56 | -0.61 | 5.46 | A | A | A | 0.13 |
| 321 | 141.64 | 0.43 | -3.84 | 4.49 | A | A | A | 0.79 |
| 322 | 140.0 | 12.6 | -4.96 | 10.05 | A | A | A | 1.02 |
| 323 | 144.9 | 1.3 | -1.63 | 4.57 | A | A | A | 0.33 |
| 324 | 137.2 | 0.89 | -6.86 | 4.53 | A | A | A | 1.41 |
| 325 | 151.9 | 12.2 | 3.12 | 9.20 | A | A | A | 0.64 |

## Reported Results for Cs-137 Determination in Sample 2

TABLE 13. Target values

| Parameter | Value |
| --- | --- |
| Target Value, Bq/kg | 147.4 |
| UNC, Bq/kg | 6.6 |
| MARB, % | 20 |

FIGURE 13. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

FIGURE 14. Reported results and evaluation parameters for Cs-137. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

TABLE 14. Evaluation Results for Cs-137 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 149.23 | 4.29 | 1.24 | 5.32 | A | A | A | 0.31 |
| 2 | 144.4 | 4.3 | -2.04 | 5.38 | A | A | A | 0.51 |
| 3 | 141.8 | 7.6 | -3.80 | 6.98 | A | A | A | 0.95 |
| 4 | 140.6 | 7.0 | -4.61 | 6.70 | A | A | A | 1.15 |
| 5 | 142.4 | 8.8 | -3.39 | 7.63 | A | A | A | 0.85 |
| 6 | 131.7 | 9 | -10.65 | 8.17 | A | A | A | 2.66 |
| 7 | 148.27 | 4.89 | 0.59 | 5.56 | A | A | A | 0.15 |
| 8 | 134.32 | 1.47 | -8.87 | 4.61 | A | A | A | 2.22 |
| 9 | 150 | 17 | 1.76 | 12.19 | A | A | A | 0.44 |
| 10 | 138.8 | 1.68 | -5.83 | 4.64 | A | A | A | 1.46 |
| 11 | 147.4 | 4.8 | 0.00 | 5.54 | A | A | A | 0.00 |
| 12 | 169.9 | 1.9 | 15.26 | 4.62 | A | N | W | 3.81 |
| 13 | 144.02 | 5.38 | -2.29 | 5.83 | A | A | A | 0.57 |
| 14 | 146.2 | 8.0 | -0.81 | 7.07 | A | A | A | 0.20 |
| 15 | 155.0 | 7.7 | 5.16 | 6.69 | A | A | A | 1.29 |
| 16 | 145.2 | 3.2 | -1.49 | 4.99 | A | A | A | 0.37 |
| 17 | 145.40 | 4.31 | -1.36 | 5.37 | A | A | A | 0.34 |
| 18 | 143.6 | 8.5 | -2.58 | 7.42 | A | A | A | 0.64 |
| 19 | 137.5 | 6.0 | -6.72 | 6.25 | A | A | A | 1.68 |
| 20 | 144 | 3 | -2.31 | 4.94 | A | A | A | 0.58 |
| 21 | 146 | 10 | -0.95 | 8.18 | A | A | A | 0.24 |
| 22 | 155.7 | 2.9 | 5.63 | 4.85 | A | A | A | 1.41 |
| 25 | 145 | 4 | -1.63 | 5.26 | A | A | A | 0.41 |
| 26 | 158.64 | 8.54 | 7.63 | 7.00 | A | A | A | 1.91 |
| 27 | 149.4 | 3.9 | 1.36 | 5.18 | A | A | A | 0.34 |
| 29 | 157.04 | 3.66 | 6.54 | 5.05 | A | A | A | 1.63 |
| 31 | 147.00 | 6.65 | -0.27 | 6.37 | A | A | A | 0.07 |
| 32 | 150.1 | 16.5 | 1.83 | 11.87 | A | A | A | 0.46 |
| 33 | 140 | 6 | -5.02 | 6.20 | A | A | A | 1.25 |
| 34 | 141.8 | 6.1 | -3.80 | 6.21 | A | A | A | 0.95 |

TABLE 14. Evaluation Results for Cs-137 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 35 | 139 | 7 | -5.70 | 6.74 | A | A | A | 1.42 |
| 36 | 144.3 | 6.4 | -2.10 | 6.30 | A | A | A | 0.53 |
| 38 | 140 | 8 | -5.02 | 7.26 | A | A | A | 1.25 |
| 39 | 142.7 | 3.4 | -3.19 | 5.07 | A | A | A | 0.80 |
| 41 | 139.4 | 5.0 | -5.43 | 5.74 | A | A | A | 1.36 |
| 42 | 154.63 | 2.42 | 4.91 | 4.74 | A | A | A | 1.23 |
| 43 | 148.74 | 2.98 | 0.91 | 4.91 | A | A | A | 0.23 |
| 45 | 72.0 | 3.7 | -51.15 | 6.82 | N | N | N | 12.78 |
| 46 | 288.77 | 6.30 | 95.91 | 4.98 | N | N | N | 23.96 |
| 47 | 139 | 2.2 | -5.70 | 4.75 | A | A | A | 1.42 |
| 48 | 147.2 | 5.3 | -0.14 | 5.75 | A | A | A | 0.03 |
| 50 | 149.76 | 4.10 | 1.60 | 5.25 | A | A | A | 0.40 |
| 51 | 126 | 2 | -14.52 | 4.75 | A | N | W | 3.63 |
| 52 | 148.9 | 5.7 | 1.02 | 5.89 | A | A | A | 0.25 |
| 53 | 145.7 | 7.7 | -1.15 | 6.93 | A | A | A | 0.29 |
| 55 | 144.3 | 3.3 | -2.10 | 5.03 | A | A | A | 0.53 |
| 56 | 147.1 | 4.35 | -0.20 | 5.37 | A | A | A | 0.05 |
| 59 | 146 | 10.3 | -0.95 | 8.36 | A | A | A | 0.24 |
| 60 | 132.8 | 18.8 | -9.91 | 14.85 | A | A | A | 2.47 |
| 61 | 144.73 | 4.30 | -1.81 | 5.37 | A | A | A | 0.45 |
| 63 | 141.5 | 4.0 | -4.00 | 5.30 | A | A | A | 1.00 |
| 64 | 149.8 | 8 | 1.63 | 6.97 | A | A | A | 0.41 |
| 66 | 157 | 9.4 | 6.51 | 7.48 | A | A | A | 1.63 |
| 68 | 144.8 | 7.9 | -1.76 | 7.06 | A | A | A | 0.44 |
| 69 | 148.59 | 2.25 | 0.81 | 4.73 | A | A | A | 0.20 |
| 70 | 147.53 | 1.73 | 0.09 | 4.63 | A | A | A | 0.02 |
| 72 | 147.792 | 9.961 | 0.27 | 8.09 | A | A | A | 0.07 |
| 74 | 158.2 | 6.5 | 7.33 | 6.08 | A | A | A | 1.83 |
| 75 | 147.4 | 4.7 | 0.00 | 5.50 | A | A | A | 0.00 |
| 76 | 143.00 | 0.52 | -2.99 | 4.49 | A | A | A | 0.75 |
| 77 | 175 | 1.6 | 18.72 | 4.57 | A | N | W | 4.68 |
| 80 | 148.6 | 9.5 | 0.81 | 7.81 | A | A | A | 0.20 |
| 81 | 147.0 | 10.4 | -0.27 | 8.37 | A | A | A | 0.07 |
| 84 | 152.76 | 3.37 | 3.64 | 4.99 | A | A | A | 0.91 |
| 85 | 150.83 | 12.60 | 2.33 | 9.48 | A | A | A | 0.58 |
| 86 | 143 | 4.1 | -2.99 | 5.32 | A | A | A | 0.75 |
| 87 | 157 | 6 | 6.51 | 5.89 | A | A | A | 1.63 |
| 88 | 141.3 | 8.0 | -4.14 | 7.22 | A | A | A | 1.03 |
| 89 | 149 | 5 | 1.09 | 5.60 | A | A | A | 0.27 |
| 92 | 149 | 16 | 1.09 | 11.63 | A | A | A | 0.27 |
| 93 | 145.70 | 2.2 | -1.15 | 4.73 | A | A | A | 0.29 |
| 94 | 139.22 | 0.73 | -5.55 | 4.51 | A | A | A | 1.39 |
| 95 | 139.89 | 0.74 | -5.09 | 4.51 | A | A | A | 1.27 |
| 96 | 148.6 | 8.9 | 0.81 | 7.48 | A | A | A | 0.20 |
| 98 | 146.2 | 7.0 | -0.81 | 6.56 | A | A | A | 0.20 |
| 99 | 147 | 4.6 | -0.27 | 5.46 | A | A | A | 0.07 |
| 101 | 150.2 | 4.0 | 1.90 | 5.21 | A | A | A | 0.47 |
| 102 | 148.0 | 4.5 | 0.41 | 5.41 | A | A | A | 0.10 |
| 103 | 144.12 | 2.99 | -2.23 | 4.93 | A | A | A | 0.56 |
| 105 | 160.4 | 1.9 | 8.82 | 4.63 | A | A | A | 2.20 |
| 106 | 145.6 | 0.5 | -1.22 | 4.49 | A | A | A | 0.31 |
| 107 | 171 | 2 | 16.01 | 4.63 | A | N | W | 4.00 |
| 108 | 146.2275 | 0.8734 | -0.80 | 4.52 | A | A | A | 0.20 |
| 109 | 157.6 | 6.0 | 6.92 | 5.88 | A | A | A | 1.73 |

TABLE 14. Evaluation Results for Cs-137 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 110 | 145.321 | 3.135 | -1.41 | 4.97 | A | A | A | 0.35 |
| 111 | 145.6 | 9.7 | -1.22 | 8.03 | A | A | A | 0.31 |
| 112 | 111.2964 | 3.429 | -24.49 | 5.44 | N | N | N | 6.12 |
| 113 | 150.0 | 7.5 | 1.76 | 6.71 | A | A | A | 0.44 |
| 116 | 142.8 | 1.8 | -3.12 | 4.65 | A | A | A | 0.78 |
| 118 | 142.0 | 12.0 | -3.66 | 9.56 | A | A | A | 0.92 |
| 120 | 142 | 4 | -3.66 | 5.29 | A | A | A | 0.92 |
| 121 | 140 | 11 | -5.02 | 9.04 | A | A | A | 1.25 |
| 123 | 140 | 9 | -5.02 | 7.83 | A | A | A | 1.25 |
| 124 | 148.4 | 2.8 | 0.68 | 4.86 | A | A | A | 0.17 |
| 125 | 143 | 3.3 | -2.99 | 5.04 | A | A | A | 0.75 |
| 126 | 138 | 8.3 | -6.38 | 7.50 | A | A | A | 1.59 |
| 129 | 147 | 4 | -0.27 | 5.24 | A | A | A | 0.07 |
| 130 | 148 | 5 | 0.41 | 5.61 | A | A | A | 0.10 |
| 132 | 143 | 8.6 | -2.99 | 7.50 | A | A | A | 0.75 |
| 133 | 143.0 | 7.2 | -2.99 | 6.74 | A | A | A | 0.75 |
| 134 | 107.32 | 6.49 | -27.19 | 7.52 | N | N | N | 6.79 |
| 135 | 142.5 | 4.6 | -3.32 | 5.52 | A | A | A | 0.83 |
| 137 | 135.9 | 3.39 | -7.80 | 5.13 | A | A | A | 1.95 |
| 139 | 149.56 | 1.1499 | 1.47 | 4.54 | A | A | A | 0.37 |
| 140 | 144.3 | 4.4 | -2.10 | 5.42 | A | A | A | 0.53 |
| 141 | 152 | 8.6 | 3.12 | 7.22 | A | A | A | 0.78 |
| 143 | 140 | 8.2 | -5.02 | 7.37 | A | A | A | 1.25 |
| 144 | 145.57 | 7.8 | -1.24 | 6.98 | A | A | A | 0.31 |
| 146 | 143.9 | 8.9 | -2.37 | 7.64 | A | A | A | 0.59 |
| 147 | 142 | 13 | -3.66 | 10.19 | A | A | A | 0.92 |
| 148 | 152.87 | 4.67 | 3.71 | 5.42 | A | A | A | 0.93 |
| 149 | 146.4 | 3.3 | -0.68 | 5.01 | A | A | A | 0.17 |
| 150 | 148.3 | 4.5 | 0.61 | 5.41 | A | A | A | 0.15 |
| 151 | 145.0 | 4.0 | -1.63 | 5.26 | A | A | A | 0.41 |
| 152 | 147 | 5 | -0.27 | 5.62 | A | A | A | 0.07 |
| 153 | 145.2 | 10.2 | -1.49 | 8.33 | A | A | A | 0.37 |
| 154 | 149.8 | 2.7 | 1.63 | 4.83 | A | A | A | 0.41 |
| 155 | 148.3 | 6.9 | 0.61 | 6.46 | A | A | A | 0.15 |
| 156 | 143 | 9 | -2.99 | 7.72 | A | A | A | 0.75 |
| 157 | 140.04 | 7.46 | -4.99 | 6.96 | A | A | A | 1.25 |
| 158 | 175.7 | 14.7 | 19.20 | 9.49 | A | A | A | 4.80 |
| 159 | 133.0 | 1.7 | -9.77 | 4.66 | A | A | A | 2.44 |
| 160 | 147 | 3 | -0.27 | 4.92 | A | A | A | 0.07 |
| 161 | 143.6 | 4.1 | -2.58 | 5.31 | A | A | A | 0.64 |
| 162 | 148.4 | 2.8 | 0.68 | 4.86 | A | A | A | 0.17 |
| 163 | 148.0 | 8.8 | 0.41 | 7.44 | A | A | A | 0.10 |
| 164 | 150 | 8 | 1.76 | 6.96 | A | A | A | 0.44 |
| 165 | 162.3 | 9.1 | 10.11 | 7.18 | A | A | A | 2.53 |
| 166 | 145.2 | 7.9 | -1.49 | 7.05 | A | A | A | 0.37 |
| 168 | 139.57 | 8.60 | -5.31 | 7.62 | A | A | A | 1.33 |
| 169 | 145.87 | 1.81 | -1.04 | 4.65 | A | A | A | 0.26 |
| 171 | 148.00 | 13.32 | 0.41 | 10.05 | A | A | A | 0.10 |
| 172 | 163.9 | 21.6 | 11.19 | 13.92 | A | A | A | 2.80 |
| 173 | 156.0 | 5.0 | 5.83 | 5.51 | A | A | A | 1.46 |
| 176 | 147 | 8.31 | -0.27 | 7.21 | A | A | A | 0.07 |
| 177 | 162 | 10 | 9.91 | 7.63 | A | A | A | 2.47 |
| 179 | 148.4 | 3.4 | 0.68 | 5.03 | A | A | A | 0.17 |
| 180 | 146.79 | 5.08 | -0.41 | 5.66 | A | A | A | 0.10 |

TABLE 14. Evaluation Results for Cs-137 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 181 | 141.394 | 3.073 | -4.07 | 4.98 | A | A | A | 1.02 |
| 182 | 159.13 | 7.09 | 7.96 | 6.32 | A | A | A | 1.99 |
| 183 | 148.8 | 2.82 | 0.95 | 4.86 | A | A | A | 0.24 |
| 184 | 149.31 | 0.43 | 1.30 | 4.49 | A | A | A | 0.32 |
| 185 | 146.2 | 8.5 | -0.81 | 7.34 | A | A | A | 0.20 |
| 186 | 143.4 | 8.7 | -2.71 | 7.54 | A | A | A | 0.68 |
| 187 | 180.2 | 4.4 | 22.25 | 5.10 | N | N | N | 5.56 |
| 189 | 147.4 | 9.1 | 0.00 | 7.63 | A | A | A | 0.00 |
| 190 | 150 | 15 | 1.76 | 10.96 | A | A | A | 0.44 |
| 191 | 146.8 | 5.5 | -0.41 | 5.84 | A | A | A | 0.10 |
| 192 | 0.086 | 0.05 | -99.94 | 58.31 | N | N | N | 24.97 |
| 193 | 143.7 | 14.37 | -2.51 | 10.96 | A | A | A | 0.63 |
| 195 | 149.7 | 6.00 | 1.56 | 6.01 | A | A | A | 0.39 |
| 196 | 137 | 11 | -7.06 | 9.19 | A | A | A | 1.76 |
| 197 | 145 | 4 | -1.63 | 5.26 | A | A | A | 0.41 |
| 198 | 143.85 | 6.37 | -2.41 | 6.30 | A | A | A | 0.60 |
| 199 | 142.71 | 2.85 | -3.18 | 4.90 | A | A | A | 0.79 |
| 200 | 147.60 | 3.56 | 0.14 | 5.09 | A | A | A | 0.03 |
| 201 | 150 | 15 | 1.76 | 10.96 | A | A | A | 0.44 |
| 202 | 141.67 | 8.20 | -3.89 | 7.32 | A | A | A | 0.97 |
| 205 | 146 | 5 | -0.95 | 5.64 | A | A | A | 0.24 |
| 206 | 144.9 | 4.2 | -1.70 | 5.33 | A | A | A | 0.42 |
| 207 | 144 | 8.1 | -2.31 | 7.19 | A | A | A | 0.58 |
| 208 | 133.07 | 7.27 | -9.72 | 7.06 | A | A | A | 2.43 |
| 209 | 156 | 8.6 | 5.83 | 7.10 | A | A | A | 1.46 |
| 210 | 147 | 10 | -0.27 | 8.14 | A | A | A | 0.07 |
| 211 | 148 | 11 | 0.41 | 8.68 | A | A | A | 0.10 |
| 212 | 146 | 3 | -0.95 | 4.93 | A | A | A | 0.24 |
| 213 | 145 | 9 | -1.63 | 7.65 | A | A | A | 0.41 |
| 214 | 150 | 10.5 | 1.76 | 8.31 | A | A | A | 0.44 |
| 215 | 134 | 8.2 | -9.09 | 7.58 | A | A | A | 2.27 |
| 216 | 140.7 | 8.50 | -4.55 | 7.52 | A | A | A | 1.14 |
| 219 | 143 | 7.7 | -2.99 | 7.00 | A | A | A | 0.75 |
| 221 | 160.02 | 4.53 | 8.56 | 5.30 | A | A | A | 2.14 |
| 222 | 156 | 4.7 | 5.83 | 5.40 | A | A | A | 1.46 |
| 223 | 137.1 | 4 | -6.99 | 5.34 | A | A | A | 1.75 |
| 224 | 156.7 | 4.2 | 6.31 | 5.22 | A | A | A | 1.58 |
| 225 | 141 | 6 | -4.34 | 6.18 | A | A | A | 1.08 |
| 227 | 142.61 | 1.98 | -3.25 | 4.69 | A | A | A | 0.81 |
| 228 | 142.1 | 5.0 | -3.60 | 5.69 | A | A | A | 0.90 |
| 229 | 144 | 4.40 | -2.31 | 5.42 | A | A | A | 0.58 |
| 230 | 144 | 5 | -2.31 | 5.67 | A | A | A | 0.58 |
| 231 | 150.4 | 26.84 | 2.04 | 18.40 | A | A | A | 0.51 |
| 232 | 142.9839 | 6.3825 | -3.00 | 6.32 | A | A | A | 0.75 |
| 233 | 138.11 | 1.47 | -6.30 | 4.60 | A | A | A | 1.57 |
| 234 | 148.66 | 4.7 | 0.85 | 5.48 | A | A | A | 0.21 |
| 235 | 149.16 | 4.04 | 1.19 | 5.23 | A | A | A | 0.30 |
| 237 | 141.9 | 8.6 | -3.73 | 7.54 | A | A | A | 0.93 |
| 238 | 147.865 | 19.222 | 0.32 | 13.75 | A | A | A | 0.08 |
| 239 | 140 | 5 | -5.02 | 5.73 | A | A | A | 1.25 |
| 240 | 169.6 | 4.8 | 15.06 | 5.30 | A | N | W | 3.76 |
| 241 | 153 | 9 | 3.80 | 7.39 | A | A | A | 0.95 |
| 242 | 141.6 | 7.2 | -3.93 | 6.78 | A | A | A | 0.98 |
| 243 | 137.9 | 7.2 | -6.45 | 6.88 | A | A | A | 1.61 |

TABLE 14. Evaluation Results for Cs-137 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 245 | 141.0 | 2.17 | -4.34 | 4.73 | A | A | A | 1.08 |
| 246 | 131.05 | 3.65 | -11.09 | 5.27 | A | A | A | 2.77 |
| 248 | 143.80 | 4.31 | -2.44 | 5.39 | A | A | A | 0.61 |
| 249 | 147 | 13 | -0.27 | 9.91 | A | A | A | 0.07 |
| 250 | 143 | 5 | -2.99 | 5.68 | A | A | A | 0.75 |
| 251 | 150 | 2.0 | 1.76 | 4.67 | A | A | A | 0.44 |
| 252 | 150.2242 | 4.12 | 1.92 | 5.25 | A | A | A | 0.48 |
| 254 | 143.2 | 6.3 | -2.85 | 6.28 | A | A | A | 0.71 |
| 255 | 151.7 | 4.3 | 2.92 | 5.30 | A | A | A | 0.73 |
| 256 | 135.79 | 0.86 | -7.88 | 4.52 | A | A | A | 1.97 |
| 257 | 159 | 7 | 7.87 | 6.28 | A | A | A | 1.97 |
| 258 | 141 | 6.05 | -4.34 | 6.20 | A | A | A | 1.08 |
| 259 | 141 | 0.32 | -4.34 | 4.48 | A | A | A | 1.08 |
| 260 | 152.1 | 7.9 | 3.19 | 6.86 | A | A | A | 0.80 |
| 261 | 151 | 7.6 | 2.44 | 6.74 | A | A | A | 0.61 |
| 262 | 140 | 8.0 | -5.02 | 7.26 | A | A | A | 1.25 |
| 263 | 147.34 | 4.29 | -0.04 | 5.34 | A | A | A | 0.01 |
| 264 | 162.7 | 2.0 | 10.38 | 4.64 | A | A | A | 2.59 |
| 266 | 141.8 | 8.5 | -3.80 | 7.48 | A | A | A | 0.95 |
| 267 | 264.79 | 17.03 | 79.64 | 7.84 | N | N | N | 19.90 |
| 268 | 133.6 | 2.7 | -9.36 | 4.91 | A | A | A | 2.34 |
| 270 | 147 | 15.9 | -0.27 | 11.71 | A | A | A | 0.07 |
| 273 | 152.1 | 4.25 | 3.19 | 5.28 | A | A | A | 0.80 |
| 274 | 140.0 | 7.8 | -5.02 | 7.15 | A | A | A | 1.25 |
| 275 | 146 | 9.02 | -0.95 | 7.63 | A | A | A | 0.24 |
| 279 | 145.5 | 8.7 | -1.29 | 7.47 | A | A | A | 0.32 |
| 281 | 143 | 6 | -2.99 | 6.14 | A | A | A | 0.75 |
| 283 | 144 | 6.5 | -2.31 | 6.36 | A | A | A | 0.58 |
| 286 | 146.0 | 5.7 | -0.95 | 5.94 | A | A | A | 0.24 |
| 287 | 145 | 1.4 | -1.63 | 4.58 | A | A | A | 0.41 |
| 288 | 139 | 8 | -5.70 | 7.29 | A | A | A | 1.42 |
| 289 | 142.0 | 7.5 | -3.66 | 6.92 | A | A | A | 0.92 |
| 290 | 160.26 | 8.27 | 8.72 | 6.83 | A | A | A | 2.18 |
| 291 | 141 | 5 | -4.34 | 5.71 | A | A | A | 1.08 |
| 295 | 146.0 | 13.4 | -0.95 | 10.21 | A | A | A | 0.24 |
| 297 | 133.0 | 9.0 | -9.77 | 8.11 | A | A | A | 2.44 |
| 298 | 137.7 | 2.9 | -6.58 | 4.95 | A | A | A | 1.64 |
| 299 | 172.6 | 10.4 | 17.10 | 7.51 | A | A | A | 4.27 |
| 304 | 103.93 | 7.27 | -29.49 | 8.31 | N | N | N | 7.37 |
| 305 | 113 | 7.15 | -23.34 | 7.75 | N | N | N | 5.83 |
| 307 | 14.8 | 0.7 | -89.96 | 6.51 | N | N | N | 22.47 |
| 308 | 134.26 | 8.10 | -8.91 | 7.51 | A | A | A | 2.23 |
| 310 | 148.9 | 2.9 | 1.02 | 4.88 | A | A | A | 0.25 |
| 312 | 139.88 | 2.13 | -5.10 | 4.73 | A | A | A | 1.27 |
| 313 | 142.0 | 3.6 | -3.66 | 5.15 | A | A | A | 0.92 |
| 314 | 143.33 | 6.62 | -2.76 | 6.43 | A | A | A | 0.69 |
| 317 | 150 |  | 1.76 |  | A |  |  | 0.44 |
| 320 | 135.75 | 4.95 | -7.90 | 5.77 | A | A | A | 1.97 |
| 321 | 143.89 | 1.1 | -2.38 | 4.54 | A | A | A | 0.59 |
| 322 | 159.0 | 16.6 | 7.87 | 11.36 | A | A | A | 1.97 |
| 323 | 144.0 | 2.7 | -2.31 | 4.85 | A | A | A | 0.58 |
| 324 | 138.7 | 2.03 | -5.90 | 4.71 | A | A | A | 1.47 |
| 325 | 143.8 | 8.5 | -2.44 | 7.42 | A | A | A | 0.61 |

## Reported Results for Am-241 Determination in Sample 2

TABLE 15. Target values

| Parameter | Value |
| --- | --- |
| Target Value, Bq/kg | 108.3 |
| UNC, Bq/kg | 4.9 |
| MARB, % | 20 |

FIGURE 15. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

FIGURE 16. Reported results and evaluation parameters for Am-241. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

TABLE 16. Evaluation Results for Am-241 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 115.09 | 3.52 | 6.27 | 5.46 | A | A | A | 0.84 |
| 2 | 113.5 | 5.8 | 4.80 | 6.83 | A | A | A | 0.65 |
| 3 | 108.8 | 6.7 | 0.46 | 7.64 | A | A | A | 0.06 |
| 4 | 41.1 | 2.1 | -62.05 | 6.82 | N | N | N | 8.35 |
| 5 | 111.7 | 12 | 3.14 | 11.66 | A | A | A | 0.42 |
| 6 | 156.1 | 14.5 | 44.14 | 10.33 | N | N | N | 5.94 |
| 7 | 85.14 | 7.13 | -21.39 | 9.52 | N | A | N | 2.88 |
| 9 | 115 | 19 | 6.19 | 17.13 | A | A | A | 0.83 |
| 10 | 104.1 | 1.6 | -3.88 | 4.78 | A | A | A | 0.52 |
| 11 | 105.3 | 3.7 | -2.77 | 5.73 | A | A | A | 0.37 |
| 12 | 151.4 | 3.8 | 39.80 | 5.17 | N | N | N | 5.35 |
| 13 | 107.54 | 3.93 | -0.70 | 5.82 | A | A | A | 0.09 |
| 14 | 113.9 | 8.2 | 5.17 | 8.50 | A | A | A | 0.70 |
| 15 | 118.2 | 10.5 | 9.14 | 9.97 | A | A | A | 1.23 |
| 16 | 106.6 | 4.3 | -1.57 | 6.06 | A | A | A | 0.21 |
| 17 | 114.05 | 4.16 | 5.31 | 5.81 | A | A | A | 0.71 |
| 18 | 110.4 | 8.0 | 1.94 | 8.54 | A | A | A | 0.26 |
| 19 | 110.7 | 5.5 | 2.22 | 6.72 | A | A | A | 0.30 |
| 20 | 110 | 6 | 1.57 | 7.09 | A | A | A | 0.21 |
| 21 | 105.8 | 7.8 | -2.31 | 8.65 | A | A | A | 0.31 |
| 22 | 110.2 | 3.5 | 1.75 | 5.53 | A | A | A | 0.24 |
| 25 | 110 | 10 | 1.57 | 10.15 | A | A | A | 0.21 |
| 26 | 97.83 | 5.36 | -9.67 | 7.11 | A | A | A | 1.30 |
| 27 | 121 | 5 | 11.73 | 6.13 | A | A | A | 1.58 |
| 29 | 133.10 | 4.41 | 22.90 | 5.61 | N | N | N | 3.08 |
| 31 | 123.20 | 7.57 | 13.76 | 7.63 | A | A | A | 1.85 |
| 32 | 115.4 | 18.5 | 6.56 | 16.66 | A | A | A | 0.88 |
| 33 | 115 | 5 | 6.19 | 6.27 | A | A | A | 0.83 |
| 34 | 105.2 | 6.2 | -2.86 | 7.43 | A | A | A | 0.39 |
| 35 | 95.7 | 5.8 | -11.63 | 7.56 | A | A | A | 1.57 |

TABLE 16. Evaluation Results for Am-241 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 36 | 105 | 5.1 | -3.05 | 6.64 | A | A | A | 0.41 |
| 38 | 106 | 8 | -2.12 | 8.80 | A | A | A | 0.29 |
| 39 | 108 | 4 | -0.28 | 5.85 | A | A | A | 0.04 |
| 41 | 111.6 | 6.0 | 3.05 | 7.03 | A | A | A | 0.41 |
| 43 | 110.24 | 2.75 | 1.79 | 5.17 | A | A | A | 0.24 |
| 45 | 60.2 | 3.1 | -44.41 | 6.85 | N | N | N | 5.98 |
| 46 | 173.86 | 3.90 | 60.54 | 5.05 | N | N | N | 8.14 |
| 47 | 190 | 9.4 | 75.44 | 6.70 | N | N | N | 10.15 |
| 48 | 103.5 | 3.8 | -4.43 | 5.83 | A | A | A | 0.60 |
| 50 | 112.22 | 18.46 | 3.62 | 17.06 | A | A | A | 0.49 |
| 51 | 39 | 1 | -63.99 | 5.20 | N | N | N | 8.61 |
| 52 | 110.1 | 4.9 | 1.66 | 6.35 | A | A | A | 0.22 |
| 53 | 98.6 | 4.0 | -8.96 | 6.08 | A | A | A | 1.20 |
| 56 | 107.2 | 2.92 | -1.02 | 5.28 | A | A | A | 0.14 |
| 59 | 101 | 7.1 | -6.74 | 8.36 | A | A | A | 0.91 |
| 63 | 135.8 | 5.7 | 25.39 | 6.17 | N | N | N | 3.42 |
| 64 | 110.2 | 6.5 | 1.75 | 7.43 | A | A | A | 0.24 |
| 66 | 116 | 11.8 | 7.11 | 11.13 | A | A | A | 0.96 |
| 68 | 118.0 | 8.0 | 8.96 | 8.15 | A | A | A | 1.20 |
| 69 | 111.46 | 3.72 | 2.92 | 5.62 | A | A | A | 0.39 |
| 70 | 104.41 | 4.08 | -3.59 | 5.98 | A | A | A | 0.48 |
| 72 | 124.146 | 8.825 | 14.63 | 8.43 | A | A | A | 1.97 |
| 74 | 111.4 | 5.0 | 2.86 | 6.37 | A | A | A | 0.39 |
| 75 | 140 | 9.7 | 29.27 | 8.28 | N | N | N | 3.94 |
| 76 | 96.30 | 2.91 | -11.08 | 5.44 | A | A | A | 1.49 |
| 80 | 154.1 | 10.2 | 42.29 | 8.02 | N | N | N | 5.69 |
| 81 | 108.0 | 7.9 | -0.28 | 8.60 | A | A | A | 0.04 |
| 84 | 104.19 | 2.5 | -3.80 | 5.12 | A | A | A | 0.51 |
| 85 | 112.43 | 10.21 | 3.81 | 10.15 | A | A | A | 0.51 |
| 86 | 101 | 6.1 | -6.74 | 7.55 | A | A | A | 0.91 |
| 87 | 110 | 6 | 1.57 | 7.09 | A | A | A | 0.21 |
| 88 | 102.6 | 6.5 | -5.26 | 7.79 | A | A | A | 0.71 |
| 91 | 108 | 13 | -0.28 | 12.86 | A | A | A | 0.04 |
| 92 | 113 | 13 | 4.34 | 12.36 | A | A | A | 0.58 |
| 93 | 101.1 | 6.05 | -6.65 | 7.50 | A | A | A | 0.89 |
| 94 | 100.08 | 1.06 | -7.59 | 4.65 | A | A | A | 1.02 |
| 95 | 99.02 | 1.13 | -8.57 | 4.67 | A | A | A | 1.15 |
| 96 | 108.0 | 12.0 | -0.28 | 12.00 | A | A | A | 0.04 |
| 98 | 109.8 | 7.2 | 1.39 | 7.97 | A | A | A | 0.19 |
| 99 | 114 | 10.2 | 5.26 | 10.03 | A | A | A | 0.71 |
| 101 | 109.6 | 3.5 | 1.20 | 5.54 | A | A | A | 0.16 |
| 102 | 104.2 | 3.4 | -3.79 | 5.58 | A | A | A | 0.51 |
| 103 | 110.78 | 3.85 | 2.29 | 5.71 | A | A | A | 0.31 |
| 105 | 102.7 | 2.1 | -5.17 | 4.97 | A | A | A | 0.70 |
| 106 | 106.0 | 1.4 | -2.12 | 4.71 | A | A | A | 0.29 |
| 107 | 112 | 3 | 3.42 | 5.26 | A | A | A | 0.46 |
| 109 | 107.8 | 5.0 | -0.46 | 6.48 | A | A | A | 0.06 |
| 111 | 133.9 | 8.4 | 23.64 | 7.73 | N | N | N | 3.18 |
| 113 | 107.3 | 5.6 | -0.92 | 6.91 | A | A | A | 0.12 |
| 116 | 107.4 | 2.8 | -0.83 | 5.22 | A | A | A | 0.11 |
| 118 | 109.6 | 10.3 | 1.20 | 10.43 | A | A | A | 0.16 |
| 120 | 105 | 4 | -3.05 | 5.91 | A | A | A | 0.41 |
| 121 | 110 | 12 | 1.57 | 11.81 | A | A | A | 0.21 |
| 123 | 113 | 7 | 4.34 | 7.67 | A | A | A | 0.58 |

TABLE 16. Evaluation Results for Am-241 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 124 | 104.7 | 3.2 | -3.32 | 5.46 | A | A | A | 0.45 |
| 125 | 99.3 | 2.6 | -8.31 | 5.23 | A | A | A | 1.12 |
| 126 | 104 | 10.5 | -3.97 | 11.06 | A | A | A | 0.53 |
| 129 | 107 | 3 | -1.20 | 5.32 | A | A | A | 0.16 |
| 130 | 99 | 4 | -8.59 | 6.07 | A | A | A | 1.16 |
| 132 | 115 | 7.4 | 6.19 | 7.87 | A | A | A | 0.83 |
| 133 | 105.8 | 5.4 | -2.31 | 6.82 | A | A | A | 0.31 |
| 134 | 81.35 | 8.27 | -24.88 | 11.13 | N | A | N | 3.35 |
| 135 | 113.2 | 7.9 | 4.52 | 8.32 | A | A | A | 0.61 |
| 137 | 104.7 | 4.13 | -3.32 | 6.00 | A | A | A | 0.45 |
| 139 | 95.794 | 1.3659 | -11.55 | 4.74 | A | A | A | 1.55 |
| 144 | 106.6 | 15.5 | -1.57 | 15.23 | A | A | A | 0.21 |
| 146 | 106.6 | 10.9 | -1.57 | 11.18 | A | A | A | 0.21 |
| 147 | 102 | 10 | -5.82 | 10.80 | A | A | A | 0.78 |
| 148 | 100.68 | 8.67 | -7.04 | 9.73 | A | A | A | 0.95 |
| 149 | 110.1 | 2.5 | 1.66 | 5.06 | A | A | A | 0.22 |
| 150 | 102.8 | 3.4 | -5.08 | 5.60 | A | A | A | 0.68 |
| 151 | 107.0 | 3.0 | -1.20 | 5.32 | A | A | A | 0.16 |
| 152 | 102 | 4 | -5.82 | 5.99 | A | A | A | 0.78 |
| 153 | 112.8 | 8.0 | 4.16 | 8.41 | A | A | A | 0.56 |
| 154 | 90.3 | 7.0 | -16.62 | 8.98 | A | A | A | 2.24 |
| 155 | 120.3 | 6.7 | 11.08 | 7.18 | A | A | A | 1.49 |
| 156 | 107 | 7 | -1.20 | 7.95 | A | A | A | 0.16 |
| 157 | 96.21 | 6.30 | -11.16 | 7.96 | A | A | A | 1.50 |
| 159 | 100.9 | 1.9 | -6.83 | 4.90 | A | A | A | 0.92 |
| 160 | 105 | 3 | -3.05 | 5.35 | A | A | A | 0.41 |
| 161 | 94.8 | 3.6 | -12.47 | 5.91 | A | A | A | 1.68 |
| 162 | 107.4 | 2.7 | -0.83 | 5.18 | A | A | A | 0.11 |
| 163 | 95.5 | 9.6 | -11.82 | 11.02 | A | A | A | 1.59 |
| 164 | 115 | 6 | 6.19 | 6.91 | A | A | A | 0.83 |
| 165 | 109.9 | 7.7 | 1.48 | 8.34 | A | A | A | 0.20 |
| 166 | 119.4 | 7.5 | 10.25 | 7.74 | A | A | A | 1.38 |
| 168 | 90.11 | 10.31 | -16.80 | 12.30 | A | A | A | 2.26 |
| 169 | 112.55 | 1.38 | 3.92 | 4.69 | A | A | A | 0.53 |
| 171 | 106.20 | 10.62 | -1.94 | 10.98 | A | A | A | 0.26 |
| 172 | 137.0 | 19.8 | 26.50 | 15.14 | N | A | N | 3.57 |
| 177 | 175 | 15 | 61.59 | 9.69 | N | N | N | 8.29 |
| 179 | 115.5 | 2.8 | 6.65 | 5.13 | A | A | A | 0.89 |
| 180 | 107.4 | 5.16 | -0.83 | 6.60 | A | A | A | 0.11 |
| 181 | 105.958 | 6.122 | -2.16 | 7.34 | A | A | A | 0.29 |
| 182 | 191.35 | 8.53 | 76.69 | 6.35 | N | N | N | 10.32 |
| 183 | 109.28 | 2.88 | 0.90 | 5.24 | A | A | A | 0.12 |
| 184 | 109.37 | 0.69 | 0.99 | 4.57 | A | A | A | 0.13 |
| 185 | 112.5 | 7.3 | 3.88 | 7.91 | A | A | A | 0.52 |
| 186 | 107.4 | 6.5 | -0.83 | 7.56 | A | A | A | 0.11 |
| 189 | 117.7 | 12.0 | 8.68 | 11.15 | A | A | A | 1.17 |
| 190 | 105 | 11 | -3.05 | 11.41 | A | A | A | 0.41 |
| 191 | 105.9 | 3.8 | -2.22 | 5.77 | A | A | A | 0.30 |
| 192 | 0.034 | 0.002 | -99.97 | 7.42 | N | N | N | 13.45 |
| 193 | 98.13 | 13.74 | -9.39 | 14.71 | A | A | A | 1.26 |
| 195 | 110.0 | 8.00 | 1.57 | 8.57 | A | A | A | 0.21 |
| 196 | 116 | 11 | 7.11 | 10.51 | A | A | A | 0.96 |
| 197 | 110 | 4 | 1.57 | 5.80 | A | A | A | 0.21 |
| 198 | 105.29 | 6.50 | -2.78 | 7.65 | A | A | A | 0.37 |

TABLE 16. Evaluation Results for Am-241 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 199 | 166.49 | 3.50 | 53.73 | 4.99 | N | N | N | 7.23 |
| 205 | 107 | 4 | -1.20 | 5.87 | A | A | A | 0.16 |
| 206 | 131.9 | 13.2 | 21.79 | 10.98 | N | A | N | 2.93 |
| 207 | 107 | 6.1 | -1.20 | 7.28 | A | A | A | 0.16 |
| 208 | 105.17 | 5.6 | -2.89 | 6.99 | A | A | A | 0.39 |
| 209 | 116 | 8.6 | 7.11 | 8.69 | A | A | A | 0.96 |
| 211 | 107 | 10 | -1.20 | 10.38 | A | A | A | 0.16 |
| 212 | 106 | 4 | -2.12 | 5.89 | A | A | A | 0.29 |
| 216 | 110.1 | 5.27 | 1.66 | 6.59 | A | A | A | 0.22 |
| 219 | 111 | 24.6 | 2.49 | 22.62 | A | N | W | 0.34 |
| 222 | 56.4 | 3.7 | -47.92 | 7.97 | N | N | N | 6.45 |
| 223 | 104.5 | 3 | -3.51 | 5.36 | A | A | A | 0.47 |
| 224 | 108.2 | 3.8 | -0.09 | 5.73 | A | A | A | 0.01 |
| 225 | 100 | 5 | -7.66 | 6.74 | A | A | A | 1.03 |
| 227 | 95.48 | 2.4 | -11.84 | 5.18 | A | A | A | 1.59 |
| 228 | 104.8 | 4.2 | -3.23 | 6.04 | A | A | A | 0.43 |
| 230 | 110 | 4 | 1.57 | 5.80 | A | A | A | 0.21 |
| 231 | 99.48 | 21.34 | -8.14 | 21.92 | A | N | W | 1.10 |
| 232 | 112.3524 | 6.6606 | 3.74 | 7.46 | A | A | A | 0.50 |
| 233 | 112.83 | 1.88 | 4.18 | 4.82 | A | A | A | 0.56 |
| 234 | 124.59 | 4.3 | 15.04 | 5.69 | A | N | W | 2.02 |
| 235 | 101.73 | 6.50 | -6.07 | 7.83 | A | A | A | 0.82 |
| 237 | 109.5 | 11.1 | 1.11 | 11.10 | A | A | A | 0.15 |
| 238 | 99.113 | 12.885 | -8.48 | 13.77 | A | A | A | 1.14 |
| 239 | 104 | 4 | -3.97 | 5.94 | A | A | A | 0.53 |
| 240 | 126.6 | 4.4 | 16.90 | 5.71 | A | N | W | 2.27 |
| 241 | 120 | 14 | 10.80 | 12.51 | A | A | A | 1.45 |
| 242 | 126.8 | 6.6 | 17.08 | 6.90 | A | A | A | 2.30 |
| 243 | 106.1 | 6.8 | -2.03 | 7.85 | A | A | A | 0.27 |
| 246 | 85.56 | 3.88 | -21.00 | 6.41 | N | N | N | 2.82 |
| 248 | 118.83 | 3.0 | 9.72 | 5.18 | A | A | A | 1.31 |
| 249 | 131 | 9 | 20.96 | 8.23 | N | A | N | 2.82 |
| 250 | 102 | 5 | -5.82 | 6.67 | A | A | A | 0.78 |
| 251 | 104.6 | 5.7 | -3.42 | 7.08 | A | A | A | 0.46 |
| 254 | 107.1 | 8.0 | -1.11 | 8.73 | A | A | A | 0.15 |
| 255 | 94.6 | 3.0 | -12.65 | 5.53 | A | A | A | 1.70 |
| 256 | 102.02 | 1.42 | -5.80 | 4.73 | A | A | A | 0.78 |
| 257 | 132 | 5.5 | 21.88 | 6.15 | N | N | N | 2.94 |
| 258 | 108 | 4.89 | -0.28 | 6.40 | A | A | A | 0.04 |
| 259 | 97.1 | 0.5 | -10.34 | 4.55 | A | A | A | 1.39 |
| 261 | 110 | 8.3 | 1.57 | 8.80 | A | A | A | 0.21 |
| 262 | 123 | 7.0 | 13.57 | 7.27 | A | A | A | 1.83 |
| 263 | 99.01 | 2.75 | -8.58 | 5.31 | A | A | A | 1.15 |
| 266 | 104.7 | 10.6 | -3.32 | 11.09 | A | A | A | 0.45 |
| 267 | 86.53 | 2.6 | -20.10 | 5.43 | N | N | N | 2.70 |
| 268 | 106.1 | 4.3 | -2.03 | 6.07 | A | A | A | 0.27 |
| 270 | 117 | 8.31 | 8.03 | 8.42 | A | A | A | 1.08 |
| 274 | 109.6 | 9.3 | 1.20 | 9.62 | A | A | A | 0.16 |
| 275 | 98.9 | 6.16 | -8.68 | 7.70 | A | A | A | 1.17 |
| 279 | 107.1 | 10.9 | -1.11 | 11.14 | A | A | A | 0.15 |
| 281 | 86 | 3 | -20.59 | 5.71 | N | N | N | 2.77 |
| 283 | 102 | 5.4 | -5.82 | 6.96 | A | A | A | 0.78 |
| 287 | 103.9 | 0.9 | -4.06 | 4.61 | A | A | A | 0.55 |
| 288 | 103.5 | 9.0 | -4.43 | 9.80 | A | A | A | 0.60 |

TABLE 16. Evaluation Results for Am-241 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 289 | 112 | 6 | 3.42 | 7.01 | A | A | A | 0.46 |
| 290 | 84.50 | 9.92 | -21.98 | 12.58 | N | A | N | 2.96 |
| 291 | 111 | 4 | 2.49 | 5.78 | A | A | A | 0.34 |
| 295 | 859.0 | 80.2 | 693.17 | 10.37 | N | N | N | 93.25 |
| 298 | 113.5 | 2.3 | 4.80 | 4.96 | A | A | A | 0.65 |
| 299 | 77.9 | 4.9 | -28.07 | 7.75 | N | N | N | 3.78 |
| 304 | 131.02 | 10.48 | 20.98 | 9.19 | N | A | N | 2.82 |
| 305 | 99.5 | 10.28 | -8.13 | 11.28 | A | A | A | 1.09 |
| 310 | 105.6 | 3.7 | -2.49 | 5.72 | A | A | A | 0.34 |
| 312 | 95.95 | 1.82 | -11.40 | 4.91 | A | A | A | 1.53 |
| 313 | 113.5 | 5.2 | 4.80 | 6.44 | A | A | A | 0.65 |
| 314 | 112.92 | 5.59 | 4.27 | 6.71 | A | A | A | 0.57 |
| 317 | 113 |  | 4.34 |  | A |  |  | 0.58 |
| 320 | 111.39 | 4.69 | 2.85 | 6.18 | A | A | A | 0.38 |
| 321 | 112.56 | 1.2 | 3.93 | 4.65 | A | A | A | 0.53 |
| 322 | 97.0 | 15.6 | -10.43 | 16.71 | A | A | A | 1.40 |
| 323 | 109.6 | 2.9 | 1.20 | 5.24 | A | A | A | 0.16 |
| 324 | 114.7 | 2.74 | 5.91 | 5.12 | A | A | A | 0.80 |
| 325 | 107.7 | 9.3 | -0.55 | 9.75 | A | A | A | 0.07 |

## Reported Results for K-40 Determination in Sample 4

TABLE 17. Target values

| Parameter | Value |
| --- | --- |
| Target Value, Bq/kg | 97.9 |
| UNC, Bq/kg | 7.2 |
| MARB, % | 30 |

FIGURE 17. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

FIGURE 18. Reported results and evaluation parameters for K-40. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

TABLE 18. Evaluation Results for K-40 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 2 | 147.5 | 22 | 50.66 | 16.63 | N | N | N | 3.06 |
| 3 | 91 | 6 | -7.05 | 9.88 | A | A | A | 0.43 |
| 4 | 97.9 | 5.3 | 0.00 | 9.13 | A | A | A | 0.00 |
| 5 | 63 | 6.1 | -35.65 | 12.16 | N | N | N | 2.15 |
| 6 | 80.9 | 10.3 | -17.36 | 14.70 | A | A | A | 1.05 |
| 7 | 71.37 | 8.92 | -27.10 | 14.50 | A | A | A | 1.64 |
| 8 | 103.92 | 6.90 | 6.15 | 9.91 | A | A | A | 0.37 |
| 9 | 110 | 23 | 12.36 | 22.16 | A | A | A | 0.75 |
| 10 | 90.2 | 8.2 | -7.87 | 11.69 | A | A | A | 0.47 |
| 11 | 129.4 | 9.2 | 32.18 | 10.23 | N | N | N | 1.94 |
| 13 | 88.66 | 5.54 | -9.44 | 9.65 | A | A | A | 0.57 |
| 16 | 73.9 | 20.5 | -24.51 | 28.70 | A | A | A | 1.48 |
| 18 | 83.7 | 5.7 | -14.50 | 10.02 | A | A | A | 0.88 |
| 19 | 93.1 | 8.2 | -4.90 | 11.47 | A | A | A | 0.30 |
| 21 | 72.2 | 12.6 | -26.25 | 18.94 | A | A | A | 1.58 |
| 25 | 140 | 20 | 43.00 | 16.07 | N | N | N | 2.60 |
| 27 | 113 | 17 | 15.42 | 16.75 | A | A | A | 0.93 |
| 29 | 106.29 | 12.38 | 8.57 | 13.77 | A | A | A | 0.52 |
| 32 | 107.0 | 22.5 | 9.30 | 22.28 | A | A | A | 0.56 |
| 35 | 475 | 26 | 385.19 | 9.17 | N | N | N | 23.25 |
| 38 | 128 | 14 | 30.75 | 13.18 | N | A | N | 1.86 |
| 39 | 84.0 | 4.5 | -14.20 | 9.10 | A | A | A | 0.86 |
| 47 | 518 | 9.4 | 429.11 | 7.58 | N | N | N | 25.90 |
| 50 | 154.93 | 59.31 | 58.25 | 38.98 | N | N | N | 3.52 |
| 52 | 106.2 | 9.5 | 8.48 | 11.58 | A | A | A | 0.51 |
| 55 | 94.2 | 12.3 | -3.78 | 14.99 | A | A | A | 0.23 |
| 66 | 132.8 | 6 | 35.65 | 8.63 | N | N | N | 2.15 |
| 68 | 98.2 | 6.9 | 0.31 | 10.17 | A | A | A | 0.02 |
| 74 | 109.0 | 10.0 | 11.34 | 11.76 | A | A | A | 0.68 |
| 84 | 92.08 | 7.7 | -5.94 | 11.14 | A | A | A | 0.36 |

TABLE 18. Evaluation Results for K-40 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 85 | 84.30 | 8.36 | -13.89 | 12.35 | A | A | A | 0.84 |
| 87 | 110 | 5 | 12.36 | 8.65 | A | A | A | 0.75 |
| 91 | 89.6 | 3.4 | -8.48 | 8.28 | A | A | A | 0.51 |
| 102 | 78.1 | 5.9 | -20.22 | 10.54 | A | A | A | 1.22 |
| 107 | 87.7 | 7.7 | -10.42 | 11.45 | A | A | A | 0.63 |
| 109 | 106.2 | 10.2 | 8.48 | 12.10 | A | A | A | 0.51 |
| 113 | 88.3 | 18.9 | -9.81 | 22.63 | A | A | A | 0.59 |
| 121 | 77.1 | 3.5 | -21.25 | 8.64 | A | A | A | 1.28 |
| 122 | 88.0 | 5.0 | -10.11 | 9.29 | A | A | A | 0.61 |
| 123 | 136 | 8 | 38.92 | 9.42 | N | N | N | 2.35 |
| 124 | 92.2 | 2.1 | -5.82 | 7.70 | A | A | A | 0.35 |
| 125 | 118 | 40 | 20.53 | 34.69 | A | N | W | 1.24 |
| 126 | 51.8 | 4.3 | -47.09 | 11.09 | N | N | N | 2.84 |
| 129 | 91 | 6 | -7.05 | 9.88 | A | A | A | 0.43 |
| 132 | 59.0 | 7.7 | -39.73 | 14.98 | N | N | N | 2.40 |
| 133 | 86.2 | 5.5 | -11.95 | 9.74 | A | A | A | 0.72 |
| 139 | 78.775 | 10.528 | -19.54 | 15.25 | A | A | A | 1.18 |
| 146 | 97.6 | 11.6 | -0.31 | 13.98 | A | A | A | 0.02 |
| 148 | 91.11 | 10.33 | -6.94 | 13.51 | A | A | A | 0.42 |
| 149 | 78.8 | 3.4 | -19.51 | 8.53 | A | A | A | 1.18 |
| 150 | 80.9 | 4.2 | -17.36 | 9.00 | A | A | A | 1.05 |
| 151 | 91.0 | 6.0 | -7.05 | 9.88 | A | A | A | 0.43 |
| 152 | 88.4 | 10.9 | -9.70 | 14.36 | A | A | A | 0.59 |
| 153 | 118 | 19 | 20.53 | 17.70 | A | A | A | 1.24 |
| 155 | 99.7 | 7.5 | 1.84 | 10.52 | A | A | A | 0.11 |
| 156 | 91 | 14 | -7.05 | 17.05 | A | A | A | 0.43 |
| 163 | 81.1 | 4.9 | -17.16 | 9.52 | A | A | A | 1.04 |
| 172 | 146.2 | 27.5 | 49.34 | 20.20 | N | A | N | 2.98 |
| 176 | 95.2 | 10.1 | -2.76 | 12.91 | A | A | A | 0.17 |
| 180 | 75.46 | 9.95 | -22.92 | 15.10 | A | A | A | 1.38 |
| 182 | 81.03 | 10.46 | -17.23 | 14.86 | A | A | A | 1.04 |
| 184 | 93.30 | 2.66 | -4.70 | 7.89 | A | A | A | 0.28 |
| 187 | 10787 | 338 | 10918.39 | 7.99 | N | N | N | 659.01 |
| 191 | 91 | 6 | -7.05 | 9.88 | A | A | A | 0.43 |
| 193 | 96.6 | 16.4 | -1.33 | 18.50 | A | A | A | 0.08 |
| 195 | 83.3 | 10.0 | -14.91 | 14.08 | A | A | A | 0.90 |
| 196 | 72.9 | 2.8 | -25.54 | 8.30 | A | N | W | 1.54 |
| 198 | 95.17 | 6.14 | -2.79 | 9.78 | A | A | A | 0.17 |
| 199 | 95.58 | 5.92 | -2.37 | 9.62 | A | A | A | 0.14 |
| 205 | 95 | 8 | -2.96 | 11.18 | A | A | A | 0.18 |
| 209 | 102.7 | 27.8 | 4.90 | 28.05 | A | A | A | 0.30 |
| 224 | 105.5 | 7.8 | 7.76 | 10.43 | A | A | A | 0.47 |
| 227 | 99.62 | 2.59 | 1.76 | 7.80 | A | A | A | 0.11 |
| 230 | 71 | 5 | -27.48 | 10.18 | A | N | W | 1.66 |
| 231 | 116.70 | 14.23 | 19.20 | 14.24 | A | A | A | 1.16 |
| 233 | 79.13 | 5.03 | -19.17 | 9.72 | A | A | A | 1.16 |
| 234 | 97.60 | 9.68 | -0.31 | 12.35 | A | A | A | 0.02 |
| 238 | 80.042 | 9.605 | -18.24 | 14.07 | A | A | A | 1.10 |
| 241 | 91.7 | 4.4 | -6.33 | 8.78 | A | A | A | 0.38 |
| 245 | 92.7 | 6.91 | -5.31 | 10.47 | A | A | A | 0.32 |
| 248 | 383.70 | 13.43 | 291.93 | 8.14 | N | N | N | 17.62 |
| 249 | 105 | 22 | 7.25 | 22.21 | A | A | A | 0.44 |
| 250 | 96 | 7 | -1.94 | 10.36 | A | A | A | 0.12 |
| 251 | 83.5 | 5.6 | -14.71 | 9.95 | A | A | A | 0.89 |

TABLE 18. Evaluation Results for K-40 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 254 | 78.7 | 8.7 | -19.61 | 13.28 | A | A | A | 1.18 |
| 256 | 73.25 | 5.95 | -25.18 | 10.96 | A | A | A | 1.52 |
| 257 | 85 | 7 | -13.18 | 11.04 | A | A | A | 0.80 |
| 260 | 121.6 | 14.0 | 24.21 | 13.66 | A | A | A | 1.46 |
| 261 | 96.5 | 9.4 | -1.43 | 12.21 | A | A | A | 0.09 |
| 268 | 80.0 | 6.5 | -18.28 | 10.96 | A | A | A | 1.10 |
| 283 | 77.8 | 9.3 | -20.53 | 14.03 | A | A | A | 1.24 |
| 290 | 97.68 | 8.53 | -0.22 | 11.42 | A | A | A | 0.01 |
| 299 | 197.8 | 11.8 | 102.04 | 9.47 | N | N | N | 6.16 |
| 308 | 13226.73 | 547.96 | 13410.45 | 8.44 | N | N | N | 809.42 |
| 316 | 106.25 | 20.19 | 8.53 | 20.38 | A | A | A | 0.51 |
| 317 | 147 |  | 50.15 |  | N |  | N | 3.03 |
| 322 | 138.2 | 23.1 | 41.16 | 18.26 | N | A | N | 2.48 |
| 323 | 80.5 | 6.7 | -17.77 | 11.11 | A | A | A | 1.07 |
| 324 | 79.5 | 0.14 | -18.79 | 7.36 | A | A | A | 1.13 |
| 325 | 87.4 | 6.5 | -10.73 | 10.46 | A | A | A | 0.65 |

## Reported Results for Cs-134 Determination in Sample 4

TABLE 19. Target values

| Parameter | Value |
| --- | --- |
| Target Value, Bq/kg | 86.6 |
| UNC, Bq/kg | 4 |
| MARB, % | 25 |

FIGURE 19. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

FIGURE 20. Reported results and evaluation parameters for Cs-134. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

TABLE 20. Evaluation Results for Cs-134 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 63.78 | 2.57 | -26.35 | 6.13 | N | N | N | 2.28 |
| 2 | 83.5 | 6.5 | -3.58 | 9.05 | A | A | A | 0.31 |
| 3 | 83.5 | 4.5 | -3.58 | 7.10 | A | A | A | 0.31 |
| 4 | 89.0 | 4.5 | 2.77 | 6.85 | A | A | A | 0.24 |
| 5 | 68 | 5 | -21.48 | 8.68 | A | A | A | 1.86 |
| 6 | 62.8 | 1.7 | -27.48 | 5.35 | N | N | N | 2.38 |
| 7 | 83.50 | 2.09 | -3.58 | 5.25 | A | A | A | 0.31 |
| 8 | 77.97 | 0.97 | -9.97 | 4.78 | A | A | A | 0.86 |
| 9 | 90 | 15 | 3.93 | 17.29 | A | A | A | 0.34 |
| 10 | 88.82 | 7.06 | 2.56 | 9.19 | A | A | A | 0.22 |
| 11 | 70.2 | 2.3 | -18.94 | 5.66 | A | N | W | 1.64 |
| 12 | 61.5 | 4.4 | -28.98 | 8.52 | N | N | N | 2.51 |
| 13 | 87.27 | 4.12 | 0.77 | 6.60 | A | A | A | 0.07 |
| 14 | 2089 | 111 | 2312.24 | 7.04 | N | N | N | 200.24 |
| 15 | 83.8 | 4.3 | -3.23 | 6.90 | A | A | A | 0.28 |
| 16 | 94.6 | 2 | 9.24 | 5.08 | A | A | A | 0.80 |
| 17 | 88.01 | 4.96 | 1.63 | 7.29 | A | A | A | 0.14 |
| 18 | 84.4 | 6.8 | -2.54 | 9.29 | A | A | A | 0.22 |
| 19 | 85.0 | 10.9 | -1.85 | 13.63 | A | A | A | 0.16 |
| 21 | 80.6 | 5.9 | -6.93 | 8.66 | A | A | A | 0.60 |
| 25 | 89 | 5 | 2.77 | 7.27 | A | A | A | 0.24 |
| 26 | 91.64 | 4.82 | 5.82 | 7.00 | A | A | A | 0.50 |
| 27 | 136 | 20 | 57.04 | 15.41 | N | N | N | 4.94 |
| 29 | 75.53 | 2.47 | -12.78 | 5.66 | A | A | A | 1.11 |
| 31 | 70.10 | 4.21 | -19.05 | 7.58 | A | A | A | 1.65 |
| 32 | 90.1 | 14.4 | 4.04 | 16.64 | A | A | A | 0.35 |
| 33 | 85 | 4 | -1.85 | 6.59 | A | A | A | 0.16 |
| 34 | 81.7 | 6.8 | -5.66 | 9.52 | A | A | A | 0.49 |
| 35 | 108 | 7 | 24.71 | 7.96 | A | N | W | 2.14 |
| 36 | 84.6 | 4 | -2.31 | 6.61 | A | A | A | 0.20 |

TABLE 20. Evaluation Results for Cs-134 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 38 | 93.8 | 5.6 | 8.31 | 7.55 | A | A | A | 0.72 |
| 39 | 73.7 | 3.5 | -14.90 | 6.62 | A | A | A | 1.29 |
| 41 | 87.5 | 4.0 | 1.04 | 6.50 | A | A | A | 0.09 |
| 42 | 54.89 | 1.23 | -36.62 | 5.13 | N | N | N | 3.17 |
| 43 | 85.59 | 3.31 | -1.17 | 6.02 | A | A | A | 0.10 |
| 45 | 94.7 | 8.5 | 9.35 | 10.09 | A | A | A | 0.81 |
| 46 | 144.59 | 4.37 | 66.96 | 5.52 | N | N | N | 5.80 |
| 47 | 65.6 | 3.7 | -24.25 | 7.29 | A | N | W | 2.10 |
| 48 | 81.8 | 3.5 | -5.54 | 6.30 | A | A | A | 0.48 |
| 50 | 88.93 | 7.79 | 2.69 | 9.90 | A | A | A | 0.23 |
| 51 | 82 | 2 | -5.31 | 5.22 | A | A | A | 0.46 |
| 52 | 81.2 | 4.2 | -6.24 | 6.93 | A | A | A | 0.54 |
| 53 | 89.0 | 5.0 | 2.77 | 7.27 | A | A | A | 0.24 |
| 55 | 76.3 | 1.8 | -11.89 | 5.19 | A | A | A | 1.03 |
| 56 | 67.23 | 4.79 | -22.37 | 8.49 | A | N | W | 1.94 |
| 59 | 86.5 | 6.2 | -0.12 | 8.53 | A | A | A | 0.01 |
| 61 | 76.45 | 3.33 | -11.72 | 6.35 | A | A | A | 1.01 |
| 63 | 90.9 | 2.7 | 4.97 | 5.49 | A | A | A | 0.43 |
| 64 | 77.1 | 5 | -10.97 | 7.96 | A | A | A | 0.95 |
| 66 | 88.3 | 2.2 | 1.96 | 5.25 | A | A | A | 0.17 |
| 68 | 86.7 | 7.1 | 0.12 | 9.40 | A | A | A | 0.01 |
| 70 | 84.47 | 1.06 | -2.46 | 4.79 | A | A | A | 0.21 |
| 72 | 90.173 | 4.564 | 4.13 | 6.85 | A | A | A | 0.36 |
| 74 | 91.9 | 4.3 | 6.12 | 6.57 | A | A | A | 0.53 |
| 76 | 83.50 | 1.62 | -3.58 | 5.01 | A | A | A | 0.31 |
| 77 | 66 | 1.2 | -23.79 | 4.96 | A | N | W | 2.06 |
| 78 | 75.6 | 1.6 | -12.70 | 5.08 | A | A | A | 1.10 |
| 80 | 85.1 | 8.3 | -1.73 | 10.79 | A | A | A | 0.15 |
| 81 | 88.4 | 7.5 | 2.08 | 9.66 | A | A | A | 0.18 |
| 84 | 91.92 | 3.89 | 6.14 | 6.26 | A | A | A | 0.53 |
| 85 | 84.93 | 6.88 | -1.93 | 9.33 | A | A | A | 0.17 |
| 86 | 81.3 | 4.0 | -6.12 | 6.75 | A | A | A | 0.53 |
| 87 | 90 | 4 | 3.93 | 6.41 | A | A | A | 0.34 |
| 88 | 81.1 | 7.8 | -6.35 | 10.67 | A | A | A | 0.55 |
| 89 | 277 | 19 | 219.86 | 8.27 | N | N | N | 19.04 |
| 92 | 72.8 | 5.6 | -15.94 | 8.97 | A | A | A | 1.38 |
| 93 | 56.08 | 0.90 | -35.24 | 4.89 | N | N | N | 3.05 |
| 94 | 83.93 | 0.52 | -3.08 | 4.66 | A | A | A | 0.27 |
| 95 | 85.50 | 0.52 | -1.27 | 4.66 | A | A | A | 0.11 |
| 99 | 98.4 | 6.2 | 13.63 | 7.81 | A | A | A | 1.18 |
| 101 | 86.7 | 3.0 | 0.12 | 5.77 | A | A | A | 0.01 |
| 102 | 83 | 6 | -4.16 | 8.58 | A | A | A | 0.36 |
| 103 | 97.56 | 2.12 | 12.66 | 5.10 | A | A | A | 1.10 |
| 106 | 86.5 | 1.5 | -0.12 | 4.93 | A | A | A | 0.01 |
| 107 | 61.9 | 1.1 | -28.52 | 4.95 | N | N | N | 2.47 |
| 108 | 87.9296 | 1.0036 | 1.54 | 4.76 | A | A | A | 0.13 |
| 109 | 89.9 | 4.2 | 3.81 | 6.57 | A | A | A | 0.33 |
| 110 | 72.015 | 1.366 | -16.84 | 4.99 | A | N | W | 1.46 |
| 111 | 66.1 | 4.6 | -23.67 | 8.35 | A | N | W | 2.05 |
| 112 | 140.4195 | 7.920 | 62.15 | 7.29 | N | N | N | 5.38 |
| 113 | 75.6 | 4.2 | -12.70 | 7.22 | A | A | A | 1.10 |
| 116 | 105.0 | 3.3 | 21.25 | 5.59 | A | N | W | 1.84 |
| 120 | 83.0 | 4.1 | -4.16 | 6.76 | A | A | A | 0.36 |
| 121 | 78.8 | 4.2 | -9.01 | 7.05 | A | A | A | 0.78 |

TABLE 20. Evaluation Results for Cs-134 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 122 | 80.0 | 3.2 | -7.62 | 6.11 | A | A | A | 0.66 |
| 123 | 83 | 6 | -4.16 | 8.58 | A | A | A | 0.36 |
| 124 | 78.7 | 2.9 | -9.12 | 5.91 | A | A | A | 0.79 |
| 125 | 75.0 | 4.0 | -13.39 | 7.06 | A | A | A | 1.16 |
| 126 | 62.5 | 3.8 | -27.83 | 7.64 | N | N | N | 2.41 |
| 129 | 81 | 3 | -6.47 | 5.92 | A | A | A | 0.56 |
| 130 | 70 | 3 | -19.17 | 6.30 | A | N | W | 1.66 |
| 132 | 69.7 | 4.2 | -19.52 | 7.59 | A | A | A | 1.69 |
| 133 | 79.0 | 6.0 | -8.78 | 8.89 | A | A | A | 0.76 |
| 135 | 84.4 | 4.1 | -2.54 | 6.70 | A | A | A | 0.22 |
| 137 | 78.38 | 1.154 | -9.49 | 4.85 | A | A | A | 0.82 |
| 139 | 62.501 | 1.3945 | -27.83 | 5.13 | N | N | N | 2.41 |
| 140 | 84.2 | 3.2 | -2.77 | 5.98 | A | A | A | 0.24 |
| 141 | 79.8 | 5.4 | -7.85 | 8.19 | A | A | A | 0.68 |
| 143 | 84.3 | 5.2 | -2.66 | 7.71 | A | A | A | 0.23 |
| 144 | 83.92 | 5.61 | -3.09 | 8.13 | A | A | A | 0.27 |
| 146 | 90.5 | 2.2 | 4.50 | 5.22 | A | A | A | 0.39 |
| 147 | 84.8 | 7.5 | -2.08 | 9.98 | A | A | A | 0.18 |
| 148 | 73.1 | 3.01 | -15.59 | 6.19 | A | A | A | 1.35 |
| 149 | 66.7 | 3.2 | -22.98 | 6.66 | A | N | W | 1.99 |
| 150 | 84.3 | 3.5 | -2.66 | 6.21 | A | A | A | 0.23 |
| 151 | 79.0 | 5.0 | -8.78 | 7.84 | A | A | A | 0.76 |
| 152 | 88.1 | 3.5 | 1.73 | 6.09 | A | A | A | 0.15 |
| 153 | 78.2 | 5.3 | -9.70 | 8.20 | A | A | A | 0.84 |
| 154 | 71.9 | 2.7 | -16.97 | 5.95 | A | N | W | 1.47 |
| 155 | 84.1 | 3.4 | -2.89 | 6.14 | A | A | A | 0.25 |
| 156 | 88.7 | 6.0 | 2.42 | 8.19 | A | A | A | 0.21 |
| 157 | 79.77 | 4.33 | -7.89 | 7.13 | A | A | A | 0.68 |
| 160 | 67.7 | 3.5 | -21.82 | 6.93 | A | N | W | 1.89 |
| 161 | 79.0 | 2.4 | -8.78 | 5.53 | A | A | A | 0.76 |
| 162 | 78.8 | 1.9 | -9.01 | 5.21 | A | A | A | 0.78 |
| 163 | 69.9 | 2.8 | -19.28 | 6.11 | A | N | W | 1.67 |
| 164 | 89 | 5 | 2.77 | 7.27 | A | A | A | 0.24 |
| 165 | 65.3 | 3.9 | -24.60 | 7.55 | A | N | W | 2.13 |
| 166 | 88.2 | 5.0 | 1.85 | 7.31 | A | A | A | 0.16 |
| 168 | 76.98 | 4.29 | -11.11 | 7.24 | A | A | A | 0.96 |
| 169 | 85.19 | 3.78 | -1.63 | 6.40 | A | A | A | 0.14 |
| 171 | 82.50 | 8.30 | -4.73 | 11.07 | A | A | A | 0.41 |
| 172 | 79.9 | 11.9 | -7.74 | 15.59 | A | A | A | 0.67 |
| 173 | 113.0 | 4.0 | 30.48 | 5.82 | N | N | N | 2.64 |
| 176 | 2100 | 115 | 2324.94 | 7.16 | N | N | N | 201.34 |
| 180 | 82.59 | 3.42 | -4.63 | 6.20 | A | A | A | 0.40 |
| 181 | 76.25 | 2.45 | -11.95 | 5.63 | A | A | A | 1.03 |
| 182 | 80.93 | 5.46 | -6.55 | 8.18 | A | A | A | 0.57 |
| 183 | 84.81 | 3.4 | -2.07 | 6.12 | A | A | A | 0.18 |
| 184 | 79.39 | 1.18 | -8.33 | 4.85 | A | A | A | 0.72 |
| 185 | 94.6 | 7.8 | 9.24 | 9.45 | A | A | A | 0.80 |
| 186 | 87.4 | 6.0 | 0.92 | 8.27 | A | A | A | 0.08 |
| 187 | 55.3 | 0.7 | -36.14 | 4.79 | N | N | N | 3.13 |
| 189 | 81.0 | 4.9 | -6.47 | 7.61 | A | A | A | 0.56 |
| 190 | 100 | 12 | 15.47 | 12.86 | A | A | A | 1.34 |
| 191 | 82.4 | 4.0 | -4.85 | 6.70 | A | A | A | 0.42 |
| 193 | 100.9 | 10.1 | 16.51 | 11.02 | A | A | A | 1.43 |
| 195 | 81.1 | 4.50 | -6.35 | 7.22 | A | A | A | 0.55 |

TABLE 20. Evaluation Results for Cs-134 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 196 | 76.6 | 3.9 | -11.55 | 6.87 | A | A | A | 1.00 |
| 197 | 88.2 | 2.6 | 1.85 | 5.48 | A | A | A | 0.16 |
| 198 | 84.78 | 4.56 | -2.10 | 7.09 | A | A | A | 0.18 |
| 199 | 53.80 | 1.67 | -37.88 | 5.57 | N | N | N | 3.28 |
| 200 | 77.74 | 1.38 | -10.23 | 4.95 | A | A | A | 0.89 |
| 201 | 76.2 | 7.6 | -12.01 | 10.99 | A | A | A | 1.04 |
| 202 | 79.94 | 4.69 | -7.69 | 7.47 | A | A | A | 0.67 |
| 205 | 82.1 | 3.1 | -5.20 | 5.97 | A | A | A | 0.45 |
| 206 | 88.4 | 4.3 | 2.08 | 6.71 | A | A | A | 0.18 |
| 207 | 77 | 4.7 | -11.09 | 7.65 | A | A | A | 0.96 |
| 208 | 64.2 | 2.89 | -25.87 | 6.45 | N | N | N | 2.24 |
| 209 | 68.8 | 7.3 | -20.55 | 11.57 | A | A | A | 1.78 |
| 210 | 95.1 | 7.2 | 9.82 | 8.87 | A | A | A | 0.85 |
| 212 | 90.7 | 6.8 | 4.73 | 8.81 | A | A | A | 0.41 |
| 213 | 92.2 | 6.1 | 6.47 | 8.07 | A | A | A | 0.56 |
| 214 | 67.5 | 4.7 | -22.06 | 8.36 | A | N | W | 1.91 |
| 215 | 81.4 | 5.1 | -6.00 | 7.78 | A | A | A | 0.52 |
| 216 | 79.83 | 3.97 | -7.82 | 6.79 | A | A | A | 0.68 |
| 219 | 90.8 | 8.4 | 4.85 | 10.34 | A | A | A | 0.42 |
| 221 | 68.65 | 2.42 | -20.73 | 5.81 | A | N | W | 1.79 |
| 222 | 62.2 | 1.77 | -28.18 | 5.43 | N | N | N | 2.44 |
| 224 | 88.2 | 4.1 | 1.85 | 6.55 | A | A | A | 0.16 |
| 225 | 93 | 5 | 7.39 | 7.09 | A | A | A | 0.64 |
| 227 | 86.41 | 1.54 | -0.22 | 4.95 | A | A | A | 0.02 |
| 230 | 68.5 | 3 | -20.90 | 6.37 | A | N | W | 1.81 |
| 231 | 79.9 | 12.91 | -7.74 | 16.80 | A | A | A | 0.67 |
| 232 | 80.4733 | 3.4637 | -7.07 | 6.31 | A | A | A | 0.61 |
| 233 | 81.53 | 1.86 | -5.85 | 5.15 | A | A | A | 0.51 |
| 234 | 79.01 | 4.21 | -8.76 | 7.05 | A | A | A | 0.76 |
| 237 | 81.0 | 3.1 | -6.47 | 6.00 | A | A | A | 0.56 |
| 238 | 64.968 | 7.796 | -24.98 | 12.86 | A | A | A | 2.16 |
| 239 | 76.3 | 2.8 | -11.89 | 5.90 | A | A | A | 1.03 |
| 240 | 82.2 | 1.6 | -5.08 | 5.01 | A | A | A | 0.44 |
| 241 | 82.8 | 2.7 | -4.39 | 5.65 | A | A | A | 0.38 |
| 242 | 78.7 | 4.4 | -9.12 | 7.25 | A | A | A | 0.79 |
| 245 | 73.4 | 2.14 | -15.24 | 5.46 | A | N | W | 1.32 |
| 246 | 60.33 | 1.33 | -30.33 | 5.12 | N | N | N | 2.63 |
| 248 | 56.39 | 1.41 | -34.88 | 5.25 | N | N | N | 3.02 |
| 249 | 90.2 | 3.4 | 4.16 | 5.96 | A | A | A | 0.36 |
| 250 | 75.8 | 2.2 | -12.47 | 5.46 | A | A | A | 1.08 |
| 251 | 80.7 | 1.0 | -6.81 | 4.78 | A | A | A | 0.59 |
| 252 | 65.77 | 3.34 | -24.05 | 6.86 | A | N | W | 2.08 |
| 254 | 83.7 | 4.6 | -3.35 | 7.18 | A | A | A | 0.29 |
| 255 | 73.0 | 2.3 | -15.70 | 5.59 | A | N | W | 1.36 |
| 256 | 63.58 | 1.77 | -26.58 | 5.39 | N | N | N | 2.30 |
| 257 | 70 | 5 | -19.17 | 8.51 | A | A | A | 1.66 |
| 258 | 87.7 | 3.71 | 1.27 | 6.26 | A | A | A | 0.11 |
| 259 | 71.8 | 1.13 | -17.09 | 4.88 | A | N | W | 1.48 |
| 260 | 83.0 | 7.1 | -4.16 | 9.72 | A | A | A | 0.36 |
| 261 | 85.0 | 4.7 | -1.85 | 7.20 | A | A | A | 0.16 |
| 262 | 88.9 | 5.3 | 2.66 | 7.54 | A | A | A | 0.23 |
| 263 | 0.079 | 0.0045 | -99.91 | 7.33 | N | N | N | 8.65 |
| 266 | 69.6 | 2.6 | -19.63 | 5.94 | A | N | W | 1.70 |
| 267 | 66.79 | 1.0 | -22.88 | 4.86 | A | N | W | 1.98 |

TABLE 20. Evaluation Results for Cs-134 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 268 | 92.4 | 3.7 | 6.70 | 6.11 | A | A | A | 0.58 |
| 273 | 80.43 | 3.61 | -7.12 | 6.44 | A | A | A | 0.62 |
| 274 | 88.7 | 2.6 | 2.42 | 5.47 | A | A | A | 0.21 |
| 279 | 83.8 | 3.6 | -3.23 | 6.31 | A | A | A | 0.28 |
| 281 | 67 | 3 | -22.63 | 6.43 | A | N | W | 1.96 |
| 283 | 74 | 2.5 | -14.55 | 5.72 | A | A | A | 1.26 |
| 287 | 80.6 | 3 | -6.93 | 5.93 | A | A | A | 0.60 |
| 289 | 84.5 | 4.5 | -2.42 | 7.05 | A | A | A | 0.21 |
| 290 | 46.45 | 2.75 | -46.36 | 7.51 | N | N | N | 4.01 |
| 291 | 93.6 | 3.8 | 8.08 | 6.15 | A | A | A | 0.70 |
| 295 | 70.0 | 4.2 | -19.17 | 7.57 | A | A | A | 1.66 |
| 297 | 60.3 | 3.8 | -30.37 | 7.81 | N | N | N | 2.63 |
| 299 | 94.8 | 5.6 | 9.47 | 7.50 | A | A | A | 0.82 |
| 305 | 58.8 | 4.10 | -32.10 | 8.36 | N | N | N | 2.78 |
| 307 | 55.1 | 0.8 | -36.37 | 4.84 | N | N | N | 3.15 |
| 308 | 47.02 | 2.15 | -45.70 | 6.50 | N | N | N | 3.96 |
| 310 | 71.8 | 1.9 | -17.09 | 5.32 | A | N | W | 1.48 |
| 312 | 92.56 | 1.37 | 6.88 | 4.85 | A | A | A | 0.60 |
| 313 | 78.7 | 2.0 | -9.12 | 5.27 | A | A | A | 0.79 |
| 314 | 83.23 | 11.27 | -3.89 | 14.31 | A | A | A | 0.34 |
| 316 | 73.125 | 4.39 | -15.56 | 7.57 | A | A | A | 1.35 |
| 317 | 76 |  | -12.24 |  | A |  |  | 1.06 |
| 322 | 98.3 | 7.1 | 13.51 | 8.57 | A | A | A | 1.17 |
| 323 | 91.9 | 2.8 | 6.12 | 5.53 | A | A | A | 0.53 |
| 324 | 68.49 | 1.34 | -20.91 | 5.02 | A | N | W | 1.81 |
| 325 | 84.7 | 6.8 | -2.19 | 9.26 | A | A | A | 0.19 |

## Reported Results for Cs-137 Determination in Sample 4

TABLE 21. Target values

| Parameter | Value |
| --- | --- |
| Target Value, Bq/kg | 2063 |
| UNC, Bq/kg | 93 |
| MARB, % | 20 |

FIGURE 21. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

FIGURE 22. Reported results and evaluation parameters for Cs-137. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

TABLE 22. Evaluation Results for Cs-137 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 1996.20 | 86.26 | -3.24 | 6.24 | A | A | A | 0.37 |
| 2 | 1885.4 | 115 | -8.61 | 7.58 | A | A | A | 0.98 |
| 3 | 1930 | 96 | -6.45 | 6.71 | A | A | A | 0.74 |
| 4 | 2206 | 110 | 6.93 | 6.72 | A | A | A | 0.79 |
| 5 | 1571 | 93 | -23.85 | 7.44 | N | N | N | 2.73 |
| 6 | 1990.5 | 69.2 | -3.51 | 5.69 | A | A | A | 0.40 |
| 7 | 2189.87 | 61.01 | 6.15 | 5.30 | A | A | A | 0.70 |
| 8 | 1940.50 | 16.91 | -5.94 | 4.59 | A | A | A | 0.68 |
| 9 | 2110 | 340 | 2.28 | 16.73 | A | A | A | 0.26 |
| 10 | 2001.8 | 10.9 | -2.97 | 4.54 | A | A | A | 0.34 |
| 11 | 1652 | 54 | -19.92 | 5.57 | A | N | W | 2.28 |
| 12 | 2033 | 40 | -1.45 | 4.92 | A | A | A | 0.17 |
| 13 | 2025.0 | 94.76 | -1.84 | 6.50 | A | A | A | 0.21 |
| 14 | 87.9 | 4.9 | -95.74 | 7.17 | N | N | N | 10.95 |
| 15 | 1924.1 | 98.1 | -6.73 | 6.81 | A | A | A | 0.77 |
| 16 | 2124.5 | 46.2 | 2.98 | 5.01 | A | A | A | 0.34 |
| 17 | 2079.0 | 110.8 | 0.78 | 6.98 | A | A | A | 0.09 |
| 18 | 1940 | 120 | -5.96 | 7.65 | A | A | A | 0.68 |
| 19 | 2123 | 204 | 2.91 | 10.61 | A | A | A | 0.33 |
| 21 | 1884 | 133 | -8.68 | 8.38 | A | A | A | 0.99 |
| 25 | 2000 | 100 | -3.05 | 6.73 | A | A | A | 0.35 |
| 26 | 1878.10 | 95.04 | -8.96 | 6.78 | A | A | A | 1.03 |
| 27 | 2423 | 244 | 17.45 | 11.03 | A | A | A | 2.00 |
| 29 | 1756.95 | 53.06 | -14.84 | 5.43 | A | N | W | 1.70 |
| 31 | 2013.7 | 94.00 | -2.39 | 6.49 | A | A | A | 0.27 |
| 32 | 2109.8 | 337.6 | 2.27 | 16.62 | A | A | A | 0.26 |
| 33 | 1750 | 80 | -15.17 | 6.42 | A | A | A | 1.74 |
| 34 | 1961.4 | 110.1 | -4.92 | 7.20 | A | A | A | 0.56 |
| 35 | 2470 | 140 | 19.73 | 7.24 | A | N | W | 2.26 |
| 36 | 1960 | 88 | -4.99 | 6.36 | A | A | A | 0.57 |

TABLE 22. Evaluation Results for Cs-137 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 38 | 2190 | 140 | 6.16 | 7.82 | A | A | A | 0.70 |
| 39 | 1999 | 47 | -3.10 | 5.08 | A | A | A | 0.35 |
| 41 | 2002.0 | 61.0 | -2.96 | 5.44 | A | A | A | 0.34 |
| 42 | 1219.2 | 1.85 | -40.90 | 4.51 | N | N | N | 4.68 |
| 43 | 2093.2 | 44.6 | 1.46 | 4.99 | A | A | A | 0.17 |
| 45 | 2175.0 | 175.0 | 5.43 | 9.22 | A | A | A | 0.62 |
| 46 | 3032.20 | 67.26 | 46.98 | 5.02 | N | N | N | 5.37 |
| 47 | 1750 | 14.8 | -15.17 | 4.59 | A | N | W | 1.74 |
| 48 | 1916 | 55 | -7.13 | 5.34 | A | A | A | 0.81 |
| 50 | 2236.88 | 33.25 | 8.43 | 4.75 | A | A | A | 0.96 |
| 51 | 1877 | 10 | -9.02 | 4.54 | A | A | A | 1.03 |
| 52 | 2170 | 83 | 5.19 | 5.91 | A | A | A | 0.59 |
| 53 | 2030 | 100 | -1.60 | 6.68 | A | A | A | 0.18 |
| 55 | 1843.1 | 40.9 | -10.66 | 5.02 | A | A | A | 1.22 |
| 56 | 1930 | 121.8 | -6.45 | 7.76 | A | A | A | 0.74 |
| 59 | 2060 | 143 | -0.15 | 8.28 | A | A | A | 0.02 |
| 60 | 1925.8 | 247.3 | -6.65 | 13.61 | A | A | A | 0.76 |
| 61 | 1683.93 | 126.47 | -18.37 | 8.76 | A | A | A | 2.10 |
| 63 | 2422.2 | 57.6 | 17.41 | 5.10 | A | N | W | 1.99 |
| 64 | 11531 | 95.2 | 458.94 | 4.58 | N | N | N | 52.49 |
| 66 | 2014.2 | 108 | -2.37 | 7.01 | A | A | A | 0.27 |
| 68 | 2007 | 120 | -2.71 | 7.49 | A | A | A | 0.31 |
| 70 | 1970 | 19 | -4.51 | 4.61 | A | A | A | 0.52 |
| 72 | 2080.46 | 140.24 | 0.85 | 8.11 | A | A | A | 0.10 |
| 74 | 2380.0 | 110 | 15.37 | 6.46 | A | A | A | 1.76 |
| 75 | 1340 | 40 | -35.05 | 5.41 | N | N | N | 4.01 |
| 76 | 2130.0 | 6.0 | 3.25 | 4.52 | A | A | A | 0.37 |
| 77 | 1663 | 4 | -19.39 | 4.51 | A | N | W | 2.22 |
| 78 | 1950 | 15 | -5.48 | 4.57 | A | A | A | 0.63 |
| 80 | 2014 | 103 | -2.38 | 6.82 | A | A | A | 0.27 |
| 81 | 2030 | 188 | -1.60 | 10.30 | A | A | A | 0.18 |
| 84 | 2052 | 123 | -0.53 | 7.50 | A | A | A | 0.06 |
| 85 | 1978.10 | 161.44 | -4.12 | 9.32 | A | A | A | 0.47 |
| 86 | 197 | 7.7 | -90.45 | 5.97 | N | N | N | 10.34 |
| 87 | 2350 | 90 | 13.91 | 5.92 | A | A | A | 1.59 |
| 88 | 1910.0 | 95.5 | -7.42 | 6.73 | A | A | A | 0.85 |
| 89 | 960 | 29 | -53.47 | 5.43 | N | N | N | 6.11 |
| 92 | 1820 | 190 | -11.78 | 11.37 | A | A | A | 1.35 |
| 93 | 1677.7 | 15.6 | -18.68 | 4.60 | A | N | W | 2.14 |
| 94 | 1917.32 | 9.83 | -7.06 | 4.54 | A | A | A | 0.81 |
| 95 | 1926.92 | 9.87 | -6.60 | 4.54 | A | A | A | 0.75 |
| 98 | 849 | 51 | -58.85 | 7.51 | N | N | N | 6.73 |
| 99 | 2250 | 132 | 9.06 | 7.40 | A | A | A | 1.04 |
| 101 | 1960 | 120 | -4.99 | 7.60 | A | A | A | 0.57 |
| 102 | 1960 | 60 | -4.99 | 5.45 | A | A | A | 0.57 |
| 103 | 2149.22 | 44.43 | 4.18 | 4.96 | A | A | A | 0.48 |
| 105 | 1797.8 | 33.5 | -12.86 | 4.88 | A | N | W | 1.47 |
| 106 | 2004.4 | 4.2 | -2.84 | 4.51 | A | A | A | 0.32 |
| 107 | 1672 | 33 | -18.95 | 4.92 | A | N | W | 2.17 |
| 108 | 1993.8990 | 4.8703 | -3.35 | 4.51 | A | A | A | 0.38 |
| 109 | 2371.6 | 110.6 | 14.96 | 6.49 | A | A | A | 1.71 |
| 110 | 1791.600 | 37.972 | -13.16 | 4.98 | A | N | W | 1.50 |
| 111 | 65.9 | 4.5 | -96.81 | 8.18 | N | N | N | 11.07 |
| 112 | 1724.3509 | 17.606 | -16.42 | 4.62 | A | N | W | 1.88 |

TABLE 22. Evaluation Results for Cs-137 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 113 | 1950.0 | 97.8 | -5.48 | 6.74 | A | A | A | 0.63 |
| 116 | 2404 | 75 | 16.53 | 5.48 | A | N | W | 1.89 |
| 120 | 1879 | 77 | -8.92 | 6.09 | A | A | A | 1.02 |
| 121 | 1970 | 65 | -4.51 | 5.59 | A | A | A | 0.52 |
| 122 | 1803 | 72 | -12.60 | 6.02 | A | A | A | 1.44 |
| 123 | 1920 | 140 | -6.93 | 8.57 | A | A | A | 0.79 |
| 124 | 1798 | 42 | -12.85 | 5.08 | A | A | A | 1.47 |
| 125 | 1990 | 70 | -3.54 | 5.72 | A | A | A | 0.40 |
| 126 | 1460 | 87.5 | -29.23 | 7.50 | N | N | N | 3.34 |
| 129 | 1865 | 40 | -9.60 | 4.99 | A | A | A | 1.10 |
| 130 | 1943 | 89 | -5.82 | 6.43 | A | A | A | 0.67 |
| 132 | 1770 | 100 | -14.20 | 7.23 | A | A | A | 1.62 |
| 133 | 1727 | 87 | -16.29 | 6.76 | A | A | A | 1.86 |
| 135 | 1982 | 100 | -3.93 | 6.77 | A | A | A | 0.45 |
| 137 | 1671.33 | 28.526 | -18.99 | 4.82 | A | N | W | 2.17 |
| 139 | 1844.3 | 7.3256 | -10.60 | 4.53 | A | A | A | 1.21 |
| 140 | 1993.6 | 57.5 | -3.36 | 5.35 | A | A | A | 0.38 |
| 141 | 1930 | 120 | -6.45 | 7.68 | A | A | A | 0.74 |
| 142 | 1339 | 37 | -35.09 | 5.29 | N | N | N | 4.01 |
| 143 | 1950 | 13 | -5.48 | 4.56 | A | A | A | 0.63 |
| 144 | 1994.67 | 108.5 | -3.31 | 7.06 | A | A | A | 0.38 |
| 146 | 2120 | 59 | 2.76 | 5.30 | A | A | A | 0.32 |
| 147 | 1954 | 172 | -5.28 | 9.89 | A | A | A | 0.60 |
| 148 | 1770.23 | 53.51 | -14.19 | 5.43 | A | N | W | 1.62 |
| 149 | 1769 | 59 | -14.25 | 5.61 | A | A | A | 1.63 |
| 150 | 1930 | 60 | -6.45 | 5.48 | A | A | A | 0.74 |
| 151 | 1860.0 | 50.0 | -9.84 | 5.25 | A | A | A | 1.13 |
| 152 | 2000 | 61 | -3.05 | 5.44 | A | A | A | 0.35 |
| 153 | 1877 | 93 | -9.02 | 6.70 | A | A | A | 1.03 |
| 154 | 1795 | 63.2 | -12.99 | 5.72 | A | A | A | 1.49 |
| 155 | 2032 | 73 | -1.50 | 5.76 | A | A | A | 0.17 |
| 156 | 2077 | 137 | 0.68 | 7.99 | A | A | A | 0.08 |
| 157 | 2093.03 | 109.58 | 1.46 | 6.91 | A | A | A | 0.17 |
| 160 | 1830 | 55 | -11.29 | 5.42 | A | A | A | 1.29 |
| 161 | 1752.8 | 48.3 | -15.04 | 5.28 | A | N | W | 1.72 |
| 162 | 1852.2 | 49.8 | -10.22 | 5.25 | A | A | A | 1.17 |
| 163 | 1640.0 | 98.5 | -20.50 | 7.51 | N | N | N | 2.34 |
| 164 | 2100 | 100 | 1.79 | 6.56 | A | A | A | 0.21 |
| 165 | 1783 | 99 | -13.57 | 7.15 | A | A | A | 1.55 |
| 166 | 2125.0 | 114.5 | 3.01 | 7.03 | A | A | A | 0.34 |
| 168 | 2013.46 | 82.39 | -2.40 | 6.09 | A | A | A | 0.27 |
| 169 | 1960.07 | 112.57 | -4.99 | 7.30 | A | A | A | 0.57 |
| 171 | 2014.80 | 180.00 | -2.34 | 10.01 | A | A | A | 0.27 |
| 172 | 2574.0 | 339.0 | 24.77 | 13.92 | N | A | N | 2.83 |
| 173 | 2460 | 101 | 19.24 | 6.10 | A | N | W | 2.20 |
| 176 | 91.7 | 5.42 | -95.56 | 7.43 | N | N | N | 10.93 |
| 180 | 1901.4 | 49.9 | -7.83 | 5.22 | A | A | A | 0.90 |
| 181 | 1996.49 | 44.69 | -3.22 | 5.03 | A | A | A | 0.37 |
| 182 | 1956.69 | 39.87 | -5.15 | 4.95 | A | A | A | 0.59 |
| 183 | 2092.2 | 43.42 | 1.42 | 4.96 | A | A | A | 0.16 |
| 184 | 2065.5 | 4.5 | 0.12 | 4.51 | A | A | A | 0.01 |
| 185 | 1977 | 138 | -4.17 | 8.31 | A | A | A | 0.48 |
| 186 | 2080 | 140 | 0.82 | 8.10 | A | A | A | 0.09 |
| 187 | 1398.2 | 3.6 | -32.22 | 4.52 | N | N | N | 3.69 |

TABLE 22. Evaluation Results for Cs-137 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 189 | 1852.8 | 111.8 | -10.19 | 7.53 | A | A | A | 1.17 |
| 190 | 1915 | 230 | -7.17 | 12.83 | A | A | A | 0.82 |
| 191 | 1871 | 49 | -9.31 | 5.21 | A | A | A | 1.06 |
| 192 | 2.200 | 0.13 | -99.89 | 7.43 | N | N | N | 11.42 |
| 193 | 2285 | 228 | 10.76 | 10.95 | A | A | A | 1.23 |
| 195 | 1930 | 80.0 | -6.45 | 6.12 | A | A | A | 0.74 |
| 196 | 2038 | 45 | -1.21 | 5.02 | A | A | A | 0.14 |
| 197 | 1993 | 54 | -3.39 | 5.26 | A | A | A | 0.39 |
| 198 | 1958.4 | 86.5 | -5.07 | 6.31 | A | A | A | 0.58 |
| 199 | 1529.14 | 30.58 | -25.88 | 4.93 | N | N | N | 2.96 |
| 200 | 1837.22 | 41.74 | -10.94 | 5.05 | A | A | A | 1.25 |
| 201 | 1941 | 194 | -5.91 | 10.96 | A | A | A | 0.68 |
| 202 | 1798 | 103 | -12.85 | 7.29 | A | A | A | 1.47 |
| 205 | 1860 | 45 | -9.84 | 5.12 | A | A | A | 1.13 |
| 206 | 2125.5 | 101.5 | 3.03 | 6.57 | A | A | A | 0.35 |
| 207 | 1840 | 117 | -10.81 | 7.79 | A | A | A | 1.24 |
| 208 | 1752 | 52.57 | -15.08 | 5.42 | A | N | W | 1.72 |
| 209 | 1610 | 183 | -21.96 | 12.23 | N | A | N | 2.51 |
| 210 | 1919 | 124 | -6.98 | 7.88 | A | A | A | 0.80 |
| 212 | 2150 | 70 | 4.22 | 5.56 | A | A | A | 0.48 |
| 213 | 1984 | 119 | -3.83 | 7.50 | A | A | A | 0.44 |
| 214 | 1903.7 | 114.2 | -7.72 | 7.50 | A | A | A | 0.88 |
| 215 | 182 | 110 | -91.18 | 60.61 | N | N | N | 10.43 |
| 216 | 1849 | 111 | -10.37 | 7.51 | A | A | A | 1.19 |
| 219 | 2009 | 96.0 | -2.62 | 6.57 | A | A | A | 0.30 |
| 221 | 1833.6 | 25.2 | -11.12 | 4.71 | A | A | A | 1.27 |
| 222 | 1867.2 | 52 | -9.49 | 5.30 | A | A | A | 1.09 |
| 224 | 2175 | 110 | 5.43 | 6.77 | A | A | A | 0.62 |
| 225 | 2290 | 95 | 11.00 | 6.13 | A | A | A | 1.26 |
| 227 | 2053.03 | 46.90 | -0.48 | 5.05 | A | A | A | 0.06 |
| 230 | 1590 | 70 | -22.93 | 6.30 | N | N | N | 2.62 |
| 231 | 1875 | 334.2 | -9.11 | 18.39 | A | A | A | 1.04 |
| 232 | 1896.594 | 84.3145 | -8.07 | 6.33 | A | A | A | 0.92 |
| 233 | 1948.47 | 58.58 | -5.55 | 5.42 | A | A | A | 0.63 |
| 234 | 1821.79 | 82.87 | -11.69 | 6.40 | A | A | A | 1.34 |
| 237 | 1887.0 | 113.0 | -8.53 | 7.50 | A | A | A | 0.98 |
| 238 | 1500.054 | 180.007 | -27.29 | 12.82 | N | A | N | 3.12 |
| 239 | 1724 | 58 | -16.43 | 5.62 | A | N | W | 1.88 |
| 240 | 1912 | 19 | -7.32 | 4.62 | A | A | A | 0.84 |
| 241 | 1963 | 117 | -4.85 | 7.47 | A | A | A | 0.55 |
| 242 | 1880 | 100 | -8.87 | 6.97 | A | A | A | 1.01 |
| 245 | 1886.5 | 51.9 | -8.56 | 5.28 | A | A | A | 0.98 |
| 246 | 1552.59 | 43.89 | -24.74 | 5.32 | N | N | N | 2.83 |
| 248 | 2117.10 | 63.51 | 2.62 | 5.41 | A | A | A | 0.30 |
| 249 | 2180 | 80 | 5.67 | 5.81 | A | A | A | 0.65 |
| 250 | 1943 | 73 | -5.82 | 5.87 | A | A | A | 0.67 |
| 251 | 2117 | 47 | 2.62 | 5.03 | A | A | A | 0.30 |
| 252 | 1886.59 | 14.4 | -8.55 | 4.57 | A | A | A | 0.98 |
| 253 | 1760 | 118 | -14.69 | 8.08 | A | A | A | 1.68 |
| 254 | 1920 | 130 | -6.93 | 8.13 | A | A | A | 0.79 |
| 255 | 1795.7 | 50.8 | -12.96 | 5.32 | A | A | A | 1.48 |
| 256 | 1695.31 | 101.91 | -17.82 | 7.51 | A | A | A | 2.04 |
| 257 | 1800 | 145 | -12.75 | 9.23 | A | A | A | 1.46 |
| 258 | 2120 | 81.1 | 2.76 | 5.91 | A | A | A | 0.32 |

TABLE 22. Evaluation Results for Cs-137 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 259 | 1750 | 3.45 | -15.17 | 4.51 | A | N | W | 1.74 |
| 260 | 1876.4 | 158.1 | -9.05 | 9.56 | A | A | A | 1.03 |
| 261 | 1990 | 99 | -3.54 | 6.71 | A | A | A | 0.40 |
| 262 | 2141 | 122 | 3.78 | 7.27 | A | A | A | 0.43 |
| 263 | 2.1 | 0.14 | -99.90 | 8.05 | N | N | N | 11.42 |
| 264 | 1892 | 6.3 | -8.29 | 4.52 | A | A | A | 0.95 |
| 266 | 1934.8 | 116.3 | -6.21 | 7.51 | A | A | A | 0.71 |
| 267 | 1788.38 | 17.88 | -13.31 | 4.62 | A | N | W | 1.52 |
| 268 | 2057 | 62 | -0.29 | 5.42 | A | A | A | 0.03 |
| 273 | 1872 | 13.1 | -9.26 | 4.56 | A | A | A | 1.06 |
| 274 | 2045 | 61 | -0.87 | 5.41 | A | A | A | 0.10 |
| 279 | 1970.0 | 118.2 | -4.51 | 7.50 | A | A | A | 0.52 |
| 281 | 1860 | 70 | -9.84 | 5.87 | A | A | A | 1.13 |
| 283 | 1820 | 81 | -11.78 | 6.33 | A | A | A | 1.35 |
| 287 | 2292 | 13 | 11.10 | 4.54 | A | A | A | 1.27 |
| 289 | 1827 | 95 | -11.44 | 6.88 | A | A | A | 1.31 |
| 290 | 2175.00 | 23.33 | 5.43 | 4.63 | A | A | A | 0.62 |
| 291 | 2237 | 81 | 8.43 | 5.78 | A | A | A | 0.96 |
| 295 | 1978.0 | 143.1 | -4.12 | 8.52 | A | A | A | 0.47 |
| 297 | 1570 | 94 | -23.90 | 7.49 | N | N | N | 2.73 |
| 299 | 2385.0 | 143.0 | 15.61 | 7.50 | A | A | A | 1.79 |
| 305 | 1215 | 104 | -41.11 | 9.67 | N | N | N | 4.70 |
| 307 | 1391 | 1.5 | -32.57 | 4.51 | N | N | N | 3.73 |
| 308 | 1632.03 | 98.09 | -20.89 | 7.51 | N | N | N | 2.39 |
| 310 | 1940 | 35 | -5.96 | 4.86 | A | A | A | 0.68 |
| 312 | 2072.63 | 19.85 | 0.47 | 4.61 | A | A | A | 0.05 |
| 313 | 1812 | 41 | -12.17 | 5.04 | A | A | A | 1.39 |
| 314 | 1905.54 | 257.87 | -7.63 | 14.26 | A | A | A | 0.87 |
| 316 | 2018.75 | 161.5 | -2.14 | 9.18 | A | A | A | 0.25 |
| 317 | 1990 |  | -3.54 |  | A |  |  | 0.40 |
| 322 | 2597 | 244 | 25.88 | 10.42 | N | A | N | 2.96 |
| 323 | 2127 | 38 | 3.10 | 4.85 | A | A | A | 0.35 |
| 324 | 1618.75 | 21.44 | -21.53 | 4.70 | N | N | N | 2.46 |
| 325 | 1899 | 112.8 | -7.95 | 7.46 | A | A | A | 0.91 |

## Reported Results for Cs-134 Determination in Sample 5

TABLE 23. Target values

| Parameter | Value |
| --- | --- |
| Target Value, Bq/kg | 19.05 |
| UNC, Bq/kg | 0.86 |
| MARB, % | 30 |

FIGURE 23. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

FIGURE 24. Reported results and evaluation parameters for Cs-134. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

TABLE 24. Evaluation Results for Cs-134 in Sample 5, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 2 | 18.7 | 0.5 | -1.84 | 5.25 | A | A | A | 0.30 |
| 3 | 18.5 | 1.1 | -2.89 | 7.47 | A | A | A | 0.47 |
| 4 | 17.5 | 0.9 | -8.14 | 6.84 | A | A | A | 1.31 |
| 5 | 17.7 | 1.49 | -7.09 | 9.55 | A | A | A | 1.14 |
| 6 | 14.9 | 0.5 | -21.78 | 5.63 | A | N | W | 3.52 |
| 7 | 17.94 | 0.42 | -5.83 | 5.09 | A | A | A | 0.94 |
| 8 | 15.74 | 0.35 | -17.38 | 5.03 | A | N | W | 2.81 |
| 9 | 20 | 3 | 4.99 | 15.66 | A | A | A | 0.81 |
| 10 | 19.05 | 1.46 | 0.00 | 8.89 | A | A | A | 0.00 |
| 11 | 19.1 | 0.5 | 0.26 | 5.22 | A | A | A | 0.04 |
| 12 | 21.3 | 1.5 | 11.81 | 8.37 | A | A | A | 1.91 |
| 13 | 18.14 | 0.67 | -4.78 | 5.83 | A | A | A | 0.77 |
| 14 | 19.1 | 1.1 | 0.26 | 7.32 | A | A | A | 0.04 |
| 15 | 20.3 | 1.0 | 6.56 | 6.68 | A | A | A | 1.06 |
| 16 | 20.1 | 0.4 | 5.51 | 4.93 | A | A | A | 0.89 |
| 17 | 18.43 | 1.02 | -3.25 | 7.14 | A | A | A | 0.53 |
| 18 | 19.9 | 1.6 | 4.46 | 9.22 | A | A | A | 0.72 |
| 19 | 18.9 | 1.9 | -0.79 | 11.02 | A | A | A | 0.13 |
| 20 | 19.5 | 0.6 | 2.36 | 5.46 | A | A | A | 0.38 |
| 21 | 16.1 | 1.2 | -15.49 | 8.71 | A | A | A | 2.50 |
| 22 | 19.2 | 0.5 | 0.79 | 5.21 | A | A | A | 0.13 |
| 25 | 19 | 1 | -0.26 | 6.93 | A | A | A | 0.04 |
| 26 | 19.91 | 1.09 | 4.51 | 7.10 | A | A | A | 0.73 |
| 27 | 19.0 | 2 | -0.26 | 11.45 | A | A | A | 0.04 |
| 29 | 44.06 | 0.66 | 131.29 | 4.76 | N | N | N | 21.19 |
| 31 | 17.00 | 0.64 | -10.76 | 5.88 | A | A | A | 1.74 |
| 32 | 19.6 | 2.2 | 2.89 | 12.10 | A | A | A | 0.47 |
| 34 | 18.4 | 1.6 | -3.41 | 9.80 | A | A | A | 0.55 |
| 35 | 19 | 1.5 | -0.26 | 9.09 | A | A | A | 0.04 |
| 36 | 18.7 | 0.9 | -1.84 | 6.60 | A | A | A | 0.30 |

TABLE 24. Evaluation Results for Cs-134 in Sample 5, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 38 | 18.2 | 1.0 | -4.46 | 7.11 | A | A | A | 0.72 |
| 39 | 15.9 | 0.6 | -16.54 | 5.88 | A | N | W | 2.67 |
| 41 | 19.1 | 0.8 | 0.26 | 6.16 | A | A | A | 0.04 |
| 42 | 4.32 | 0.37 | -77.32 | 9.68 | N | N | N | 12.48 |
| 43 | 19.80 | 0.80 | 3.94 | 6.06 | A | A | A | 0.64 |
| 45 | 19.2 | 0.9 | 0.79 | 6.51 | A | A | A | 0.13 |
| 46 | 38.79 | 1.52 | 103.62 | 5.98 | N | N | N | 16.73 |
| 48 | 19.5 | 0.7 | 2.36 | 5.77 | A | A | A | 0.38 |
| 50 | 17.60 | 1.28 | -7.61 | 8.56 | A | A | A | 1.23 |
| 51 | 17 | 1 | -10.76 | 7.41 | A | A | A | 1.74 |
| 52 | 16.14 | 0.93 | -15.28 | 7.32 | A | A | A | 2.47 |
| 53 | 19.9 | 1.2 | 4.46 | 7.53 | A | A | A | 0.72 |
| 55 | 18.4 | 0.6 | -3.41 | 5.57 | A | A | A | 0.55 |
| 56 | 17.64 | 0.72 | -7.40 | 6.09 | A | A | A | 1.19 |
| 59 | 18.9 | 1.4 | -0.79 | 8.67 | A | A | A | 0.13 |
| 61 | 17.60 | 0.49 | -7.61 | 5.30 | A | A | A | 1.23 |
| 63 | 18.0 | 0.5 | -5.51 | 5.30 | A | A | A | 0.89 |
| 64 | 19.0 | 0.5 | -0.26 | 5.23 | A | A | A | 0.04 |
| 66 | 20.2 | 0.6 | 6.04 | 5.40 | A | A | A | 0.97 |
| 68 | 19.2 | 1.5 | 0.79 | 9.02 | A | A | A | 0.13 |
| 69 | 17.96 | 0.83 | -5.72 | 6.46 | A | A | A | 0.92 |
| 70 | 19.38 | 0.24 | 1.73 | 4.68 | A | A | A | 0.28 |
| 72 | 18.644 | 0.950 | -2.13 | 6.81 | A | A | A | 0.34 |
| 74 | 18.8 | 0.8 | -1.31 | 6.20 | A | A | A | 0.21 |
| 75 | 17.1 | 0.4 | -10.24 | 5.08 | A | A | A | 1.65 |
| 76 | 16.00 | 0.24 | -16.01 | 4.76 | A | N | W | 2.58 |
| 77 | 24.8 | 0.4 | 30.18 | 4.79 | N | N | N | 4.87 |
| 78 | 17.8 | 0.24 | -6.56 | 4.71 | A | A | A | 1.06 |
| 80 | 18.97 | 1.61 | -0.42 | 9.61 | A | A | A | 0.07 |
| 81 | 19.9 | 1.5 | 4.46 | 8.79 | A | A | A | 0.72 |
| 84 | 22.36 | 0.33 | 17.38 | 4.75 | A | N | W | 2.81 |
| 85 | 19.38 | 1.62 | 1.73 | 9.50 | A | A | A | 0.28 |
| 86 | 18.8 | 0.8 | -1.31 | 6.20 | A | A | A | 0.21 |
| 87 | 25 | 1 | 31.23 | 6.03 | N | N | N | 5.04 |
| 88 | 18.4 | 2.0 | -3.41 | 11.77 | A | A | A | 0.55 |
| 92 | 18.3 | 1.5 | -3.94 | 9.36 | A | A | A | 0.64 |
| 93 | 16.57 | 0.25 | -13.02 | 4.76 | A | N | W | 2.10 |
| 96 | 20.0 | 1.2 | 4.99 | 7.51 | A | A | A | 0.81 |
| 98 | 17.7 | 0.6 | -7.09 | 5.65 | A | A | A | 1.14 |
| 99 | 19.5 | 0.8 | 2.36 | 6.10 | A | A | A | 0.38 |
| 101 | 19.34 | 0.57 | 1.52 | 5.39 | A | A | A | 0.25 |
| 102 | 19 | 1 | -0.26 | 6.93 | A | A | A | 0.04 |
| 103 | 19.86 | 2.25 | 4.25 | 12.20 | A | A | A | 0.69 |
| 105 | 18.33 | 0.28 | -3.78 | 4.77 | A | A | A | 0.61 |
| 107 | 19.0 | 0.2 | -0.26 | 4.64 | A | A | A | 0.04 |
| 108 | 18.6506 | 0.2838 | -2.10 | 4.76 | A | A | A | 0.34 |
| 109 | 18.9 | 0.8 | -0.79 | 6.19 | A | A | A | 0.13 |
| 110 | 18.131 | 0.359 | -4.82 | 4.93 | A | A | A | 0.78 |
| 111 | 18.9 | 1.3 | -0.79 | 8.23 | A | A | A | 0.13 |
| 112 | 13.6454 | 2.060 | -28.37 | 15.76 | A | A | A | 4.58 |
| 113 | 17.3 | 0.9 | -9.19 | 6.89 | A | A | A | 1.48 |
| 116 | 19.5 | 0.5 | 2.36 | 5.19 | A | A | A | 0.38 |
| 120 | 19.1 | 0.7 | 0.26 | 5.81 | A | A | A | 0.04 |
| 121 | 18.8 | 1.4 | -1.31 | 8.71 | A | A | A | 0.21 |

TABLE 24. Evaluation Results for Cs-134 in Sample 5, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 123 | 16.5 | 1.1 | -13.39 | 8.05 | A | A | A | 2.16 |
| 124 | 19.1 | 1.0 | 0.26 | 6.91 | A | A | A | 0.04 |
| 125 | 16.5 | 0.6 | -13.39 | 5.80 | A | A | A | 2.16 |
| 126 | 16.5 | 1 | -13.39 | 7.56 | A | A | A | 2.16 |
| 129 | 19.5 | 1.1 | 2.36 | 7.23 | A | A | A | 0.38 |
| 130 | 17.2 | 0.7 | -9.71 | 6.08 | A | A | A | 1.57 |
| 133 | 19.0 | 1.1 | -0.26 | 7.34 | A | A | A | 0.04 |
| 134 | 14.26 | 0.54 | -25.14 | 5.89 | A | N | W | 4.06 |
| 135 | 18.5 | 0.7 | -2.89 | 5.89 | A | A | A | 0.47 |
| 137 | 19.047 | 0.314 | -0.02 | 4.81 | A | A | A | 0.00 |
| 139 | 16.306 | 0.52603 | -14.40 | 5.55 | A | N | W | 2.33 |
| 140 | 18.5 | 0.9 | -2.89 | 6.64 | A | A | A | 0.47 |
| 141 | 19.3 | 1.2 | 1.31 | 7.68 | A | A | A | 0.21 |
| 144 | 18.6 | 2.5 | -2.36 | 14.18 | A | A | A | 0.38 |
| 146 | 18.7 | 0.9 | -1.84 | 6.60 | A | A | A | 0.30 |
| 147 | 18.5 | 1.7 | -2.89 | 10.24 | A | A | A | 0.47 |
| 148 | 18.52 | 0.595 | -2.78 | 5.54 | A | A | A | 0.45 |
| 149 | 18.9 | 0.4 | -0.79 | 4.99 | A | A | A | 0.13 |
| 150 | 19.2 | 1.0 | 0.79 | 6.89 | A | A | A | 0.13 |
| 152 | 18.4 | 0.8 | -3.41 | 6.27 | A | A | A | 0.55 |
| 153 | 18.7 | 1.3 | -1.84 | 8.29 | A | A | A | 0.30 |
| 154 | 18.2 | 0.8 | -4.46 | 6.30 | A | A | A | 0.72 |
| 156 | 18.6 | 1.2 | -2.36 | 7.87 | A | A | A | 0.38 |
| 157 | 16.03 | 0.92 | -15.85 | 7.30 | A | A | A | 2.56 |
| 158 | 16.4 | 2.1 | -13.91 | 13.58 | A | A | A | 2.25 |
| 160 | 17.2 | 0.5 | -9.71 | 5.37 | A | A | A | 1.57 |
| 161 | 18.3 | 0.6 | -3.94 | 5.58 | A | A | A | 0.64 |
| 162 | 18.9 | 0.5 | -0.79 | 5.23 | A | A | A | 0.13 |
| 163 | 19.2 | 0.8 | 0.79 | 6.14 | A | A | A | 0.13 |
| 164 | 19 | 1 | -0.26 | 6.93 | A | A | A | 0.04 |
| 165 | 19.7 | 1.9 | 3.41 | 10.65 | A | A | A | 0.55 |
| 166 | 19.5 | 1.1 | 2.36 | 7.23 | A | A | A | 0.38 |
| 168 | 15.54 | 0.95 | -18.43 | 7.60 | A | A | A | 2.97 |
| 169 | 18.85 | 0.35 | -1.05 | 4.88 | A | A | A | 0.17 |
| 171 | 19.07 | 2.10 | 0.10 | 11.90 | A | A | A | 0.02 |
| 172 | 17.0 | 2.2 | -10.76 | 13.71 | A | A | A | 1.74 |
| 173 | 18.5 | 2.0 | -2.89 | 11.72 | A | A | A | 0.47 |
| 176 | 17.6 | 1.17 | -7.61 | 8.04 | A | A | A | 1.23 |
| 177 | 16.1 | 1.6 | -15.49 | 10.92 | A | A | A | 2.50 |
| 179 | 19.5 | 0.3 | 2.36 | 4.77 | A | A | A | 0.38 |
| 180 | 18.66 | 1.0 | -2.05 | 7.01 | A | A | A | 0.33 |
| 181 | 16.192 | 0.337 | -15.00 | 4.97 | A | N | W | 2.42 |
| 182 | 18.84 | 1.60 | -1.10 | 9.62 | A | A | A | 0.18 |
| 183 | 18.88 | 0.9 | -0.89 | 6.57 | A | A | A | 0.14 |
| 184 | 16.76 | 0.32 | -12.02 | 4.90 | A | A | A | 1.94 |
| 185 | 18.5 | 1.5 | -2.89 | 9.28 | A | A | A | 0.47 |
| 186 | 20.1 | 1.3 | 5.51 | 7.89 | A | A | A | 0.89 |
| 189 | 18.8 | 1.1 | -1.31 | 7.39 | A | A | A | 0.21 |
| 190 | 18.1 | 2.2 | -4.99 | 12.97 | A | A | A | 0.81 |
| 191 | 19.2 | 1.0 | 0.79 | 6.89 | A | A | A | 0.13 |
| 193 | 19.22 | 1.92 | 0.89 | 10.96 | A | A | A | 0.14 |
| 195 | 17.9 | 0.50 | -6.04 | 5.31 | A | A | A | 0.97 |
| 196 | 19.1 | 1.2 | 0.26 | 7.74 | A | A | A | 0.04 |
| 197 | 18.5 | 0.8 | -2.89 | 6.25 | A | A | A | 0.47 |

TABLE 24. Evaluation Results for Cs-134 in Sample 5, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 198 | 18.58 | 1.07 | -2.47 | 7.32 | A | A | A | 0.40 |
| 199 | 16.34 | 0.36 | -14.23 | 5.02 | A | N | W | 2.30 |
| 200 | 19.67 | 0.43 | 3.25 | 5.02 | A | A | A | 0.53 |
| 201 | 18.3 | 1.8 | -3.94 | 10.82 | A | A | A | 0.64 |
| 202 | 18.42 | 1.15 | -3.31 | 7.70 | A | A | A | 0.53 |
| 205 | 18.9 | 1.2 | -0.79 | 7.79 | A | A | A | 0.13 |
| 206 | 18.8 | 1.0 | -1.31 | 6.98 | A | A | A | 0.21 |
| 207 | 18.7 | 1.1 | -1.84 | 7.41 | A | A | A | 0.30 |
| 208 | 12.96 | 0.78 | -31.97 | 7.52 | N | N | N | 5.16 |
| 209 | 18.1 | 1.4 | -4.99 | 8.96 | A | A | A | 0.81 |
| 210 | 19.3 | 1.4 | 1.31 | 8.54 | A | A | A | 0.21 |
| 211 | 19.5 | 1.7 | 2.36 | 9.82 | A | A | A | 0.38 |
| 212 | 19.3 | 0.4 | 1.31 | 4.97 | A | A | A | 0.21 |
| 213 | 19.0 | 1.2 | -0.26 | 7.76 | A | A | A | 0.04 |
| 214 | 16.4 | 1.1 | -13.91 | 8.09 | A | A | A | 2.25 |
| 215 | 17.3 | 1.79 | -9.19 | 11.29 | A | A | A | 1.48 |
| 216 | 18.04 | 0.88 | -5.30 | 6.65 | A | A | A | 0.86 |
| 219 | 18.3 | 1.8 | -3.94 | 10.82 | A | A | A | 0.64 |
| 221 | 18.72 | 0.42 | -1.73 | 5.04 | A | A | A | 0.28 |
| 222 | 12.2 | 0.423 | -35.96 | 5.69 | N | N | N | 5.81 |
| 223 | 17.60 | 0.60 | -7.61 | 5.66 | A | A | A | 1.23 |
| 224 | 18.8 | 0.8 | -1.31 | 6.20 | A | A | A | 0.21 |
| 225 | 18.4 | 0.8 | -3.41 | 6.27 | A | A | A | 0.55 |
| 227 | 18.54 | 0.50 | -2.68 | 5.26 | A | A | A | 0.43 |
| 228 | 18.3 | 0.7 | -3.94 | 5.92 | A | A | A | 0.64 |
| 229 | 19.1 | 0.809 | 0.26 | 6.19 | A | A | A | 0.04 |
| 230 | 19.65 | 0.8 | 3.15 | 6.08 | A | A | A | 0.51 |
| 231 | 17.99 | 2.943 | -5.56 | 16.97 | A | A | A | 0.90 |
| 232 | 18.1849 | 0.7926 | -4.54 | 6.28 | A | A | A | 0.73 |
| 233 | 73.74 | 0.18 | 287.09 | 4.52 | N | N | N | 46.35 |
| 234 | 19.80 | 0.70 | 3.94 | 5.73 | A | A | A | 0.64 |
| 235 | 18.57 | 5.45 | -2.52 | 29.69 | A | A | A | 0.41 |
| 237 | 18.8 | 0.8 | -1.31 | 6.20 | A | A | A | 0.21 |
| 238 | 19.4 | 2.52 | 1.84 | 13.75 | A | A | A | 0.30 |
| 239 | 18.1 | 0.7 | -4.99 | 5.94 | A | A | A | 0.81 |
| 240 | 19.2 | 0.6 | 0.79 | 5.49 | A | A | A | 0.13 |
| 241 | 19.3 | 0.6 | 1.31 | 5.48 | A | A | A | 0.21 |
| 242 | 18.7 | 1.0 | -1.84 | 7.00 | A | A | A | 0.30 |
| 243 | 19.4 | 1.6 | 1.84 | 9.40 | A | A | A | 0.30 |
| 246 | 15.34 | 0.248 | -19.48 | 4.80 | A | N | W | 3.14 |
| 248 | 16.84 | 0.42 | -11.60 | 5.16 | A | A | A | 1.87 |
| 249 | 18.5 | 0.8 | -2.89 | 6.25 | A | A | A | 0.47 |
| 250 | 16.8 | 0.5 | -11.81 | 5.41 | A | A | A | 1.91 |
| 251 | 17.6 | 0.3 | -7.61 | 4.83 | A | A | A | 1.23 |
| 252 | 18.35 | 1.4 | -3.67 | 8.87 | A | A | A | 0.59 |
| 254 | 19.69 | 0.88 | 3.36 | 6.35 | A | A | A | 0.54 |
| 255 | 19.1 | 0.6 | 0.26 | 5.50 | A | A | A | 0.04 |
| 256 | 17.43 | 0.13 | -8.50 | 4.58 | A | A | A | 1.37 |
| 257 | 18.7 | 0.8 | -1.84 | 6.22 | A | A | A | 0.30 |
| 260 | 21.5 | 1.1 | 12.86 | 6.82 | A | A | A | 2.08 |
| 261 | 20.0 | 1.2 | 4.99 | 7.51 | A | A | A | 0.81 |
| 262 | 19.7 | 1.2 | 3.41 | 7.58 | A | A | A | 0.55 |
| 263 | 15.86 | 0.54 | -16.75 | 5.65 | A | N | W | 2.70 |
| 266 | 16.0 | 0.6 | -16.01 | 5.87 | A | N | W | 2.58 |

TABLE 24. Evaluation Results for Cs-134 in Sample 5, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 267 | 30.42 | 2.03 | 59.69 | 8.06 | N | N | N | 9.64 |
| 268 | 18.57 | 0.66 | -2.52 | 5.75 | A | A | A | 0.41 |
| 270 | 18.1 | 1.99 | -4.99 | 11.89 | A | A | A | 0.81 |
| 273 | 17.68 | 0.976 | -7.19 | 7.13 | A | A | A | 1.16 |
| 274 | 17.8 | 1.0 | -6.56 | 7.21 | A | A | A | 1.06 |
| 275 | 19.7 | 1.24 | 3.41 | 7.75 | A | A | A | 0.55 |
| 279 | 18.8 | 0.8 | -1.31 | 6.20 | A | A | A | 0.21 |
| 283 | 18.9 | 0.65 | -0.79 | 5.68 | A | A | A | 0.13 |
| 286 | 18.9 | 0.8 | -0.79 | 6.19 | A | A | A | 0.13 |
| 287 | 19.16 | 0.4 | 0.58 | 4.97 | A | A | A | 0.09 |
| 289 | 19.4 | 1.1 | 1.84 | 7.25 | A | A | A | 0.30 |
| 290 | 10.41 | 0.38 | -45.35 | 5.81 | N | N | N | 7.32 |
| 291 | 18.9 | 0.8 | -0.79 | 6.19 | A | A | A | 0.13 |
| 294 | 21.1 | 0.4 | 10.76 | 4.90 | A | A | A | 1.74 |
| 295 | 15.0 | 0.9 | -21.26 | 7.51 | A | N | W | 3.43 |
| 297 | 17.5 | 1.1 | -8.14 | 7.74 | A | A | A | 1.31 |
| 299 | 21.2 | 1.3 | 11.29 | 7.61 | A | A | A | 1.82 |
| 305 | 16.1 | 0.66 | -15.49 | 6.10 | A | A | A | 2.50 |
| 307 | 14.9 | 1 | -21.78 | 8.09 | A | N | W | 3.52 |
| 308 | 12.04 | 0.67 | -36.80 | 7.17 | N | N | N | 5.94 |
| 310 | 17.2 | 0.7 | -9.71 | 6.08 | A | A | A | 1.57 |
| 312 | 18.11 | 0.46 | -4.93 | 5.18 | A | A | A | 0.80 |
| 313 | 18.96 | 0.47 | -0.47 | 5.15 | A | A | A | 0.08 |
| 314 | 18.67 | 0.87 | -1.99 | 6.49 | A | A | A | 0.32 |
| 316 | 14.032 | 1.122 | -26.34 | 9.18 | A | N | W | 4.25 |
| 317 | 17.6 |  | -7.61 |  | A |  |  | 1.23 |
| 320 | 19.86 | 0.64 | 4.25 | 5.55 | A | A | A | 0.69 |
| 321 | 15.34 | 0.33 | -19.48 | 5.00 | A | N | W | 3.14 |
| 322 | 17.9 | 1.3 | -6.04 | 8.55 | A | A | A | 0.97 |
| 323 | 19.01 | 0.38 | -0.21 | 4.94 | A | A | A | 0.03 |
| 324 | 17.82 | 0.29 | -6.46 | 4.80 | A | A | A | 1.04 |
| 325 | 18.5 | 1.5 | -2.89 | 9.28 | A | A | A | 0.47 |

## Reported Results for Cs-137 Determination in Sample 5

TABLE 25. Target values

| Parameter | Value |
| --- | --- |
| Target Value, Bq/kg | 26.02 |
| UNC, Bq/kg | 1.17 |
| MARB, % | 25 |

FIGURE 25. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

FIGURE 26. Reported results and evaluation parameters for Cs-137. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

TABLE 26. Evaluation Results for Cs-137 in Sample 5, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 2 | 25.5 | 0.7 | -2.00 | 5.27 | A | A | A | 0.40 |
| 3 | 25.6 | 1.4 | -1.61 | 7.08 | A | A | A | 0.33 |
| 4 | 23.7 | 1.2 | -8.92 | 6.77 | A | A | A | 1.80 |
| 5 | 23.9 | 2 | -8.15 | 9.50 | A | A | A | 1.64 |
| 6 | 22.2 | 0.9 | -14.68 | 6.05 | A | A | A | 2.96 |
| 7 | 26.70 | 0.94 | 2.61 | 5.71 | A | A | A | 0.53 |
| 8 | 22.90 | 0.52 | -11.99 | 5.04 | A | A | A | 2.42 |
| 9 | 27 | 3 | 3.77 | 11.99 | A | A | A | 0.76 |
| 10 | 23.73 | 0.95 | -8.80 | 6.02 | A | A | A | 1.78 |
| 11 | 26.0 | 0.8 | -0.08 | 5.45 | A | A | A | 0.02 |
| 12 | 32.7 | 0.8 | 25.67 | 5.12 | N | N | N | 5.18 |
| 13 | 25.70 | 0.97 | -1.23 | 5.87 | A | A | A | 0.25 |
| 14 | 26.6 | 1.5 | 2.23 | 7.21 | A | A | A | 0.45 |
| 15 | 27.2 | 1.4 | 4.53 | 6.83 | A | A | A | 0.91 |
| 16 | 26.3 | 0.7 | 1.08 | 5.23 | A | A | A | 0.22 |
| 17 | 25.98 | 0.89 | -0.15 | 5.65 | A | A | A | 0.03 |
| 18 | 27.0 | 2.3 | 3.77 | 9.63 | A | A | A | 0.76 |
| 19 | 24.7 | 1.1 | -5.07 | 6.33 | A | A | A | 1.02 |
| 20 | 26.8 | 0.6 | 3.00 | 5.02 | A | A | A | 0.60 |
| 21 | 26.4 | 1.9 | 1.46 | 8.49 | A | A | A | 0.29 |
| 22 | 27.1 | 0.6 | 4.15 | 5.01 | A | A | A | 0.84 |
| 25 | 26 | 1 | -0.08 | 5.92 | A | A | A | 0.02 |
| 26 | 28.83 | 1.57 | 10.80 | 7.06 | A | A | A | 2.18 |
| 27 | 26.79 | 0.95 | 2.96 | 5.73 | A | A | A | 0.60 |
| 29 | 70.14 | 1.03 | 169.56 | 4.73 | N | N | N | 34.20 |
| 31 | 26.2 | 1.00 | 0.69 | 5.90 | A | A | A | 0.14 |
| 32 | 26.5 | 2.9 | 1.84 | 11.83 | A | A | A | 0.37 |
| 34 | 25.7 | 1.8 | -1.23 | 8.32 | A | A | A | 0.25 |
| 35 | 25 | 1.7 | -3.92 | 8.15 | A | A | A | 0.79 |
| 36 | 25.8 | 1.1 | -0.85 | 6.20 | A | A | A | 0.17 |

TABLE 26. Evaluation Results for Cs-137 in Sample 5, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 38 | 24.7 | 1.5 | -5.07 | 7.56 | A | A | A | 1.02 |
| 39 | 25.6 | 0.6 | -1.61 | 5.07 | A | A | A | 0.33 |
| 41 | 25.7 | 1.2 | -1.23 | 6.48 | A | A | A | 0.25 |
| 42 | 26.94 | 0.57 | 3.54 | 4.97 | A | A | A | 0.71 |
| 43 | 25.83 | 0.90 | -0.73 | 5.69 | A | A | A | 0.15 |
| 45 | 25.5 | 1.7 | -2.00 | 8.04 | A | A | A | 0.40 |
| 46 | 52.33 | 1.93 | 101.11 | 5.82 | N | N | N | 20.40 |
| 48 | 27.8 | 1.1 | 6.84 | 5.99 | A | A | A | 1.38 |
| 50 | 26.88 | 1.68 | 3.31 | 7.70 | A | A | A | 0.67 |
| 51 | 24 | 2 | -7.76 | 9.47 | A | A | A | 1.57 |
| 52 | 27.0 | 1.0 | 3.77 | 5.83 | A | A | A | 0.76 |
| 53 | 25.6 | 1.5 | -1.61 | 7.39 | A | A | A | 0.33 |
| 55 | 24.9 | 0.8 | -4.30 | 5.53 | A | A | A | 0.87 |
| 56 | 26.85 | 0.61 | 3.19 | 5.04 | A | A | A | 0.64 |
| 59 | 26.0 | 1.9 | -0.08 | 8.58 | A | A | A | 0.02 |
| 60 | 20.4 | 6 | -21.60 | 29.75 | A | N | W | 4.36 |
| 61 | 25.83 | 0.83 | -0.73 | 5.53 | A | A | A | 0.15 |
| 63 | 26.0 | 0.9 | -0.08 | 5.67 | A | A | A | 0.02 |
| 64 | 26.0 | 2 | -0.08 | 8.91 | A | A | A | 0.02 |
| 66 | 28 | 1.7 | 7.61 | 7.56 | A | A | A | 1.53 |
| 68 | 25.7 | 1.5 | -1.23 | 7.37 | A | A | A | 0.25 |
| 69 | 26.47 | 0.98 | 1.73 | 5.82 | A | A | A | 0.35 |
| 70 | 25.99 | 0.40 | -0.12 | 4.75 | A | A | A | 0.02 |
| 72 | 25.691 | 1.681 | -1.26 | 7.94 | A | A | A | 0.26 |
| 74 | 26.5 | 1.3 | 1.84 | 6.65 | A | A | A | 0.37 |
| 75 | 27.3 | 1.02 | 4.92 | 5.85 | A | A | A | 0.99 |
| 76 | 26.20 | 0.23 | 0.69 | 4.58 | A | A | A | 0.14 |
| 77 | 31.5 | 0.3 | 21.06 | 4.60 | A | N | W | 4.25 |
| 78 | 26 | 0.3 | -0.08 | 4.64 | A | A | A | 0.02 |
| 80 | 28.10 | 0.96 | 7.99 | 5.65 | A | A | A | 1.61 |
| 81 | 27.0 | 2.1 | 3.77 | 8.98 | A | A | A | 0.76 |
| 84 | 30.96 | 0.52 | 18.99 | 4.80 | A | N | W | 3.83 |
| 85 | 25.94 | 2.21 | -0.31 | 9.63 | A | A | A | 0.06 |
| 86 | 24.9 | 1.1 | -4.30 | 6.30 | A | A | A | 0.87 |
| 87 | 18.4 | 0.7 | -29.29 | 5.89 | N | N | N | 5.91 |
| 88 | 26.3 | 2.5 | 1.08 | 10.52 | A | A | A | 0.22 |
| 89 | 25.9 | 1.0 | -0.46 | 5.93 | A | A | A | 0.09 |
| 92 | 27.2 | 3.0 | 4.53 | 11.91 | A | A | A | 0.91 |
| 93 | 26.14 | 0.55 | 0.46 | 4.96 | A | A | A | 0.09 |
| 96 | 26.9 | 1.6 | 3.38 | 7.46 | A | A | A | 0.68 |
| 98 | 26.0 | 1.4 | -0.08 | 7.02 | A | A | A | 0.02 |
| 99 | 25.9 | 0.9 | -0.46 | 5.68 | A | A | A | 0.09 |
| 101 | 26.69 | 0.74 | 2.57 | 5.28 | A | A | A | 0.52 |
| 102 | 26.6 | 0.8 | 2.23 | 5.41 | A | A | A | 0.45 |
| 103 | 25.80 | 0.55 | -0.85 | 4.98 | A | A | A | 0.17 |
| 105 | 28.5 | 0.4 | 9.53 | 4.71 | A | A | A | 1.92 |
| 107 | 30.6 | 0.5 | 17.60 | 4.78 | A | N | W | 3.55 |
| 108 | 27.7910 | 0.3639 | 6.81 | 4.68 | A | A | A | 1.37 |
| 109 | 26.5 | 1.2 | 1.84 | 6.38 | A | A | A | 0.37 |
| 110 | 28.517 | 0.684 | 9.60 | 5.10 | A | A | A | 1.94 |
| 111 | 25.5 | 1.7 | -2.00 | 8.04 | A | A | A | 0.40 |
| 112 | 20.7639 | 2.818 | -20.20 | 14.30 | A | A | A | 4.07 |
| 113 | 27.0 | 1.4 | 3.77 | 6.86 | A | A | A | 0.76 |
| 116 | 26.4 | 0.8 | 1.46 | 5.42 | A | A | A | 0.29 |

TABLE 26. Evaluation Results for Cs-137 in Sample 5, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 120 | 25.9 | 0.8 | -0.46 | 5.46 | A | A | A | 0.09 |
| 121 | 26.5 | 1.8 | 1.84 | 8.15 | A | A | A | 0.37 |
| 123 | 22.2 | 1.5 | -14.68 | 8.12 | A | A | A | 2.96 |
| 124 | 26.1 | 0.9 | 0.31 | 5.67 | A | A | A | 0.06 |
| 125 | 26.9 | 1.2 | 3.38 | 6.33 | A | A | A | 0.68 |
| 126 | 23.7 | 1.4 | -8.92 | 7.42 | A | A | A | 1.80 |
| 129 | 25.8 | 1.0 | -0.85 | 5.94 | A | A | A | 0.17 |
| 130 | 25.3 | 0.9 | -2.77 | 5.73 | A | A | A | 0.56 |
| 133 | 25.3 | 1.3 | -2.77 | 6.83 | A | A | A | 0.56 |
| 134 | 18.98 | 1.17 | -27.06 | 7.63 | N | N | N | 5.46 |
| 135 | 25.3 | 0.8 | -2.77 | 5.50 | A | A | A | 0.56 |
| 137 | 25.1 | 0.602 | -3.54 | 5.10 | A | A | A | 0.71 |
| 139 | 27.954 | 0.67057 | 7.43 | 5.10 | A | A | A | 1.50 |
| 140 | 26.1 | 1.0 | 0.31 | 5.91 | A | A | A | 0.06 |
| 141 | 27.0 | 1.6 | 3.77 | 7.44 | A | A | A | 0.76 |
| 144 | 26.5 | 3.4 | 1.84 | 13.60 | A | A | A | 0.37 |
| 146 | 25.5 | 1.6 | -2.00 | 7.72 | A | A | A | 0.40 |
| 147 | 25.6 | 2.3 | -1.61 | 10.05 | A | A | A | 0.33 |
| 148 | 27.56 | 0.932 | 5.92 | 5.63 | A | A | A | 1.19 |
| 149 | 26.0 | 0.6 | -0.08 | 5.05 | A | A | A | 0.02 |
| 150 | 26.1 | 0.9 | 0.31 | 5.67 | A | A | A | 0.06 |
| 152 | 25.8 | 1.0 | -0.85 | 5.94 | A | A | A | 0.17 |
| 153 | 25.3 | 1.8 | -2.77 | 8.42 | A | A | A | 0.56 |
| 154 | 27.4 | 1.3 | 5.30 | 6.54 | A | A | A | 1.07 |
| 156 | 25.6 | 1.6 | -1.61 | 7.70 | A | A | A | 0.33 |
| 157 | 25.64 | 1.52 | -1.46 | 7.44 | A | A | A | 0.29 |
| 158 | 30.5 | 2.7 | 17.22 | 9.93 | A | A | A | 3.47 |
| 160 | 25.5 | 1.5 | -2.00 | 7.40 | A | A | A | 0.40 |
| 161 | 25.6 | 0.8 | -1.61 | 5.48 | A | A | A | 0.33 |
| 162 | 25.9 | 0.8 | -0.46 | 5.46 | A | A | A | 0.09 |
| 163 | 26.5 | 1.6 | 1.84 | 7.53 | A | A | A | 0.37 |
| 164 | 26.8 | 1.4 | 3.00 | 6.89 | A | A | A | 0.60 |
| 165 | 29.5 | 1.8 | 13.37 | 7.58 | A | A | A | 2.70 |
| 166 | 26.8 | 1.5 | 3.00 | 7.18 | A | A | A | 0.60 |
| 168 | 25.23 | 1.74 | -3.04 | 8.23 | A | A | A | 0.61 |
| 169 | 25.93 | 0.39 | -0.35 | 4.74 | A | A | A | 0.07 |
| 171 | 26.49 | 2.93 | 1.81 | 11.94 | A | A | A | 0.36 |
| 172 | 28.6 | 4.0 | 9.92 | 14.69 | A | A | A | 2.00 |
| 173 | 29.3 | 3.1 | 12.61 | 11.50 | A | A | A | 2.54 |
| 176 | 23.6 | 1.45 | -9.30 | 7.61 | A | A | A | 1.88 |
| 177 | 28.9 | 2.8 | 11.07 | 10.68 | A | A | A | 2.23 |
| 179 | 26.5 | 0.7 | 1.84 | 5.22 | A | A | A | 0.37 |
| 180 | 26.23 | 0.85 | 0.81 | 5.54 | A | A | A | 0.16 |
| 181 | 25.282 | 0.601 | -2.84 | 5.09 | A | A | A | 0.57 |
| 182 | 28.60 | 2.25 | 9.92 | 9.06 | A | A | A | 2.00 |
| 183 | 25.93 | 1.0 | -0.35 | 5.92 | A | A | A | 0.07 |
| 184 | 25.47 | 0.35 | -2.11 | 4.70 | A | A | A | 0.43 |
| 185 | 26.6 | 2.1 | 2.23 | 9.09 | A | A | A | 0.45 |
| 186 | 25.3 | 1.7 | -2.77 | 8.09 | A | A | A | 0.56 |
| 187 | 1112 | 25 | 4173.64 | 5.03 | N | N | N | 841.84 |
| 189 | 25.8 | 1.8 | -0.85 | 8.30 | A | A | A | 0.17 |
| 190 | 26.5 | 2.8 | 1.84 | 11.48 | A | A | A | 0.37 |
| 191 | 26.1 | 1.0 | 0.31 | 5.91 | A | A | A | 0.06 |
| 192 | 0.016 | 0.001 | -99.94 | 7.70 | N | N | N | 20.16 |

TABLE 26. Evaluation Results for Cs-137 in Sample 5, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 193 | 26.37 | 2.64 | 1.35 | 10.97 | A | A | A | 0.27 |
| 195 | 27.3 | 1.30 | 4.92 | 6.55 | A | A | A | 0.99 |
| 196 | 28.8 | 1.7 | 10.68 | 7.42 | A | A | A | 2.16 |
| 197 | 25.6 | 1.0 | -1.61 | 5.96 | A | A | A | 0.33 |
| 198 | 25.69 | 1.16 | -1.27 | 6.37 | A | A | A | 0.26 |
| 199 | 24.56 | 0.51 | -5.61 | 4.95 | A | A | A | 1.13 |
| 200 | 26.51 | 1.04 | 1.88 | 5.97 | A | A | A | 0.38 |
| 201 | 26.2 | 2.6 | 0.69 | 10.89 | A | A | A | 0.14 |
| 202 | 25.77 | 1.58 | -0.96 | 7.60 | A | A | A | 0.19 |
| 205 | 26.2 | 1.1 | 0.69 | 6.15 | A | A | A | 0.14 |
| 206 | 25.6 | 1.5 | -1.61 | 7.39 | A | A | A | 0.33 |
| 207 | 25.2 | 1.6 | -3.15 | 7.78 | A | A | A | 0.64 |
| 208 | 22.72 | 2.86 | -12.68 | 13.37 | A | A | A | 2.56 |
| 209 | 26.9 | 2.1 | 3.38 | 9.01 | A | A | A | 0.68 |
| 210 | 27.5 | 1.9 | 5.69 | 8.24 | A | A | A | 1.15 |
| 211 | 26.2 | 1.9 | 0.69 | 8.53 | A | A | A | 0.14 |
| 212 | 26.2 | 0.5 | 0.69 | 4.88 | A | A | A | 0.14 |
| 213 | 26.1 | 1.6 | 0.31 | 7.60 | A | A | A | 0.06 |
| 214 | 23.2 | 1.6 | -10.84 | 8.23 | A | A | A | 2.19 |
| 215 | 23.6 | 1.58 | -9.30 | 8.06 | A | A | A | 1.88 |
| 216 | 25.62 | 1.61 | -1.54 | 7.73 | A | A | A | 0.31 |
| 219 | 25.7 | 2.3 | -1.23 | 10.02 | A | A | A | 0.25 |
| 221 | 28.55 | 1.38 | 9.72 | 6.60 | A | A | A | 1.96 |
| 222 | 20 | 1.03 | -23.14 | 6.84 | A | N | W | 4.67 |
| 223 | 24.84 | 0.55 | -4.53 | 5.01 | A | A | A | 0.91 |
| 224 | 26.4 | 0.9 | 1.46 | 5.64 | A | A | A | 0.29 |
| 225 | 25.3 | 1.1 | -2.77 | 6.25 | A | A | A | 0.56 |
| 227 | 25.78 | 0.76 | -0.92 | 5.38 | A | A | A | 0.19 |
| 228 | 25.8 | 1.0 | -0.85 | 5.94 | A | A | A | 0.17 |
| 229 | 26.7 | 0.924 | 2.61 | 5.67 | A | A | A | 0.53 |
| 230 | 26.45 | 1.0 | 1.65 | 5.87 | A | A | A | 0.33 |
| 231 | 27.77 | 5.010 | 6.73 | 18.59 | A | A | A | 1.36 |
| 232 | 25.204 | 1.1542 | -3.14 | 6.42 | A | A | A | 0.63 |
| 233 | 25.15 | 0.32 | -3.34 | 4.67 | A | A | A | 0.67 |
| 234 | 26.45 | 1.7 | 1.65 | 7.84 | A | A | A | 0.33 |
| 235 | 34.31 | 1.25 | 31.86 | 5.79 | N | N | N | 6.43 |
| 237 | 25.7 | 1.6 | -1.23 | 7.68 | A | A | A | 0.25 |
| 238 | 26.5 | 3.45 | 1.84 | 13.77 | A | A | A | 0.37 |
| 239 | 25.2 | 0.8 | -3.15 | 5.50 | A | A | A | 0.64 |
| 240 | 27.0 | 0.9 | 3.77 | 5.60 | A | A | A | 0.76 |
| 241 | 26.4 | 1.6 | 1.46 | 7.55 | A | A | A | 0.29 |
| 242 | 25.6 | 1.3 | -1.61 | 6.78 | A | A | A | 0.33 |
| 243 | 27.1 | 2.3 | 4.15 | 9.60 | A | A | A | 0.84 |
| 246 | 22.28 | 0.646 | -14.37 | 5.35 | A | N | W | 2.90 |
| 248 | 25.92 | 0.78 | -0.38 | 5.41 | A | A | A | 0.08 |
| 249 | 26.9 | 1.1 | 3.38 | 6.08 | A | A | A | 0.68 |
| 250 | 25.4 | 1.0 | -2.38 | 5.98 | A | A | A | 0.48 |
| 251 | 27.1 | 0.5 | 4.15 | 4.86 | A | A | A | 0.84 |
| 252 | 28.56 | 1.96 | 9.76 | 8.20 | A | A | A | 1.97 |
| 254 | 26.9 | 1.2 | 3.38 | 6.33 | A | A | A | 0.68 |
| 255 | 27.4 | 0.8 | 5.30 | 5.36 | A | A | A | 1.07 |
| 256 | 24.71 | 0.22 | -5.03 | 4.58 | A | A | A | 1.02 |
| 257 | 28.2 | 1.0 | 8.38 | 5.73 | A | A | A | 1.69 |
| 260 | 27.2 | 1.4 | 4.53 | 6.83 | A | A | A | 0.91 |

TABLE 26. Evaluation Results for Cs-137 in Sample 5, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 261 | 27.3 | 1.5 | 4.92 | 7.10 | A | A | A | 0.99 |
| 262 | 25.8 | 1.5 | -0.85 | 7.35 | A | A | A | 0.17 |
| 263 | 25.7 | 0.78 | -1.23 | 5.42 | A | A | A | 0.25 |
| 264 | 30.7 | 1.7 | 17.99 | 7.13 | A | A | A | 3.63 |
| 266 | 26.0 | 1.6 | -0.08 | 7.62 | A | A | A | 0.02 |
| 267 | 48.34 | 3.13 | 85.78 | 7.88 | N | N | N | 17.30 |
| 268 | 24.91 | 0.63 | -4.27 | 5.16 | A | A | A | 0.86 |
| 270 | 26.5 | 2.87 | 1.84 | 11.73 | A | A | A | 0.37 |
| 273 | 25.29 | 1.00 | -2.81 | 5.99 | A | A | A | 0.57 |
| 274 | 24.7 | 1.4 | -5.07 | 7.24 | A | A | A | 1.02 |
| 275 | 26.1 | 1.63 | 0.31 | 7.70 | A | A | A | 0.06 |
| 279 | 25.9 | 1.6 | -0.46 | 7.64 | A | A | A | 0.09 |
| 283 | 27.6 | 1.35 | 6.07 | 6.64 | A | A | A | 1.22 |
| 286 | 26.2 | 1.1 | 0.69 | 6.15 | A | A | A | 0.14 |
| 287 | 24.7 | 0.6 | -5.07 | 5.11 | A | A | A | 1.02 |
| 289 | 25.8 | 1.3 | -0.85 | 6.75 | A | A | A | 0.17 |
| 290 | 30.72 | 0.53 | 18.06 | 4.82 | A | N | W | 3.64 |
| 291 | 25.9 | 1.0 | -0.46 | 5.93 | A | A | A | 0.09 |
| 294 | 28.7 | 2.1 | 10.30 | 8.59 | A | A | A | 2.08 |
| 295 | 26.0 | 1.9 | -0.08 | 8.58 | A | A | A | 0.02 |
| 297 | 24.1 | 1.5 | -7.38 | 7.68 | A | A | A | 1.49 |
| 299 | 32.5 | 1.9 | 24.90 | 7.38 | A | N | W | 5.02 |
| 304 | 9.81 | 0.69 | -62.30 | 8.35 | N | N | N | 12.57 |
| 305 | 19.1 | 1.25 | -26.59 | 7.94 | N | N | N | 5.36 |
| 307 | 30.12 | 0.9 | 15.76 | 5.40 | A | N | W | 3.18 |
| 308 | 23.96 | 1.52 | -7.92 | 7.78 | A | A | A | 1.60 |
| 310 | 25.6 | 0.8 | -1.61 | 5.48 | A | A | A | 0.33 |
| 312 | 25.28 | 0.43 | -2.84 | 4.81 | A | A | A | 0.57 |
| 313 | 25.69 | 0.63 | -1.27 | 5.12 | A | A | A | 0.26 |
| 314 | 25.59 | 1.22 | -1.65 | 6.55 | A | A | A | 0.33 |
| 316 | 22.106 | 1.989 | -15.04 | 10.06 | A | A | A | 3.03 |
| 317 | 27 |  | 3.77 |  | A |  |  | 0.76 |
| 320 | 24.15 | 0.95 | -7.19 | 5.97 | A | A | A | 1.45 |
| 321 | 26.54 | 0.36 | 2.00 | 4.70 | A | A | A | 0.40 |
| 322 | 28.6 | 1.7 | 9.92 | 7.45 | A | A | A | 2.00 |
| 323 | 25.68 | 0.58 | -1.31 | 5.03 | A | A | A | 0.26 |
| 324 | 24.66 | 0.37 | -5.23 | 4.74 | A | A | A | 1.05 |
| 325 | 25.2 | 1.5 | -3.15 | 7.46 | A | A | A | 0.64 |

## Reported Results for Cs-137 Determination in Sample 7

TABLE 27. Target values

| Parameter | Value |
| --- | --- |
| Target Value, Bq/sample | 11.2 |
| UNC, Bq/sample | 0.6 |
| MARB, % | 20 |

FIGURE 27. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

FIGURE 28. Reported results and evaluation parameters for Cs-137. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

TABLE 28. Evaluation Results for Cs-137 in Sample 7, Rep. Value and Rep. Unc. in Bq/sample

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 11.86 | 0.34 | 5.89 | 6.08 | A | A | A | 0.84 |
| 2 | 11.01 | 0.9 | -1.70 | 9.77 | A | A | A | 0.24 |
| 4 | 11.17 | 0.42 | -0.27 | 6.55 | A | A | A | 0.04 |
| 5 | 10.9 | 0.7 | -2.68 | 8.36 | A | A | A | 0.38 |
| 6 | 12.2 | 0.4 | 8.93 | 6.28 | A | A | A | 1.27 |
| 7 | 24.55 | 1.39 | 119.20 | 7.79 | N | N | N | 16.90 |
| 9 | 12 | 2 | 7.14 | 17.51 | A | A | A | 1.01 |
| 10 | 11.15 | 0.32 | -0.45 | 6.08 | A | A | A | 0.06 |
| 14 | 10.5 | 0.5 | -6.25 | 7.17 | A | A | A | 0.89 |
| 15 | 11.7 | 0.7 | 4.46 | 8.03 | A | A | A | 0.63 |
| 16 | 12.1 | 0.6 | 8.04 | 7.30 | A | A | A | 1.14 |
| 17 | 11.22 | 0.54 | 0.18 | 7.20 | A | A | A | 0.03 |
| 18 | 10.60 | 0.63 | -5.36 | 8.00 | A | A | A | 0.76 |
| 19 | 11.5 | 0.7 | 2.68 | 8.11 | A | A | A | 0.38 |
| 21 | 10.89 | 0.77 | -2.77 | 8.87 | A | A | A | 0.39 |
| 22 | 12.7 | 0.3 | 13.39 | 5.85 | A | A | A | 1.90 |
| 26 | 11.66 | 0.47 | 4.11 | 6.70 | A | A | A | 0.58 |
| 27 | 11.0 | 0.6 | -1.79 | 7.65 | A | A | A | 0.25 |
| 32 | 11.2 | 1.8 | 0.00 | 16.94 | A | A | A | 0.00 |
| 34 | 11.5 | 0.5 | 2.68 | 6.90 | A | A | A | 0.38 |
| 35 | 11.2 | 0.75 | 0.00 | 8.58 | A | A | A | 0.00 |
| 36 | 11.6 | 0.6 | 3.57 | 7.45 | A | A | A | 0.51 |
| 38 | 12.8 | 0.8 | 14.29 | 8.23 | A | A | A | 2.03 |
| 41 | 10.9 | 0.5 | -2.68 | 7.05 | A | A | A | 0.38 |
| 43 | 11.35 | 0.35 | 1.34 | 6.18 | A | A | A | 0.19 |
| 45 | 11.1 | 0.4 | -0.89 | 6.46 | A | A | A | 0.13 |
| 46 | 8.95 | 0.22 | -20.09 | 5.89 | N | N | N | 2.85 |
| 48 | 11.4 | 0.4 | 1.79 | 6.40 | A | A | A | 0.25 |
| 50 | 11.93 | 1.79 | 6.52 | 15.93 | A | A | A | 0.92 |
| 51 | 12.5 | 0.5 | 11.61 | 6.69 | A | A | A | 1.65 |

TABLE 28. Evaluation Results for Cs-137 in Sample 7, Rep. Value and Rep. Unc. in Bq/sample

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 52 | 11.55 | 0.45 | 3.13 | 6.62 | A | A | A | 0.44 |
| 53 | 11.7 | 1.4 | 4.46 | 13.11 | A | A | A | 0.63 |
| 55 | 11.4 | 0.6 | 1.79 | 7.51 | A | A | A | 0.25 |
| 59 | 11.3 | 1.0 | 0.89 | 10.34 | A | A | A | 0.13 |
| 60 | 12.9 | 3.6 | 15.18 | 28.42 | A | N | W | 2.15 |
| 61 | 11.81 | 0.38 | 5.45 | 6.25 | A | A | A | 0.77 |
| 63 | 24.30 | 0.70 | 116.96 | 6.08 | N | N | N | 16.58 |
| 64 | 10.1 | 0.52 | -9.82 | 7.43 | A | A | A | 1.39 |
| 66 | 9.8 | 0.6 | -12.50 | 8.14 | A | A | A | 1.77 |
| 68 | 11.4 | 0.7 | 1.79 | 8.15 | A | A | A | 0.25 |
| 70 | 23.21 | 0.64 | 107.23 | 6.03 | N | N | N | 15.20 |
| 72 | 22.248 | 1.503 | 98.64 | 8.62 | N | N | N | 13.98 |
| 75 | 12.2 | 0.36 | 8.93 | 6.12 | A | A | A | 1.27 |
| 76 | 11.50 | 0.08 | 2.68 | 5.40 | A | A | A | 0.38 |
| 77 | 11.4 | 0.4 | 1.79 | 6.40 | A | A | A | 0.25 |
| 78 | 11.7 | 0.18 | 4.46 | 5.57 | A | A | A | 0.63 |
| 80 | 12.58 | 1.24 | 12.32 | 11.22 | A | A | A | 1.75 |
| 81 | 11.0 | 0.9 | -1.79 | 9.78 | A | A | A | 0.25 |
| 84 | 10.6 | 0.4 | -5.36 | 6.55 | A | A | A | 0.76 |
| 88 | 20.8 | 1.6 | 85.71 | 9.37 | N | N | N | 12.15 |
| 89 | 11.6 | 0.4 | 3.57 | 6.37 | A | A | A | 0.51 |
| 92 | 11.7 | 1.2 | 4.46 | 11.57 | A | A | A | 0.63 |
| 93 | 11.32 | 0.25 | 1.07 | 5.79 | A | A | A | 0.15 |
| 96 | 10.6 | 0.6 | -5.36 | 7.79 | A | A | A | 0.76 |
| 98 | 9.5 | 0.5 | -15.18 | 7.51 | A | A | A | 2.15 |
| 99 | 12.2 | 1.3 | 8.93 | 11.93 | A | A | A | 1.27 |
| 101 | 12.57 | 0.56 | 12.23 | 6.97 | A | A | A | 1.73 |
| 102 | 10.92 | 0.25 | -2.50 | 5.83 | A | A | A | 0.35 |
| 103 | 11.43 | 0.33 | 2.05 | 6.09 | A | A | A | 0.29 |
| 106 | 11.3 | 0.1 | 0.89 | 5.43 | A | A | A | 0.13 |
| 108 | 11.1537 | 0.2381 | -0.41 | 5.77 | A | A | A | 0.06 |
| 110 | 7.406 | 0.198 | -33.88 | 5.99 | N | N | N | 4.80 |
| 111 | 11.1 | 0.5 | -0.89 | 7.00 | A | A | A | 0.13 |
| 112 | 21.9079 | 0.7988 | 95.61 | 6.48 | N | N | N | 13.55 |
| 113 | 11.3 | 0.6 | 0.89 | 7.54 | A | A | A | 0.13 |
| 116 | 11.7 | 1.5 | 4.46 | 13.89 | A | A | A | 0.63 |
| 120 | 11.9 | 0.3 | 6.25 | 5.92 | A | A | A | 0.89 |
| 121 | 11.88 | 0.60 | 6.07 | 7.36 | A | A | A | 0.86 |
| 122 | 11.7 | 0.4 | 4.46 | 6.36 | A | A | A | 0.63 |
| 123 | 12.0 | 0.5 | 7.14 | 6.79 | A | A | A | 1.01 |
| 124 | 10.9 | 0.4 | -2.68 | 6.49 | A | A | A | 0.38 |
| 126 | 19 | 0.82 | 69.64 | 6.88 | N | N | N | 9.87 |
| 129 | 11.0 | 0.3 | -1.79 | 6.01 | A | A | A | 0.25 |
| 130 | 11.7 | 0.5 | 4.46 | 6.85 | A | A | A | 0.63 |
| 132 | 10.8 | 0.6 | -3.57 | 7.72 | A | A | A | 0.51 |
| 133 | 12.18 | 0.61 | 8.75 | 7.33 | A | A | A | 1.24 |
| 134 | 10.31 | 0.62 | -7.95 | 8.05 | A | A | A | 1.13 |
| 137 | 10.85 | 0.274 | -3.12 | 5.92 | A | A | A | 0.44 |
| 139 | 15.518 | 0.095 | 38.55 | 5.39 | N | N | N | 5.47 |
| 144 | 23.3 | 2.8 | 108.04 | 13.16 | N | N | N | 15.32 |
| 146 | 11.1 | 0.4 | -0.89 | 6.46 | A | A | A | 0.13 |
| 147 | 10.2 | 1.00 | -8.93 | 11.17 | A | A | A | 1.27 |
| 148 | 10.89 | 0.363 | -2.77 | 6.31 | A | A | A | 0.39 |
| 150 | 10.79 | 0.42 | -3.66 | 6.62 | A | A | A | 0.52 |

TABLE 28. Evaluation Results for Cs-137 in Sample 7, Rep. Value and Rep. Unc. in Bq/sample

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 152 | 11.4 | 0.4 | 1.79 | 6.40 | A | A | A | 0.25 |
| 153 | 10.7 | 0.4 | -4.46 | 6.53 | A | A | A | 0.63 |
| 154 | 13.3 | 0.3 | 18.75 | 5.81 | A | N | W | 2.66 |
| 156 | 11.0 | 0.8 | -1.79 | 9.03 | A | A | A | 0.25 |
| 157 | 11.59 | 0.85 | 3.48 | 9.08 | A | A | A | 0.49 |
| 160 | 11.7 | 0.4 | 4.46 | 6.36 | A | A | A | 0.63 |
| 161 | 11.0 | 0.3 | -1.79 | 6.01 | A | A | A | 0.25 |
| 162 | 10.35 | 0.29 | -7.59 | 6.05 | A | A | A | 1.08 |
| 163 | 9.7 | 0.6 | -13.39 | 8.18 | A | A | A | 1.90 |
| 164 | 11.6 | 0.4 | 3.57 | 6.37 | A | A | A | 0.51 |
| 166 | 11.11 | 0.59 | -0.80 | 7.54 | A | A | A | 0.11 |
| 168 | 9.91 | 0.54 | -11.52 | 7.64 | A | A | A | 1.63 |
| 169 | 11.80 | 0.45 | 5.36 | 6.58 | A | A | A | 0.76 |
| 171 | 10.62 | 1.10 | -5.18 | 11.66 | A | A | A | 0.73 |
| 172 | 12.1 | 1.7 | 8.04 | 15.04 | A | A | A | 1.14 |
| 173 | 11.1 | 1.3 | -0.89 | 12.88 | A | A | A | 0.13 |
| 177 | 11.03 | 1 | -1.52 | 10.53 | A | A | A | 0.22 |
| 178 | 11.49 | 0.58 | 2.59 | 7.36 | A | A | A | 0.37 |
| 180 | 22.34 | 0.72 | 99.46 | 6.25 | N | N | N | 14.10 |
| 181 | 11.09 | 0.39 | -0.98 | 6.41 | A | A | A | 0.14 |
| 182 | 21.89 | 1.46 | 95.45 | 8.55 | N | N | N | 13.53 |
| 183 | 22.32 | 0.6 | 99.29 | 5.99 | N | N | N | 14.08 |
| 184 | 12.08 | 0.07 | 7.86 | 5.39 | A | A | A | 1.11 |
| 185 | 11.5 | 0.8 | 2.68 | 8.78 | A | A | A | 0.38 |
| 186 | 11.21 | 0.67 | 0.09 | 8.03 | A | A | A | 0.01 |
| 187 | 108.4 | 0.4 | 867.86 | 5.37 | N | N | N | 123.04 |
| 189 | 11.48 | 0.70 | 2.50 | 8.12 | A | A | A | 0.35 |
| 190 | 12.2 | 1.5 | 8.93 | 13.41 | A | A | A | 1.27 |
| 191 | 11.1 | 0.41 | -0.89 | 6.51 | A | A | A | 0.13 |
| 193 | 11.63 | 1.16 | 3.84 | 11.32 | A | A | A | 0.54 |
| 195 | 12.0 | 0.50 | 7.14 | 6.79 | A | A | A | 1.01 |
| 196 | 12.10 | 0.54 | 8.04 | 6.97 | A | A | A | 1.14 |
| 197 | 10.8 | 0.5 | -3.57 | 7.08 | A | A | A | 0.51 |
| 198 | 11.38 | 0.72 | 1.61 | 8.29 | A | A | A | 0.23 |
| 199 | 12.16 | 0.24 | 8.57 | 5.71 | A | A | A | 1.22 |
| 201 | 11.3 | 1.1 | 0.89 | 11.11 | A | A | A | 0.13 |
| 202 | 22.07 | 0.67 | 97.05 | 6.16 | N | N | N | 13.76 |
| 205 | 10.7 | 0.5 | -4.46 | 7.11 | A | A | A | 0.63 |
| 206 | 10.9 | 0.6 | -2.68 | 7.68 | A | A | A | 0.38 |
| 208 | 11.0 | 0.53 | -1.79 | 7.21 | A | A | A | 0.25 |
| 209 | 11.7 | 4.7 | 4.46 | 40.53 | A | N | W | 0.63 |
| 210 | 11.8 | 0.9 | 5.36 | 9.32 | A | A | A | 0.76 |
| 212 | 11.3 | 0.6 | 0.89 | 7.54 | A | A | A | 0.13 |
| 213 | 11.4 | 0.7 | 1.79 | 8.15 | A | A | A | 0.25 |
| 214 | 11.9 | 0.8 | 6.25 | 8.60 | A | A | A | 0.89 |
| 215 | 9.89 | 0.60 | -11.70 | 8.09 | A | A | A | 1.66 |
| 216 | 22.39 | 0.95 | 99.91 | 6.83 | N | N | N | 14.16 |
| 219 | 11.1 | 0.8 | -0.89 | 8.98 | A | A | A | 0.13 |
| 221 | 34.33 | 1.51 | 206.52 | 6.93 | N | N | N | 29.28 |
| 223 | 9.58 | 0.23 | -14.46 | 5.87 | A | A | A | 2.05 |
| 227 | 11.57 | 0.16 | 3.30 | 5.53 | A | A | A | 0.47 |
| 228 | 10.9 | 0.4 | -2.68 | 6.49 | A | A | A | 0.38 |
| 230 | 10.8 | 0.4 | -3.57 | 6.51 | A | A | A | 0.51 |
| 231 | 10.74 | 1.915 | -4.11 | 18.62 | A | A | A | 0.58 |

TABLE 28. Evaluation Results for Cs-137 in Sample 7, Rep. Value and Rep. Unc. in Bq/sample

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 232 | 22.4079 | 0.9037 | 100.07 | 6.71 | N | N | N | 14.19 |
| 233 | 9.71 | 0.12 | -13.30 | 5.50 | A | A | A | 1.89 |
| 234 | 11.71 | 0.29 | 4.55 | 5.90 | A | A | A | 0.65 |
| 237 | 10.33 | 0.62 | -7.77 | 8.05 | A | A | A | 1.10 |
| 238 | 11.55 | 1.65 | 3.13 | 15.26 | A | A | A | 0.44 |
| 239 | 10.9 | 0.4 | -2.68 | 6.49 | A | A | A | 0.38 |
| 240 | 11.9 | 0.3 | 6.25 | 5.92 | A | A | A | 0.89 |
| 242 | 12.7 | 0.7 | 13.39 | 7.69 | A | A | A | 1.90 |
| 243 | 9.3 | 0.5 | -16.96 | 7.59 | A | A | A | 2.41 |
| 248 | 11.73 | 0.35 | 4.73 | 6.13 | A | A | A | 0.67 |
| 249 | 11.2 | 0.4 | 0.00 | 6.44 | A | A | A | 0.00 |
| 251 | 11.15 | 0.41 | -0.45 | 6.50 | A | A | A | 0.06 |
| 252 | 10.52 | 0.2 | -6.07 | 5.68 | A | A | A | 0.86 |
| 254 | 10.40 | 0.49 | -7.14 | 7.13 | A | A | A | 1.01 |
| 255 | 11.1 | 0.3 | -0.89 | 6.00 | A | A | A | 0.13 |
| 256 | 10.30 | 0.62 | -8.04 | 8.06 | A | A | A | 1.14 |
| 257 | 10.5 | 0.4 | -6.25 | 6.57 | A | A | A | 0.89 |
| 259 | 11.76 | 0.18 | 5.00 | 5.57 | A | A | A | 0.71 |
| 260 | 11.4 | 0.7 | 1.79 | 8.15 | A | A | A | 0.25 |
| 261 | 11.3 | 0.56 | 0.89 | 7.30 | A | A | A | 0.13 |
| 262 | 10.3 | 0.3 | -8.04 | 6.10 | A | A | A | 1.14 |
| 263 | 10.74 | 0.35 | -4.11 | 6.27 | A | A | A | 0.58 |
| 264 | 11.4 | 0.1 | 1.79 | 5.43 | A | A | A | 0.25 |
| 266 | 8.5 | 0.5 | -24.11 | 7.96 | N | N | N | 3.42 |
| 267 | 12.77 | 0.07 | 14.02 | 5.39 | A | N | W | 1.99 |
| 268 | 12.3 | 0.8 | 9.82 | 8.43 | A | A | A | 1.39 |
| 270 | 9.75 | 0.828 | -12.95 | 10.04 | A | A | A | 1.84 |
| 274 | 11.2 | 0.3 | 0.00 | 5.99 | A | A | A | 0.00 |
| 275 | 11.0 | 0.66 | -1.79 | 8.04 | A | A | A | 0.25 |
| 279 | 11.61 | 0.69 | 3.66 | 8.00 | A | A | A | 0.52 |
| 283 | 13.1 | 0.70 | 16.96 | 7.57 | A | A | A | 2.41 |
| 287 | 21.5 | 1.08 | 91.96 | 7.34 | N | N | N | 13.04 |
| 289 | 11.6 | 1.5 | 3.57 | 14.00 | A | A | A | 0.51 |
| 291 | 11.3 | 0.4 | 0.89 | 6.42 | A | A | A | 0.13 |
| 295 | 12.1 | 0.9 | 8.04 | 9.17 | A | A | A | 1.14 |
| 297 | 11.2 | 0.8 | 0.00 | 8.93 | A | A | A | 0.00 |
| 298 | 11.2 | 0.5 | 0.00 | 6.97 | A | A | A | 0.00 |
| 305 | 10.6 | 0.59 | -5.36 | 7.73 | A | A | A | 0.76 |
| 307 | 11.4 | 0.1 | 1.79 | 5.43 | A | A | A | 0.25 |
| 308 | 16.78 | 1.01 | 49.82 | 8.06 | N | N | N | 7.06 |
| 310 | 11.3 | 0.3 | 0.89 | 5.98 | A | A | A | 0.13 |
| 312 | 11.65 | 0.47 | 4.02 | 6.71 | A | A | A | 0.57 |
| 313 | 11.06 | 0.31 | -1.25 | 6.05 | A | A | A | 0.18 |
| 320 | 23.5 | 0.77 | 109.82 | 6.28 | N | N | N | 15.57 |
| 322 | 12.2 | 0.4 | 8.93 | 6.28 | A | A | A | 1.27 |
| 323 | 11.86 | 0.44 | 5.89 | 6.52 | A | A | A | 0.84 |
| 325 | 12.08 | 0.51 | 7.86 | 6.82 | A | A | A | 1.11 |
| 326 | 11.2 | 0.56 | 0.00 | 7.33 | A | A | A | 0.00 |

## Reported Results for Pu-239 Determination in Sample 7

TABLE 29. Target values

| Parameter | Value |
| --- | --- |
| Target Value, Bq/sample | 8.47 |
| UNC, Bq/sample | 0.5 |
| MARB, % | 40 |

FIGURE 29. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

FIGURE 30. Reported results and evaluation parameters for Pu-239. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

TABLE 30. Evaluation Results for Pu-239 in Sample 7, Rep. Value and Rep. Unc. in Bq/sample

| Labcode | Rep. Value | Rep. Unc. | Rel. Bias | P Test | Accuracy | Precision | Final Score | Z-Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 7 | 16.70 | 1.61 | 97.17 | 11.30 | N | N | N | 5.92 |
| 9 | 6.9 | 0.9 | -18.54 | 14.32 | A | A | A | 1.13 |
| 18 | 6.41 | 0.98 | -24.32 | 16.39 | A | A | A | 1.48 |
| 80 | 8.77 | 0.8 | 3.54 | 10.87 | A | A | A | 0.22 |
| 81 | 8.3 | 1.1 | -2.01 | 14.51 | A | A | A | 0.12 |
| 91 | 7.51 | 0.06 | -11.33 | 5.96 | A | A | A | 0.69 |
| 103 | 6.4 | 0.61 | -24.44 | 11.21 | A | A | A | 1.49 |
| 108 | 8.02732 | 0.04028 | -5.23 | 5.92 | A | A | A | 0.32 |
| 128 | 7.4 | 3.3 | -12.63 | 44.98 | A | N | W | 0.77 |
| 156 | 9.1 | 0.5 | 7.44 | 8.06 | A | A | A | 0.45 |
| 169 | 7.22 | 0.44 | -14.76 | 8.48 | A | A | A | 0.90 |
| 172 | 6.5 | 1.2 | -23.26 | 19.38 | A | A | A | 1.42 |
| 186 | 7.44 | 0.24 | -12.16 | 6.73 | A | A | A | 0.74 |
| 189 | 8.18 | 0.24 | -3.42 | 6.59 | A | A | A | 0.21 |
| 233 | 7.4555 | 0.3718 | -11.98 | 7.73 | A | A | A | 0.73 |
| 261 | 8.33 | 0.16 | -1.65 | 6.21 | A | A | A | 0.10 |
| 305 | 6.99 | 0.17 | -17.47 | 6.38 | A | N | W | 1.06 |
| 323 | 8.50 | 0.13 | 0.35 | 6.10 | A | A | A | 0.02 |

# Intercomparison  Parameters

## Reported Results for Gross Alpha Determination in Sample 1

TABLE 31. Target values

| Parameter | Value |
| --- | --- |
| Robust Mean, Bq/kg | 105 |
| Robust SD, Bq/kg | 26 |
| MARB, % | 50 |

FIGURE 31. Reported results and evaluation parameters for Gross Alpha in Sample 1.

TABLE 32. IC Evaluation Results for Gross Alpha in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 1 | 85.00 | 1.55 | 0.77 | A |
| 2 | 115.0 | 3.4 | 0.38 | A |
| 5 | 101 | 40 | 0.15 | A |
| 6 | 112.6 | 5.6 | 0.29 | A |
| 8 | 108.26 | 3.40 | 0.13 | A |
| 9 | 114 | 9 | 0.35 | A |
| 10 | 126 | 17 | 0.81 | A |
| 14 | 127.7 | 4.5 | 0.87 | A |
| 15 | 122.0 | 8.6 | 0.65 | A |
| 18 | 129.3 | 9.1 | 0.93 | A |
| 25 | 117.4 | 4.2 | 0.48 | A |
| 26 | 69.67 | 7.55 | 1.36 | A |
| 30 | 94.38 | 10.59 | 0.41 | A |
| 32 | 113.5 | 8.8 | 0.33 | A |
| 34 | 109 | 9 | 0.15 | A |
| 35 | 99.2 | 8.4 | 0.22 | A |
| 39 | 112 | 2 | 0.27 | A |
| 41 | 122.9 | 9.9 | 0.69 | A |
| 45 | 90.8 | 12.0 | 0.55 | A |
| 48 | 90.1 | 3.5 | 0.57 | A |

TABLE 32. IC Evaluation Results for Gross Alpha in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 50 | 118.31 | 16.68 | 0.51 | A |
| 51 | 113 | 5 | 0.31 | A |
| 58 | 95.92 | 9.15 | 0.35 | A |
| 59 | 118 | 6.25 | 0.50 | A |
| 60 | 118.9 | 45.1 | 0.53 | A |
| 61 | 59.97 | 2.22 | 1.73 | A |
| 63 | 112.5 | 1.4 | 0.29 | A |
| 66 | 117.5 | 0.7 | 0.48 | A |
| 69 | 123.27 | 6.30 | 0.70 | A |
| 73 | 122.52 | 0.89 | 0.67 | A |
| 76 | 91.14 | 7.21 | 0.53 | A |
| 78 | 110 | 0.76 | 0.19 | A |
| 79 | 119 | 9 | 0.54 | A |
| 84 | 55.3 | 7.66 | 1.91 | A |
| 88 | 118.1 | 14.5 | 0.50 | A |
| 90 | 21.06 | 8.3 | 3.23 | N |
| 96 | 88.5 | 9.0 | 0.63 | A |
| 98 | 114.4 | 9.1 | 0.36 | A |
| 99 | 63.8 | 11.4 | 1.58 | A |
| 100 | 301.52 | 152.12 | 7.56 | N |
| 101 | 119 | 3 | 0.54 | A |
| 102 | 105.2 | 1.1 | 0.01 | A |
| 103 | 128.11 | 19.22 | 0.89 | A |
| 108 | 37.8336 | 0.6665 | 2.58 | W |
| 111 | 110 | 40 | 0.19 | A |
| 112 | 47.5900 | 0.8100 | 2.21 | W |
| 113 | 134.02 | 6.71 | 1.12 | A |
| 116 | 158.6 | 24.0 | 2.06 | W |
| 117 | 58.11 | 5.92 | 1.80 | A |
| 120 | 85.9 | 5.0 | 0.73 | A |
| 123 | 115 | 6 | 0.38 | A |
| 125 | 56 | 1.8 | 1.88 | A |
| 135 | 120 | 2 | 0.58 | A |
| 136 | 144.42 | 35.40 | 1.52 | A |
| 137 | 87 | 2.69 | 0.69 | A |
| 148 | 69.6 | 0.871 | 1.36 | A |
| 150 | 103.6 | 1.1 | 0.05 | A |
| 151 | 94.0 | 10.0 | 0.42 | A |
| 153 | 123 | 12 | 0.69 | A |
| 154 | 113.9 | 22.7 | 0.34 | A |
| 156 | 64 | 9 | 1.58 | A |
| 157 | 36.95 | 1.30 | 2.62 | W |
| 162 | 111.14 | 1.69 | 0.24 | A |
| 164 | 122 | 6 | 0.65 | A |
| 166 | 118.76 | 8.20 | 0.53 | A |
| 169 | 83.90 | 2.56 | 0.81 | A |
| 171 | 110.24 | 13.23 | 0.20 | A |
| 172 | 85.9 | 16.3 | 0.73 | A |
| 180 | 89.9 | 6.2 | 0.58 | A |
| 181 | 87.27 | 0.52 | 0.68 | A |
| 182 | 89.26 | 0.96 | 0.61 | A |
| 183 | 28.8 | 2.0 | 2.93 | W |
| 185 | 136 | 19 | 1.19 | A |
| 186 | 122.2 | 7.0 | 0.66 | A |

TABLE 32. IC Evaluation Results for Gross Alpha in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 189 | 103.3 | 3.8 | 0.07 | A |
| 190 | 146 | 32 | 1.58 | A |
| 193 | 26.1 | 5 | 3.03 | N |
| 195 | 103.1 | 11.8 | 0.07 | A |
| 197 | 130 | 8 | 0.96 | A |
| 207 | 116 | 12 | 0.42 | A |
| 208 | 120.9 | 8.7 | 0.61 | A |
| 210 | 111 | 9 | 0.23 | A |
| 212 | 120 | 5 | 0.58 | A |
| 213 | 104 | 8 | 0.04 | A |
| 215 | 72.00 | 0.70 | 1.27 | A |
| 219 | 126.5 | 9.2 | 0.83 | A |
| 221 | 0.42 | 0.05 | 4.02 | N |
| 224 | 79.5 | 3.1 | 0.98 | A |
| 227 | 86.73 | 0.80 | 0.70 | A |
| 228 | 128 | 4.70 | 0.88 | A |
| 233 | 102.90 | 9.26 | 0.08 | A |
| 235 | 190 | 6 | 3.27 | N |
| 238 | 102.2 | 12.78 | 0.11 | A |
| 240 | 109 | 3 | 0.15 | A |
| 242 | 119 | 18 | 0.54 | A |
| 243 | 95.4 | 9.0 | 0.37 | A |
| 248 | 113.37 | 3.34 | 0.32 | A |
| 249 | 120 | 36 | 0.58 | A |
| 250 | 44.3 | 0.7 | 2.33 | W |
| 251 | 119.9 | 6.3 | 0.57 | A |
| 252 | 91.31 | 2.96 | 0.53 | A |
| 253 | 116 | 18 | 0.42 | A |
| 254 | 105 | 22 | 0.00 | A |
| 255 | 107.0 | 7.0 | 0.08 | A |
| 256 | 71.092 | 0.553 | 1.30 | A |
| 257 | 97.9 | 5.3 | 0.27 | A |
| 259 | 97.4 | 1.95 | 0.29 | A |
| 261 | 125 | 6.3 | 0.77 | A |
| 266 | 115.0 | 4.3 | 0.38 | A |
| 267 | 57.72 | 3.004 | 1.82 | A |
| 268 | 101.4 | 12.2 | 0.14 | A |
| 269 | 102 | 3 | 0.12 | A |
| 280 | 127.407 | 9.895 | 0.86 | A |
| 283 | 129 | 12.9 | 0.92 | A |
| 287 | 8 | 0.89 | 3.73 | N |
| 291 | 106 | 8 | 0.04 | A |
| 295 | 92.13 | 0.57 | 0.50 | A |
| 305 | 116 | 6.61 | 0.42 | A |
| 310 | 110.8 | 15 | 0.22 | A |
| 313 | 156.2 | 6.3 | 1.97 | A |
| 317 | 125 |  | 0.77 | A |
| 320 | 122.62 | 14.71 | 0.68 | A |
| 322 | 60 | 14 | 1.73 | A |
| 324 | 103.8 | 14.8 | 0.05 | A |

## Reported Results for Gross Beta Determination in Sample 1

TABLE 33. Target values

| Parameter | Value |
| --- | --- |
| Robust Mean, Bq/kg | 159 |
| Robust SD, Bq/kg | 45 |
| MARB, % | 55 |

FIGURE 32. Reported results and evaluation parameters for Gross Beta in Sample 1.

TABLE 34. IC Evaluation Results for Gross Beta in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 1 | 17.29 | 0.5 | 3.15 | N |
| 2 | 146.1 | 3.2 | 0.29 | A |
| 5 | 729 | 304 | 12.67 | N |
| 6 | 137.6 | 6.9 | 0.48 | A |
| 8 | 129.59 | 0.89 | 0.65 | A |
| 9 | 252 | 21 | 2.07 | W |
| 10 | 183 | 19 | 0.53 | A |
| 14 | 133.4 | 2.9 | 0.57 | A |
| 15 | 121.2 | 8.6 | 0.84 | A |
| 18 | 262 | 19 | 2.29 | W |
| 25 | 120.1 | 4.3 | 0.86 | A |
| 26 | 134.85 | 14.99 | 0.54 | A |
| 30 | 106.73 | 11.95 | 1.16 | A |
| 32 | 251.8 | 21.3 | 2.06 | W |
| 34 | 152 | 12 | 0.16 | A |
| 35 | 209 | 18 | 1.11 | A |
| 39 | 309 | 2.5 | 3.33 | N |
| 41 | 138.5 | 7.0 | 0.46 | A |
| 45 | 139.0 | 15.5 | 0.44 | A |
| 48 | 192.5 | 7.3 | 0.74 | A |
| 50 | 101.23 | 9.47 | 1.28 | A |

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 51 | 129 | 5 | 0.67 | A |
| 55 | 142.1 | 14.7 | 0.38 | A |
| 58 | 169.68 | 42.82 | 0.24 | A |
| 59 | 123 | 6.27 | 0.80 | A |
| 60 | 50.2 | 19.2 | 2.42 | W |
| 61 | 93.31 | 2.92 | 1.46 | A |
| 63 | 130.9 | 1.5 | 0.62 | A |
| 66 | 257.7 | 1.3 | 2.19 | W |
| 69 | 248.10 | 12.65 | 1.98 | A |
| 73 | 223.2 | 1.6 | 1.43 | A |
| 76 | 152.0 | 5.2 | 0.16 | A |
| 78 | 170 | 0.86 | 0.24 | A |
| 79 | 165 | 13 | 0.13 | A |
| 84 | 179.7 | 6 | 0.46 | A |
| 88 | 163.4 | 14.4 | 0.10 | A |
| 90 | 66.48 | 8.2 | 2.06 | W |
| 96 | 174.1 | 17.4 | 0.34 | A |
| 98 | 233.1 | 21.2 | 1.65 | A |
| 99 | 200.0 | 19.4 | 0.91 | A |
| 100 | 116.33 | 58.38 | 0.95 | A |
| 101 | 215 | 8 | 1.24 | A |
| 102 | 122.5 | 1.1 | 0.81 | A |
| 103 | 130.76 | 19.61 | 0.63 | A |
| 108 | 128.5346 | 2.2347 | 0.68 | A |
| 111 | 106 | 25 | 1.18 | A |
| 112 | 124.3297 | 1.7426 | 0.77 | A |
| 113 | 124.34 | 6.86 | 0.77 | A |
| 116 | 128.8 | 13.6 | 0.67 | A |
| 117 | 109.84 | 11.26 | 1.09 | A |
| 118 | 128.7 | 13.1 | 0.67 | A |
| 120 | 156 | 13 | 0.07 | A |
| 123 | 182 | 3 | 0.51 | A |
| 135 | 185 | 3 | 0.58 | A |
| 136 | 169.03 | 2.32 | 0.22 | A |
| 137 | 163 | 2.34 | 0.09 | A |
| 148 | 165.6 | 1.028 | 0.15 | A |
| 150 | 125.7 | 1.1 | 0.74 | A |
| 151 | 121.0 | 12.0 | 0.84 | A |
| 153 | 130 | 16 | 0.64 | A |
| 154 | 178.2 | 7.2 | 0.43 | A |
| 156 | 187 | 27 | 0.62 | A |
| 157 | 122.55 | 2.08 | 0.81 | A |
| 162 | 111.02 | 1.54 | 1.07 | A |
| 164 | 153 | 9 | 0.13 | A |
| 166 | 106.09 | 4.69 | 1.18 | A |
| 169 | 188.27 | 3.68 | 0.65 | A |
| 171 | 170.22 | 20.43 | 0.25 | A |
| 172 | 28.0 | 4.1 | 2.91 | W |
| 180 | 99.6 | 5.6 | 1.32 | A |
| 181 | 151.50 | 0.41 | 0.17 | A |
| 182 | 99.34 | 0.49 | 1.33 | A |
| 183 | 229.6 | 4.8 | 1.57 | A |
| 186 | 167.1 | 9.6 | 0.18 | A |
| 189 | 200.0 | 8.2 | 0.91 | A |

TABLE 34. IC Evaluation Results for Gross Beta in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

TABLE 34. IC Evaluation Results for Gross Beta in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 190 | 188 | 56 | 0.64 | A |
| 193 | 88.1 | 9 | 1.58 | A |
| 195 | 182.8 | 13.0 | 0.53 | A |
| 197 | 105 | 5 | 1.20 | A |
| 200 | 145.10 | 11.93 | 0.31 | A |
| 201 | 185 | 19 | 0.58 | A |
| 202 | 138.11 | 4.47 | 0.46 | A |
| 207 | 149 | 15 | 0.22 | A |
| 208 | 198.7 | 8.7 | 0.88 | A |
| 210 | 175 | 14 | 0.36 | A |
| 212 | 176 | 5 | 0.38 | A |
| 213 | 169 | 13 | 0.22 | A |
| 215 | 195.00 | 0.10 | 0.80 | A |
| 216 | 145.8 | 4.64 | 0.29 | A |
| 219 | 121.2 | 8.2 | 0.84 | A |
| 221 | 1.24 | 0.11 | 3.51 | N |
| 224 | 175.2 | 6.2 | 0.36 | A |
| 227 | 162.79 | 0.93 | 0.08 | A |
| 228 | 125 | 4.33 | 0.76 | A |
| 233 | 171.6 | 8.99 | 0.28 | A |
| 235 | 130 | 4.10 | 0.64 | A |
| 238 | 259.83 | 31.18 | 2.24 | W |
| 240 | 151 | 4 | 0.18 | A |
| 242 | 182 | 22 | 0.51 | A |
| 243 | 186.2 | 14.0 | 0.60 | A |
| 248 | 130.63 | 5.57 | 0.63 | A |
| 249 | 283 | 85 | 2.76 | W |
| 250 | 142.8 | 2.7 | 0.36 | A |
| 251 | 158.9 | 3.5 | 0.00 | A |
| 252 | 113.55 | 2.85 | 1.01 | A |
| 253 | 162 | 24 | 0.07 | A |
| 254 | 260 | 54 | 2.24 | W |
| 255 | 142.0 | 5.0 | 0.38 | A |
| 256 | 164.330 | 0.637 | 0.12 | A |
| 257 | 235 | 7 | 1.69 | A |
| 259 | 152 | 3.04 | 0.16 | A |
| 261 | 220 | 12 | 1.36 | A |
| 266 | 173.4 | 3.9 | 0.32 | A |
| 267 | 173.16 | 1.08 | 0.31 | A |
| 268 | 235 | 12 | 1.69 | A |
| 269 | 100 | 5 | 1.31 | A |
| 280 | 155.418 | 10.167 | 0.08 | A |
| 283 | 132 | 13.2 | 0.60 | A |
| 287 | 45 | 2.12 | 2.53 | W |
| 291 | 402 | 33 | 5.40 | N |
| 295 | 191.67 | 0.88 | 0.73 | A |
| 305 | 256 |  | 2.16 | W |
| 310 | 222.5 | 40 | 1.41 | A |
| 313 | 172.8 | 6.3 | 0.31 | A |
| 317 | 137.3 |  | 0.48 | A |
| 320 | 188.87 | 22.66 | 0.66 | A |
| 322 | 180 | 45 | 0.47 | A |
| 324 | 196.9 | 6.1 | 0.84 | A |

## Reported Results for Gross Alpha Determination in Sample 2

TABLE 35. Target values

| Parameter | Value |
| --- | --- |
| Robust Mean, Bq/kg | 97 |
| Robust SD, Bq/kg | 28 |
| MARB, % | 55 |

FIGURE 33. Reported results and evaluation parameters for Gross Alpha in Sample 2.

TABLE 36. IC Evaluation Results for Gross Alpha in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 1 | 19.49 | 0.60 | 2.77 | W |
| 2 | 111.2 | 3.3 | 0.51 | A |
| 5 | 124 | 6 | 0.96 | A |
| 6 | 79.9 | 4 | 0.61 | A |
| 8 | 95.18 | 2.99 | 0.06 | A |
| 9 | 112 | 12 | 0.54 | A |
| 10 | 110 | 10 | 0.46 | A |
| 14 | 107.8 | 5.0 | 0.39 | A |
| 15 | 107.0 | 7.6 | 0.36 | A |
| 18 | 124.4 | 8.8 | 0.98 | A |
| 25 | 106.8 | 3.9 | 0.35 | A |
| 26 | 103.36 | 5.36 | 0.23 | A |
| 30 | 11264 | 1282 | 398.82 | N |
| 32 | 112.3 | 11.4 | 0.55 | A |
| 34 | 91.2 | 7.5 | 0.21 | A |
| 35 | 112 | 9.3 | 0.54 | A |
| 39 | 52 | 1.5 | 1.61 | A |
| 41 | 111.3 | 8.9 | 0.51 | A |
| 45 | 58.4 | 8.5 | 1.38 | A |
| 48 | 94.3 | 3.7 | 0.10 | A |
| 50 | 106.11 | 15.83 | 0.33 | A |

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 51 | 164 | 5 | 2.39 | W |
| 58 | 75.22 | 7.23 | 0.78 | A |
| 59 | 110 | 7.04 | 0.46 | A |
| 60 | 100.5 | 36.9 | 0.12 | A |
| 61 | 46.37 | 1.86 | 1.81 | A |
| 63 | 93.2 | 1.2 | 0.14 | A |
| 66 | 108.7 | 0.7 | 0.42 | A |
| 69 | 108.81 | 5.57 | 0.42 | A |
| 73 | 367.31 | 3.10 | 9.65 | N |
| 76 | 80.80 | 6.20 | 0.58 | A |
| 78 | 75.4 | 0.61 | 0.77 | A |
| 79 | 104 | 8 | 0.25 | A |
| 84 | 49.7 | 9 | 1.69 | A |
| 88 | 110.8 | 11.4 | 0.49 | A |
| 90 | 14.22 | 8.15 | 2.96 | W |
| 98 | 108.7 | 8.5 | 0.42 | A |
| 99 | 53.5 | 21.1 | 1.55 | A |
| 100 | 204.39 | 103.19 | 3.84 | N |
| 101 | 92 | 4 | 0.18 | A |
| 102 | 121.1 | 1.2 | 0.86 | A |
| 103 | 121.47 | 18.22 | 0.87 | A |
| 108 | 32.5136 | 0.6182 | 2.30 | W |
| 111 | 104 | 37 | 0.25 | A |
| 112 | 40.5199 | 0.7501 | 2.02 | W |
| 113 | 94.03 | 8.61 | 0.11 | A |
| 116 | 77.0 | 11.7 | 0.71 | A |
| 117 | 44.41 | 4.97 | 1.88 | A |
| 120 | 107 | 6 | 0.36 | A |
| 123 | 105 | 6 | 0.29 | A |
| 135 | 114 | 2 | 0.61 | A |
| 136 | 118.97 | 28.90 | 0.78 | A |
| 137 | 88 | 2.76 | 0.32 | A |
| 148 | 51.0 | 0.733 | 1.64 | A |
| 150 | 118.3 | 1.2 | 0.76 | A |
| 151 | 75.0 | 6.0 | 0.79 | A |
| 153 | 115 | 11 | 0.64 | A |
| 154 | 90.0 | 18.5 | 0.25 | A |
| 156 | 55 | 8 | 1.50 | A |
| 157 | 35.54 | 1.28 | 2.19 | W |
| 164 | 111 | 6 | 0.50 | A |
| 166 | 101.49 | 7.10 | 0.16 | A |
| 169 | 68.94 | 2.34 | 1.00 | A |
| 171 | 100.12 | 12.02 | 0.11 | A |
| 172 | 121.3 | 38.1 | 0.87 | A |
| 180 | 75.2 | 3.5 | 0.78 | A |
| 181 | 94.86 | 0.54 | 0.08 | A |
| 182 | 73.45 | 0.74 | 0.84 | A |
| 183 | 27.1 | 2.0 | 2.50 | W |
| 185 | 112 | 14 | 0.54 | A |
| 186 | 100 | 10 | 0.11 | A |
| 189 | 96.7 | 3.7 | 0.01 | A |
| 190 | 135 | 35 | 1.36 | A |
| 193 | 41.6 | 8 | 1.98 | A |
| 195 | 90.8 | 13.3 | 0.22 | A |

TABLE 36. IC Evaluation Results for Gross Alpha in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

TABLE 36. IC Evaluation Results for Gross Alpha in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 197 | 120 | 7 | 0.82 | A |
| 207 | 112 | 12 | 0.54 | A |
| 208 | 99.6 | 7.97 | 0.09 | A |
| 211 | 101 | 3.6 | 0.14 | A |
| 212 | 106 | 5 | 0.32 | A |
| 215 | 55.00 | 0.60 | 1.50 | A |
| 219 | 143.1 | 9.9 | 1.65 | A |
| 221 | 0.34 | 0.031 | 3.45 | N |
| 224 | 78.5 | 2.5 | 0.66 | A |
| 227 | 73.57 | 0.70 | 0.84 | A |
| 228 | 103 | 3.81 | 0.21 | A |
| 233 | 97.29 | 8.75 | 0.01 | A |
| 235 | 316.66 | 4.8 | 7.85 | N |
| 238 | 104.16 | 13.02 | 0.26 | A |
| 240 | 86 | 3 | 0.39 | A |
| 242 | 102 | 17 | 0.18 | A |
| 243 | 75.7 | 7.4 | 0.76 | A |
| 248 | 100.83 | 3.09 | 0.14 | A |
| 249 | 125 | 38 | 1.00 | A |
| 250 | 33.5 | 0.6 | 2.27 | W |
| 251 | 121.9 | 2.4 | 0.89 | A |
| 252 | 61.34 | 2.0 | 1.27 | A |
| 253 | 97.1 | 14.6 | 0.00 | A |
| 254 | 109 | 23 | 0.43 | A |
| 255 | 100.0 | 7.0 | 0.11 | A |
| 256 | 54.260 | 0.483 | 1.53 | A |
| 257 | 92 | 5 | 0.18 | A |
| 259 | 94.5 | 1.89 | 0.09 | A |
| 261 | 106 | 5.4 | 0.32 | A |
| 266 | 104.7 | 3.9 | 0.28 | A |
| 267 | 41.29 | 2.23 | 1.99 | A |
| 268 | 155 | 19 | 2.07 | W |
| 269 | 101 | 3 | 0.14 | A |
| 280 | 120.260 | 9.343 | 0.83 | A |
| 283 | 118 | 11.8 | 0.75 | A |
| 287 | 0.2 | 0.14 | 3.46 | N |
| 291 | 154 | 10 | 2.04 | W |
| 295 | 88.08 | 0.55 | 0.32 | A |
| 305 | 97.3 | 7.70 | 0.01 | A |
| 310 | 90.6 | 14 | 0.23 | A |
| 313 | 215.1 | 8.7 | 4.22 | N |
| 317 | 124 |  | 0.96 | A |
| 320 | 103.06 | 12.37 | 0.22 | A |
| 322 | 95 | 24 | 0.07 | A |
| 324 | 85.1 | 24.7 | 0.43 | A |

## Reported Results for Gross Beta Determination in Sample 2

TABLE 37. Target values

| Parameter | Value |
| --- | --- |
| Robust Mean, Bq/kg | 457 |
| Robust SD, Bq/kg | 112 |
| MARB, % | 50 |

FIGURE 34. Reported results and evaluation parameters for Gross Beta in Sample 2.

TABLE 38. IC Evaluation Results for Gross Beta in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 1 | 245.78 | 4.08 | 1.89 | A |
| 2 | 967.0 | 10.5 | 4.55 | N |
| 6 | 442.6 | 22.1 | 0.13 | A |
| 8 | 385.80 | 2.40 | 0.64 | A |
| 9 | 600 | 70 | 1.28 | A |
| 10 | 541 | 25 | 0.75 | A |
| 14 | 491.2 | 8.7 | 0.31 | A |
| 15 | 438.1 | 31.0 | 0.17 | A |
| 18 | 691 | 50 | 2.09 | W |
| 25 | 452.5 | 13.2 | 0.04 | A |
| 26 | 423.24 | 25.64 | 0.30 | A |
| 30 | 24952 | 2794 | 218.71 | N |
| 32 | 600.0 | 72.1 | 1.28 | A |
| 34 | 435 | 35 | 0.20 | A |
| 35 | 561 | 46 | 0.93 | A |
| 39 | 1159 | 6.5 | 6.27 | N |
| 41 | 447.2 | 22.7 | 0.09 | A |
| 45 | 280.0 | 39.0 | 1.58 | A |
| 48 | 438 | 17 | 0.17 | A |
| 50 | 375.32 | 48.23 | 0.73 | A |
| 51 | 576 | 5 | 1.06 | A |

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 55 | 434.4 | 44.8 | 0.20 | A |
| 58 | 488.35 | 123.11 | 0.28 | A |
| 59 | 367 | 22.35 | 0.80 | A |
| 60 | 291.5 | 98.7 | 1.48 | A |
| 61 | 239.35 | 7.29 | 1.94 | A |
| 63 | 403.4 | 3.1 | 0.48 | A |
| 66 | 1974.4 | 3.7 | 13.55 | N |
| 69 | 1926.00 | 97.85 | 13.12 | N |
| 73 | 3148.3 | 9.9 | 24.03 | N |
| 76 | 415.0 | 14.4 | 0.38 | A |
| 78 | 470 | 1.4 | 0.12 | A |
| 79 | 497 | 37 | 0.36 | A |
| 84 | 473.75 | 26 | 0.15 | A |
| 88 | 482.9 | 11.3 | 0.23 | A |
| 90 | 169.11 | 8.6 | 2.57 | W |
| 98 | 625.7 | 56.5 | 1.51 | A |
| 99 | 508.0 | 81.9 | 0.46 | A |
| 100 | 411.81 | 206.12 | 0.40 | A |
| 101 | 550 | 39 | 0.83 | A |
| 102 | 372.8 | 2.7 | 0.75 | A |
| 103 | 753.49 | 113.02 | 2.65 | W |
| 108 | 297.4197 | 3.0517 | 1.42 | A |
| 111 | 320 | 77 | 1.22 | A |
| 112 | 405.4258 | 2.9907 | 0.46 | A |
| 113 | 805.65 | 25.01 | 3.11 | N |
| 116 | 380.0 | 38.7 | 0.69 | A |
| 117 | 287.23 | 27.81 | 1.52 | A |
| 118 | 387.1 | 47.2 | 0.62 | A |
| 120 | 491 | 42 | 0.30 | A |
| 123 | 514 | 7 | 0.51 | A |
| 135 | 987 | 11 | 4.73 | N |
| 136 | 473.46 | 12.57 | 0.15 | A |
| 137 | 458 | 3.90 | 0.01 | A |
| 148 | 441 | 1.63 | 0.14 | A |
| 150 | 371.4 | 2.7 | 0.76 | A |
| 151 | 407.0 | 40.0 | 0.45 | A |
| 153 | 379 | 39 | 0.70 | A |
| 154 | 487.4 | 17.9 | 0.27 | A |
| 156 | 481 | 68 | 0.21 | A |
| 157 | 378.93 | 3.67 | 0.70 | A |
| 164 | 485 | 30 | 0.25 | A |
| 166 | 348.05 | 15.40 | 0.97 | A |
| 169 | 502.32 | 6.42 | 0.40 | A |
| 171 | 572.50 | 74.43 | 1.03 | A |
| 172 | 92.7 | 14.6 | 3.25 | N |
| 180 | 340.1 | 7.8 | 1.04 | A |
| 181 | 419.10 | 0.62 | 0.34 | A |
| 182 | 340.73 | 1.12 | 1.04 | A |
| 183 | 494.57 | 8.24 | 0.34 | A |
| 186 | 486 | 26 | 0.26 | A |
| 189 | 1930.0 | 16.4 | 13.15 | N |
| 190 | 680 | 136 | 1.99 | A |
| 193 | 364 | 36 | 0.83 | A |
| 195 | 482.0 | 21.4 | 0.22 | A |

TABLE 38. IC Evaluation Results for Gross Beta in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

TABLE 38. IC Evaluation Results for Gross Beta in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 197 | 330 | 15 | 1.13 | A |
| 200 | 383.72 | 19.18 | 0.65 | A |
| 201 | 504 | 50 | 0.42 | A |
| 202 | 390.12 | 9.95 | 0.60 | A |
| 207 | 440 | 45 | 0.15 | A |
| 208 | 608.6 | 17.65 | 1.35 | A |
| 210 | 484 | 38 | 0.24 | A |
| 212 | 452 | 11 | 0.04 | A |
| 213 | 479 | 36 | 0.20 | A |
| 215 | 498.00 | 0.18 | 0.37 | A |
| 216 | 425.7 | 12.3 | 0.28 | A |
| 219 | 381.2 | 18.3 | 0.68 | A |
| 221 | 2.69 | 0.21 | 4.06 | N |
| 224 | 493 | 36 | 0.32 | A |
| 227 | 472.69 | 2.04 | 0.14 | A |
| 228 | 379 | 11.90 | 0.70 | A |
| 233 | 497.70 | 25.98 | 0.36 | A |
| 235 | 853.33 | 13.05 | 3.54 | N |
| 238 | 444.95 | 53.39 | 0.11 | A |
| 240 | 434 | 9 | 0.21 | A |
| 242 | 565 | 79 | 0.96 | A |
| 243 | 475.8 | 33.8 | 0.17 | A |
| 248 | 383.70 | 14.20 | 0.65 | A |
| 249 | 930 | 280 | 4.22 | N |
| 250 | 438.4 | 6.5 | 0.17 | A |
| 251 | 440.1 | 10.8 | 0.15 | A |
| 252 | 269.51 | 6.80 | 1.67 | A |
| 253 | 487 | 73 | 0.27 | A |
| 254 | 610 | 130 | 1.37 | A |
| 255 | 453.0 | 8.0 | 0.04 | A |
| 256 | 401.991 | 0.996 | 0.49 | A |
| 257 | 610 | 33 | 1.37 | A |
| 259 | 454 | 9.08 | 0.03 | A |
| 261 | 475 | 25 | 0.16 | A |
| 266 | 445.7 | 10.0 | 0.10 | A |
| 267 | 414.4 | 8.31 | 0.38 | A |
| 268 | 316 | 14 | 1.26 | A |
| 269 | 280 | 5 | 1.58 | A |
| 280 | 451.318 | 29.467 | 0.05 | A |
| 283 | 974 | 97.4 | 4.62 | N |
| 287 | 20.7 | 1.44 | 3.90 | N |
| 291 | 144 | 7 | 2.79 | W |
| 295 | 475.70 | 1.82 | 0.17 | A |
| 305 | 2202 |  | 15.58 | N |
| 310 | 353.9 | 55 | 0.92 | A |
| 313 | 411.4 | 14.9 | 0.41 | A |
| 317 | 494 |  | 0.33 | A |
| 320 | 508.67 | 61.04 | 0.46 | A |
| 322 | 440 | 110 | 0.15 | A |
| 324 | 534.3 | 34.2 | 0.69 | A |

## Reported Results for H-3 Determination in Sample 2

TABLE 39. Target values

| Parameter | Value |
| --- | --- |
| Robust Mean, Bq/kg | 1653.6 |
| Robust SD, Bq/kg | 98.2 |
| MARB, % | 30 |

FIGURE 35. Reported results and evaluation parameters for H-3 in Sample 2.

TABLE 40. IC Evaluation Results for H-3 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 2 | 600.0 | 15 | 10.73 | N |
| 5 | 1680 | 100 | 0.27 | A |
| 9 | 1600 | 195 | 0.55 | A |
| 10 | 1633 | 106 | 0.21 | A |
| 12 | 1471 | 40 | 1.86 | A |
| 14 | 1760.0 | 70.7 | 1.08 | A |
| 18 | 1737 | 91 | 0.85 | A |
| 19 | 1705.6 | 127.9 | 0.53 | A |
| 25 | 1636 | 41 | 0.18 | A |
| 26 | 1773.57 | 70.48 | 1.22 | A |
| 32 | 1605 | 193 | 0.49 | A |
| 34 | 1608 | 69 | 0.46 | A |
| 35 | 1469 | 53 | 1.88 | A |
| 38 | 1740 | 40 | 0.88 | A |
| 40 | 1884 | 222 | 2.35 | W |
| 41 | 1661.3 | 42.6 | 0.08 | A |
| 45 | 1691 | 75 | 0.38 | A |
| 50 | 1734.85 | 319.96 | 0.83 | A |
| 55 | 1710.2 | 143.5 | 0.58 | A |
| 59 | 1580 | 47.23 | 0.75 | A |
| 61 | 1670.07 | 102.75 | 0.17 | A |

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 69 | 399.59 | 0.06 | 12.77 | N |
| 70 | 1671 | 7 | 0.18 | A |
| 73 | 1363.0 | 14.4 | 2.96 | W |
| 76 | 1670.0 | 43.2 | 0.17 | A |
| 81 | 1661 | 139 | 0.08 | A |
| 85 | 1740.1 | 41.8 | 0.88 | A |
| 86 | 1479 | 185 | 1.78 | A |
| 88 | 1687 | 128 | 0.34 | A |
| 98 | 1763 | 280 | 1.11 | A |
| 99 | 1650 | 598 | 0.04 | A |
| 100 | 857.777 | 13.226 | 8.10 | N |
| 101 | 1607 | 117 | 0.47 | A |
| 102 | 1567 | 30 | 0.88 | A |
| 103 | 1677.02 | 167.70 | 0.24 | A |
| 108 | 1624.8385 | 8.6980 | 0.29 | A |
| 113 | 1577.61 | 15 | 0.77 | A |
| 116 | 1613 | 31 | 0.41 | A |
| 118 | 1679.0 | 100.9 | 0.26 | A |
| 123 | 1596 | 45 | 0.59 | A |
| 125 | 1600 | 45 | 0.55 | A |
| 132 | 1491 | 59 | 1.66 | A |
| 134 | 1525.45 | 147.82 | 1.30 | A |
| 136 | 1580 | 18 | 0.75 | A |
| 137 | 1726 | 225 | 0.74 | A |
| 141 | 1680 | 110 | 0.27 | A |
| 147 | 1700 | 144 | 0.47 | A |
| 148 | 1895 | 20.7 | 2.46 | W |
| 150 | 1612 | 31 | 0.42 | A |
| 151 | 1635.0 | 50.0 | 0.19 | A |
| 152 | 1620 | 48 | 0.34 | A |
| 155 | 1673 | 47 | 0.20 | A |
| 156 | 1580 | 40 | 0.75 | A |
| 157 | 1470.14 | 14.04 | 1.87 | A |
| 162 | 1657.8 | 74.7 | 0.04 | A |
| 166 | 1602.6 | 87.0 | 0.52 | A |
| 171 | 1645.80 | 165.00 | 0.08 | A |
| 173 | 670 | 40 | 10.02 | N |
| 176 | 1650 | 129 | 0.04 | A |
| 180 | 1774.1 | 15.6 | 1.23 | A |
| 182 | 1773.82 | 9.61 | 1.22 | A |
| 184 | 1581.0 | 17.6 | 0.74 | A |
| 185 | 1696 | 85 | 0.43 | A |
| 186 | 1611 | 93 | 0.43 | A |
| 188 | 1920 | 200 | 2.71 | W |
| 189 | 1652.4 | 95.1 | 0.01 | A |
| 195 | 1602 | 33.8 | 0.53 | A |
| 197 | 1600 | 56 | 0.55 | A |
| 200 | 1721.99 | 56.72 | 0.70 | A |
| 201 | 1719 | 172 | 0.67 | A |
| 210 | 1659 | 99 | 0.05 | A |
| 211 | 1633 | 49.0 | 0.21 | A |
| 212 | 1734 | 39 | 0.82 | A |
| 213 | 1688 | 51 | 0.35 | A |
| 215 | 1567 | 39.00 | 0.88 | A |

TABLE 40. IC Evaluation Results for H-3 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

TABLE 40. IC Evaluation Results for H-3 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 216 | 1682 | 73.2 | 0.29 | A |
| 220 | 1642 | 21 | 0.12 | A |
| 225 | 1670 | 70 | 0.17 | A |
| 227 | 1680 | 80 | 0.27 | A |
| 230 | 1699 | 24 | 0.46 | A |
| 233 | 1564 | 9.28 | 0.91 | A |
| 235 | 3571.96 | 27.39 | 19.54 | N |
| 238 | 1862.28 | 152.6 | 2.13 | W |
| 240 | 1673 | 29 | 0.20 | A |
| 242 | 1840 | 60 | 1.90 | A |
| 250 | 1680 | 60 | 0.27 | A |
| 251 | 1694 | 20 | 0.41 | A |
| 252 | 1614.15 | 34.28 | 0.40 | A |
| 253 | 1759 | 126 | 1.07 | A |
| 254 | 1750 | 120 | 0.98 | A |
| 256 | 1631 | 28 | 0.23 | A |
| 261 | 1642.6 | 65.5 | 0.11 | A |
| 266 | 1712 | 35 | 0.59 | A |
| 268 | 1900 | 228 | 2.51 | W |
| 283 | 1650 | 74 | 0.04 | A |
| 284 | 1932.34 | 17.95 | 2.84 | W |
| 287 | 14.372 | 0.04 | 16.69 | N |
| 288 | 1680 | 70 | 0.27 | A |
| 289 | 1700 | 50 | 0.47 | A |
| 291 | 1693 | 54 | 0.40 | A |
| 304 | 32.0 | 1.43 | 16.51 | N |
| 305 | 1089 |  | 5.75 | N |
| 306 | 1285.53 | 39.21 | 3.75 | N |
| 310 | 1713 | 60 | 0.60 | A |
| 313 | 1551 | 40 | 1.04 | A |
| 317 | 1679 |  | 0.26 | A |
| 318 | 1680 | 80 | 0.27 | A |
| 320 | 1684.72 | 175.88 | 0.32 | A |
| 322 | 1730 | 237 | 0.78 | A |
| 324 | 1605.07 | 8.25 | 0.49 | A |

## Reported Results for Eu-152 Determination in Sample 2

TABLE 41. Target values

| Parameter | Value |
| --- | --- |
| Robust Mean, Bq/kg | 3.01 |
| Robust SD, Bq/kg | 0.37 |
| MARB, % | 30 |

FIGURE 36. Reported results and evaluation parameters for Eu-152 in Sample 2.

TABLE 42. IC Evaluation Results for Eu-152 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 2 | 3.6 | 0.5 | 1.59 | A |
| 3 | 3.1 | 0.3 | 0.24 | A |
| 4 | 2.84 | 0.14 | 0.46 | A |
| 7 | 2.79 | 0.11 | 0.59 | A |
| 9 | 3.3 | 0.7 | 0.78 | A |
| 10 | 2.66 | 0.27 | 0.95 | A |
| 11 | 2.77 | 0.18 | 0.65 | A |
| 14 | 2.9 | 0.6 | 0.30 | A |
| 15 | 3.3 | 0.4 | 0.78 | A |
| 16 | 2.7 | 0.4 | 0.84 | A |
| 18 | 2.97 | 0.33 | 0.11 | A |
| 20 | 3.10 | 0.13 | 0.24 | A |
| 21 | 2.80 | 0.34 | 0.57 | A |
| 25 | 4.5 | 0.9 | 4.03 | N |
| 29 | 3.29 | 0.32 | 0.76 | A |
| 31 | 3.00 | 0.35 | 0.03 | A |
| 32 | 3.2 | 0.6 | 0.51 | A |
| 34 | 3.4 | 0.8 | 1.05 | A |
| 36 | 3.3 | 0.2 | 0.78 | A |
| 38 | 3.46 | 0.33 | 1.22 | A |
| 39 | 3.6 | 0.2 | 1.59 | A |

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 41 | 3.1 | 0.2 | 0.24 | A |
| 42 | 2.86 | 0.57 | 0.41 | A |
| 50 | 3.34 | 1.25 | 0.89 | A |
| 51 | 2.5 | 0.2 | 1.38 | A |
| 52 | 2.35 | 0.25 | 1.78 | A |
| 68 | 2.97 | 0.17 | 0.11 | A |
| 69 | 2.54 | 1.26 | 1.27 | A |
| 70 | 2.97 | 0.77 | 0.11 | A |
| 74 | 3.3 | 0.3 | 0.78 | A |
| 75 | 2.9 | 0.24 | 0.30 | A |
| 76 | 3.02 | 0.38 | 0.03 | A |
| 77 | 4.1 | 0.3 | 2.95 | W |
| 78 | 2.86 | 0.3 | 0.41 | A |
| 81 | 3.4 | 0.3 | 1.05 | A |
| 84 | 2.82 | 0.3 | 0.51 | A |
| 89 | 4.6 | 0.2 | 4.30 | N |
| 93 | 2.55 | 0.15 | 1.24 | A |
| 94 | 2.41 | 0.10 | 1.62 | A |
| 95 | 2.49 | 0.11 | 1.41 | A |
| 98 | 3.3 | 0.4 | 0.78 | A |
| 101 | 3.01 | 0.28 | 0.00 | A |
| 102 | 2.63 | 0.26 | 1.03 | A |
| 106 | 2.8 | 0.2 | 0.57 | A |
| 107 | 2.87 | 0.10 | 0.38 | A |
| 109 | 3.2 | 0.3 | 0.51 | A |
| 113 | 2.1 | 0.5 | 2.46 | W |
| 120 | 3.0 | 0.3 | 0.03 | A |
| 121 | 2.88 | 0.48 | 0.35 | A |
| 123 | 2.7 | 0.3 | 0.84 | A |
| 124 | 3.0 | 0.4 | 0.03 | A |
| 126 | 2.9 | 0.4 | 0.30 | A |
| 129 | 3.1 | 0.2 | 0.24 | A |
| 130 | 3.1 | 0.6 | 0.24 | A |
| 133 | 2.95 | 0.26 | 0.16 | A |
| 134 | 2.21 | 0.18 | 2.16 | W |
| 135 | 3.01 | 0.19 | 0.00 | A |
| 137 | 2.448 | 0.401 | 1.52 | A |
| 142 | 97.8 | 6.8 | 256.19 | N |
| 144 | 3.05 | 0.6 | 0.11 | A |
| 147 | 3.53 | 0.46 | 1.41 | A |
| 148 | 2.74 | 0.51 | 0.73 | A |
| 150 | 2.67 | 0.29 | 0.92 | A |
| 154 | 3.9 | 1.5 | 2.41 | W |
| 156 | 3.32 | 0.27 | 0.84 | A |
| 157 | 3.19 | 0.57 | 0.49 | A |
| 158 | 3.0 | 0.4 | 0.03 | A |
| 159 | 2.6 | 0.2 | 1.11 | A |
| 160 | 2.4 | 0.1 | 1.65 | A |
| 161 | 3.0 | 0.7 | 0.03 | A |
| 162 | 3.12 | 0.47 | 0.30 | A |
| 163 | 2.7 | 0.2 | 0.84 | A |
| 164 | 3.4 | 0.2 | 1.05 | A |
| 165 | 3.3 | 1.5 | 0.78 | A |
| 166 | 3.3 | 1.0 | 0.78 | A |

TABLE 42. IC Evaluation Results for Eu-152 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

TABLE 42. IC Evaluation Results for Eu-152 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 169 | 2.76 | 0.41 | 0.68 | A |
| 171 | 2.80 | 0.36 | 0.57 | A |
| 172 | 2.48 | 0.36 | 1.43 | A |
| 177 | 5 | 0.5 | 5.38 | N |
| 179 | 2.9 | 0.2 | 0.30 | A |
| 180 | 3.67 | 0.27 | 1.78 | A |
| 181 | 2.614 | 0.149 | 1.07 | A |
| 182 | 2.87 | 0.89 | 0.38 | A |
| 185 | 3.25 | 0.3 | 0.65 | A |
| 191 | 2.98 | 0.21 | 0.08 | A |
| 193 | 3.04 | 0.3 | 0.08 | A |
| 195 | 3.29 | 0.35 | 0.76 | A |
| 196 | 3.10 | 0.44 | 0.24 | A |
| 197 | 2.91 | 0.24 | 0.27 | A |
| 199 | 3.22 | 0.25 | 0.57 | A |
| 205 | 3.1 | 0.3 | 0.24 | A |
| 209 | 2.67 | 0.61 | 0.92 | A |
| 212 | 2.74 | 0.28 | 0.73 | A |
| 219 | 3.49 | 1.3 | 1.30 | A |
| 222 | 2.67 | 0.46 | 0.92 | A |
| 223 | 3.29 | 0.40 | 0.76 | A |
| 224 | 3.3 | 0.3 | 0.78 | A |
| 225 | 2.76 | 0.23 | 0.68 | A |
| 230 | 3.34 | 0.2 | 0.89 | A |
| 231 | 5.051 | 1.152 | 5.52 | N |
| 232 | 3.0163 | 0.2108 | 0.02 | A |
| 233 | 2.43 | 0.10 | 1.57 | A |
| 238 | 2.98 | 0.39 | 0.08 | A |
| 239 | 3.02 | 0.22 | 0.03 | A |
| 240 | 5.1 | 1.3 | 5.65 | N |
| 241 | 3.39 | 0.20 | 1.03 | A |
| 245 | 2.67 | 0.11 | 0.92 | A |
| 246 | 2.33 | 0.167 | 1.84 | A |
| 248 | 3.57 | 0.16 | 1.51 | A |
| 251 | 2.78 | 0.18 | 0.62 | A |
| 254 | 2.78 | 0.29 | 0.62 | A |
| 255 | 3.2 | 1.5 | 0.51 | A |
| 257 | 2.8 | 0.3 | 0.57 | A |
| 261 | 3.53 | 0.62 | 1.41 | A |
| 262 | 3.6 | 0.7 | 1.59 | A |
| 263 | 3.41 | 0.36 | 1.08 | A |
| 266 | 2.8 | 0.1 | 0.57 | A |
| 267 | 4.14 | 0.34 | 3.05 | N |
| 268 | 2.75 | 0.33 | 0.70 | A |
| 270 | 4.16 | 0.764 | 3.11 | N |
| 274 | 2.67 | 0.24 | 0.92 | A |
| 275 | 3.38 | 0.44 | 1.00 | A |
| 279 | 3.0 | 0.4 | 0.03 | A |
| 281 | 2.7 | 0.4 | 0.84 | A |
| 283 | 2.55 | 0.36 | 1.24 | A |
| 287 | 4.2 | 0.9 | 3.22 | N |
| 288 | 2.90 | 0.25 | 0.30 | A |
| 291 | 3.03 | 0.33 | 0.05 | A |
| 297 | 3.9 | 0.5 | 2.41 | W |

TABLE 42. IC Evaluation Results for Eu-152 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 299 | 3.0 | 0.2 | 0.03 | A |
| 305 | 2.55 | 0.51 | 1.24 | A |
| 307 | 2.04 | 1.07 | 2.62 | W |
| 312 | 4.28 | 0.37 | 3.43 | N |
| 313 | 3.52 | 0.39 | 1.38 | A |
| 314 | 2.9 | 0.29 | 0.30 | A |
| 317 | 2.83 |  | 0.49 | A |
| 320 | 2.76 | 0.37 | 0.68 | A |
| 321 | 2.94 | 0.71 | 0.19 | A |
| 322 | 2.7 | 0.5 | 0.84 | A |
| 323 | 2.6 | 0.3 | 1.11 | A |
| 324 | 2.77 | 0.13 | 0.65 | A |
| 325 | 2.66 | 0.32 | 0.95 | A |

## Reported Results for Gross Alpha Determination in Sample 5

TABLE 43. Target values

| Parameter | Value |
| --- | --- |
| Robust Mean, Bq/kg | 16.4 |
| Robust SD, Bq/kg | 5.4 |
| MARB, % | 60 |

FIGURE 37. Reported results and evaluation parameters for Gross Alpha in Sample 5.

TABLE 44. IC Evaluation Results for Gross Alpha in Sample 5, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 1 | 3.34 | 0.15 | 2.42 | W |
| 2 | 14.0 | 0.4 | 0.44 | A |
| 6 | 14.8 | 0.7 | 0.30 | A |
| 8 | 18.08 | 0.57 | 0.31 | A |
| 9 | 20.4 | 2.0 | 0.74 | A |
| 10 | 19.2 | 1.7 | 0.52 | A |
| 14 | 19.3 | 0.6 | 0.54 | A |
| 15 | 18.0 | 1.3 | 0.30 | A |
| 18 | 19.8 | 1.5 | 0.63 | A |
| 25 | 17.1 | 1.2 | 0.13 | A |
| 26 | 18.11 | 1.09 | 0.32 | A |
| 30 | 20.29 | 2.34 | 0.72 | A |
| 32 | 20.4 | 2.0 | 0.74 | A |
| 34 | 15.4 | 1.4 | 0.19 | A |
| 35 | 14.2 | 1.3 | 0.41 | A |
| 39 | 15.2 | 0.5 | 0.22 | A |
| 41 | 15.8 | 1.3 | 0.11 | A |
| 45 | 12.5 | 2.0 | 0.72 | A |
| 48 | 19.6 | 0.9 | 0.59 | A |
| 50 | 20.92 | 3.22 | 0.84 | A |
| 51 | 17.5 | 2 | 0.20 | A |

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 58 | 12.85 | 1.42 | 0.66 | A |
| 59 | 22.5 | 2.16 | 1.13 | A |
| 60 | 16.2 | 5.9 | 0.04 | A |
| 61 | 11.44 | 0.96 | 0.92 | A |
| 63 | 21.6 | 0.4 | 0.96 | A |
| 66 | 38.8 | 0.5 | 4.15 | N |
| 69 | 27.82 | 1.46 | 2.11 | W |
| 73 | 17.43 | 0.46 | 0.19 | A |
| 76 | 13.60 | 1.08 | 0.52 | A |
| 78 | 15.8 | 0.28 | 0.11 | A |
| 79 | 19.3 | 1.5 | 0.54 | A |
| 84 | 9.25 | 2 | 1.32 | A |
| 88 | 15.1 | 1.6 | 0.24 | A |
| 90 | 4.97 | 0.8 | 2.12 | W |
| 98 | 19.4 | 1.6 | 0.56 | A |
| 99 | 11.6 | 1.9 | 0.89 | A |
| 100 | 70.83 | 37.6 | 10.08 | N |
| 101 | 20 | 2 | 0.67 | A |
| 102 | 12.5 | 0.3 | 0.72 | A |
| 103 | 22.92 | 3.44 | 1.21 | A |
| 108 | 3.2126 | 0.2003 | 2.44 | W |
| 111 | 25.4 | 9.0 | 1.67 | A |
| 112 | 6.5530 | 0.3052 | 1.82 | A |
| 113 | 13.29 | 0.35 | 0.58 | A |
| 116 | 10.8 | 1.7 | 1.04 | A |
| 117 | 11.35 | 1.46 | 0.94 | A |
| 120 | 14.7 | 0.9 | 0.31 | A |
| 123 | 15.6 | 0.9 | 0.15 | A |
| 125 | 10 | 0.4 | 1.19 | A |
| 135 | 21 | 0.5 | 0.85 | A |
| 136 | 20.94 | 3.81 | 0.84 | A |
| 137 | 16 | 1.1 | 0.07 | A |
| 148 | 10.06 | 0.334 | 1.17 | A |
| 150 | 12.4 | 0.3 | 0.74 | A |
| 151 | 11.0 | 1.0 | 1.00 | A |
| 153 | 27.5 | 2.5 | 2.06 | W |
| 154 | 18.1 | 3.7 | 0.31 | A |
| 156 | 28 | 4 | 2.15 | W |
| 157 | 6.84 | 0.56 | 1.77 | A |
| 164 | 20.0 | 1.2 | 0.67 | A |
| 169 | 14.77 | 1.12 | 0.30 | A |
| 171 | 12.80 | 1.66 | 0.67 | A |
| 172 | 10.4 | 2.9 | 1.11 | A |
| 180 | 9.4 | 0.8 | 1.30 | A |
| 181 | 13.81 | 0.20 | 0.48 | A |
| 182 | 9.43 | 0.07 | 1.29 | A |
| 183 | 5.18 | 0.01 | 2.08 | W |
| 186 | 21.3 | 2.2 | 0.91 | A |
| 189 | 6.7 | 1.9 | 1.80 | A |
| 193 | 4.46 | 0.9 | 2.21 | W |
| 195 | 16.3 | 4.60 | 0.02 | A |
| 197 | 15.8 | 1.1 | 0.11 | A |
| 207 | 20.1 | 2.1 | 0.69 | A |
| 208 | 19.0 | 0.036 | 0.48 | A |

TABLE 44. IC Evaluation Results for Gross Alpha in Sample 5, Rep. Value and Rep. Unc. in Bq/kg

TABLE 44. IC Evaluation Results for Gross Alpha in Sample 5, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 210 | 20.3 | 1.5 | 0.72 | A |
| 212 | 21.2 | 1.0 | 0.89 | A |
| 213 | 19.6 | 1.6 | 0.59 | A |
| 215 | 12.00 | 0.30 | 0.81 | A |
| 219 | 23.5 | 1.88 | 1.31 | A |
| 221 | 0.056 | 0.01 | 3.03 | N |
| 224 | 10 | 1 | 1.19 | A |
| 227 | 14.68 | 0.24 | 0.32 | A |
| 228 | 17.9 | 0.87 | 0.28 | A |
| 230 | 12.65 | 2.0 | 0.69 | A |
| 233 | 17.37 | 1.60 | 0.18 | A |
| 235 | 46.66 | 1.65 | 5.60 | N |
| 238 | 5.9 | 0.738 | 1.94 | A |
| 240 | 14.9 | 0.6 | 0.28 | A |
| 242 | 17.5 | 2.6 | 0.20 | A |
| 243 | 11.0 | 1.9 | 1.00 | A |
| 248 | 16.38 | 1.17 | 0.00 | A |
| 250 | 15.1 | 0.4 | 0.24 | A |
| 251 | 12.82 | 0.21 | 0.66 | A |
| 252 | 14.44 | 0.48 | 0.36 | A |
| 253 | 20.6 | 3.1 | 0.78 | A |
| 254 | 17.1 | 4.2 | 0.13 | A |
| 255 | 18.0 | 3.0 | 0.30 | A |
| 256 | 10.990 | 0.218 | 1.00 | A |
| 257 | 22.1 | 1.2 | 1.06 | A |
| 261 | 16.3 | 0.91 | 0.02 | A |
| 266 | 14.9 | 0.6 | 0.28 | A |
| 267 | 3.86 | 0.41 | 2.32 | W |
| 268 | 24.8 | 3.0 | 1.56 | A |
| 269 | 16.2 | 0.5 | 0.04 | A |
| 280 | 22.033 | 1.763 | 1.04 | A |
| 283 | 21 | 2.1 | 0.85 | A |
| 287 | 1.1 | 0.33 | 2.83 | W |
| 291 | 21 | 2 | 0.85 | A |
| 295 | 15.52 | 0.19 | 0.16 | A |
| 305 | 17.5 | 1.80 | 0.20 | A |
| 310 | 35.0 | 5 | 3.44 | N |
| 313 | 28.6 | 1.1 | 2.26 | W |
| 317 | 18 |  | 0.30 | A |
| 320 | 14.88 | 1.79 | 0.28 | A |
| 322 | 14.0 | 3.5 | 0.44 | A |
| 324 | 16.8 | 4.8 | 0.07 | A |

## Reported Results for Gross Beta Determination in Sample 5

TABLE 45. Target values

| Parameter | Value |
| --- | --- |
| Robust Mean, Bq/kg | 33 |
| Robust SD, Bq/kg | 9 |
| MARB, % | 50 |

FIGURE 38. Reported results and evaluation parameters for Gross Beta in Sample 5.

TABLE 46. IC Evaluation Results for Gross Beta in Sample 5, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 1 | 20.97 | 0.42 | 1.34 | A |
| 2 | 35.4 | 0.6 | 0.27 | A |
| 6 | 31.6 | 1.6 | 0.16 | A |
| 8 | 28.55 | 0.24 | 0.49 | A |
| 9 | 65.0 | 8.0 | 3.56 | N |
| 10 | 34.3 | 8.4 | 0.14 | A |
| 14 | 33.7 | 0.7 | 0.08 | A |
| 15 | 30.4 | 2.2 | 0.29 | A |
| 18 | 53.7 | 3.9 | 2.30 | W |
| 25 | 30.2 | 1.7 | 0.31 | A |
| 26 | 29.33 | 1.92 | 0.41 | A |
| 30 | 20.33 | 2.48 | 1.41 | A |
| 32 | 65.0 | 7.8 | 3.56 | N |
| 34 | 37.3 | 3.0 | 0.48 | A |
| 35 | 42.8 | 3.6 | 1.09 | A |
| 39 | 66.7 | 0.9 | 3.74 | N |
| 41 | 30.3 | 1.6 | 0.30 | A |
| 45 | 40 | 3.9 | 0.78 | A |
| 48 | 43.1 | 1.6 | 1.12 | A |
| 50 | 22.86 | 2.32 | 1.13 | A |
| 51 | 30.7 | 3 | 0.26 | A |

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 55 | 29.9 | 3.1 | 0.34 | A |
| 58 | 34.96 | 8.88 | 0.22 | A |
| 59 | 25.5 | 1.33 | 0.83 | A |
| 60 | 15.3 | 5.6 | 1.97 | A |
| 61 | 19.71 | 0.71 | 1.48 | A |
| 63 | 30.3 | 0.6 | 0.30 | A |
| 66 | 56.2 | 1 | 2.58 | W |
| 69 | 41.9 | 2.16 | 0.99 | A |
| 73 | 50.5 | 1.5 | 1.94 | A |
| 76 | 21.8 | 0.8 | 1.24 | A |
| 78 | 34.6 | 0.39 | 0.18 | A |
| 79 | 38.1 | 2.8 | 0.57 | A |
| 84 | 35.21 | 2.5 | 0.25 | A |
| 88 | 31.7 | 1.8 | 0.14 | A |
| 90 | 25.15 | 6.8 | 0.87 | A |
| 98 | 39.0 | 3.7 | 0.67 | A |
| 99 | 45.7 | 8.3 | 1.41 | A |
| 100 | 35.67 | 18.19 | 0.30 | A |
| 101 | 46 | 7 | 1.44 | A |
| 102 | 27.0 | 0.4 | 0.67 | A |
| 103 | 28.99 | 4.35 | 0.45 | A |
| 108 | 16.4285 | 1.4584 | 1.84 | A |
| 111 | 17.4 | 4.0 | 1.73 | A |
| 112 | 27.5146 | 0.9536 | 0.61 | A |
| 113 | 10.56 | 0.39 | 2.49 | W |
| 116 | 27.1 | 3.7 | 0.66 | A |
| 117 | 24.51 | 2.38 | 0.94 | A |
| 120 | 32.3 | 2.8 | 0.08 | A |
| 123 | 36.9 | 0.6 | 0.43 | A |
| 135 | 40 | 1 | 0.78 | A |
| 136 | 31.07 | 0.90 | 0.21 | A |
| 137 | 37 | 1.12 | 0.44 | A |
| 148 | 33.55 | 0.47 | 0.06 | A |
| 150 | 26.2 | 0.4 | 0.76 | A |
| 151 | 23.0 | 3.5 | 1.11 | A |
| 153 | 19.4 | 2.1 | 1.51 | A |
| 154 | 33.9 | 1.5 | 0.10 | A |
| 156 | 37 | 6 | 0.44 | A |
| 157 | 27.93 | 0.99 | 0.56 | A |
| 164 | 33 | 2 | 0.00 | A |
| 169 | 40.22 | 1.83 | 0.80 | A |
| 171 | 37.50 | 4.50 | 0.50 | A |
| 172 | 22.2 | 3.1 | 1.20 | A |
| 181 | 32.35 | 0.16 | 0.07 | A |
| 183 | 43.83 | 2.27 | 1.20 | A |
| 186 | 35.1 | 1.9 | 0.23 | A |
| 189 | 30.0 | 7.7 | 0.33 | A |
| 190 | 50.6 | 10.1 | 1.96 | A |
| 193 | 17.4 | 1.7 | 1.73 | A |
| 195 | 34.4 | 5.86 | 0.16 | A |
| 197 | 19.7 | 1.2 | 1.48 | A |
| 200 | 29.76 | 1.73 | 0.36 | A |
| 201 | 34.9 | 3.5 | 0.21 | A |
| 202 | 26.38 | 1.73 | 0.74 | A |

TABLE 46. IC Evaluation Results for Gross Beta in Sample 5, Rep. Value and Rep. Unc. in Bq/kg

TABLE 46. IC Evaluation Results for Gross Beta in Sample 5, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 207 | 34 | 4 | 0.11 | A |
| 208 | 54.5 | 0.054 | 2.39 | W |
| 210 | 31.5 | 2.5 | 0.17 | A |
| 212 | 33.0 | 1.1 | 0.00 | A |
| 213 | 30.7 | 2.4 | 0.26 | A |
| 215 | 41.00 | 0.04 | 0.89 | A |
| 216 | 28.7 | 1.77 | 0.48 | A |
| 219 | 13.5 | 0.78 | 2.17 | W |
| 221 | 0.3 | 0.03 | 3.63 | N |
| 224 | 32 | 2 | 0.11 | A |
| 227 | 32.09 | 0.33 | 0.10 | A |
| 228 | 21.8 | 0.77 | 1.24 | A |
| 233 | 37.16 | 1.98 | 0.46 | A |
| 235 | 86.66 | 3.07 | 5.96 | N |
| 238 | 61.2 | 7.344 | 3.13 | N |
| 240 | 26.0 | 0.5 | 0.78 | A |
| 242 | 39.3 | 4.7 | 0.70 | A |
| 243 | 27.9 | 2.7 | 0.57 | A |
| 248 | 27.19 | 1.75 | 0.65 | A |
| 250 | 30.2 | 1.0 | 0.31 | A |
| 251 | 32.65 | 0.81 | 0.04 | A |
| 252 | 23.90 | 0.6 | 1.01 | A |
| 253 | 30.2 | 4.6 | 0.31 | A |
| 254 | 53 | 12 | 2.22 | W |
| 255 | 30.0 | 2.0 | 0.33 | A |
| 256 | 33.112 | 0.288 | 0.01 | A |
| 257 | 55.7 | 3.0 | 2.52 | W |
| 261 | 44.8 | 2.7 | 1.31 | A |
| 266 | 30.6 | 0.7 | 0.27 | A |
| 267 | 26.78 | 2.48 | 0.69 | A |
| 268 | 44.6 | 2.3 | 1.29 | A |
| 269 | 20.0 | 0.7 | 1.44 | A |
| 280 | 33.262 | 2.196 | 0.03 | A |
| 283 | 33 | 3.3 | 0.00 | A |
| 287 | 21.7 | 1.47 | 1.26 | A |
| 291 | 33 | 3 | 0.00 | A |
| 295 | 38.07 | 0.34 | 0.56 | A |
| 305 | 79.6 |  | 5.18 | N |
| 310 | 82.9 | 15 | 5.54 | N |
| 313 | 75.8 | 2.8 | 4.76 | N |
| 317 | 35.2 |  | 0.24 | A |
| 320 | 36.78 | 4.41 | 0.42 | A |
| 322 | 46.5 | 11.0 | 1.50 | A |
| 324 | 38.4 | 1.5 | 0.60 | A |

## Reported Results for Pu-239 Determination in Sample 5

TABLE 47. Target values

| Parameter | Value |
| --- | --- |
| Robust Mean, Bq/kg | 5.93 |
| Robust SD, Bq/kg | 2.27 |
| MARB, % | 25 |

FIGURE 39. Reported results and evaluation parameters for Pu-239 in Sample 5.

TABLE 48. IC Evaluation Results for Pu-239 in Sample 5, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 5 | 2.75 | 0.3 | 1.40 | A |
| 9 | 8.9 | 1.6 | 1.31 | A |
| 18 | 3.01 | 0.46 | 1.29 | A |
| 26 | 2.24 | 0.02 | 1.63 | A |
| 32 | 8.9 | 1.6 | 1.31 | A |
| 34 | 5.84 | 0.61 | 0.04 | A |
| 39 | 6.1 | 0.3 | 0.07 | A |
| 41 | 7.5 | 0.4 | 0.69 | A |
| 50 | 6.353 | 0.255 | 0.19 | A |
| 76 | 2.358 | 0.178 | 1.57 | A |
| 81 | 4.6 | 0.6 | 0.59 | A |
| 84 | 2.89 | 0.23 | 1.34 | A |
| 86 | 2.0 | 0.25 | 1.73 | A |
| 91 | 7.48 | 0.64 | 0.68 | A |
| 92 | 7.3 | 0.7 | 0.60 | A |
| 99 | 8.12 | 0.5 | 0.96 | A |
| 125 | 5.21 | 0.31 | 0.32 | A |
| 132 | 6.04 | 0.18 | 0.05 | A |
| 135 | 10.9 | 0.3 | 2.19 | W |
| 137 | 1.6665 | 0.1467 | 1.88 | A |
| 148 | 4.81 | 0.756 | 0.49 | A |

TABLE 48. IC Evaluation Results for Pu-239 in Sample 5, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 151 | 2.4 | 0.4 | 1.56 | A |
| 156 | 6.0 | 0.4 | 0.03 | A |
| 161 | 6.3 | 0.2 | 0.16 | A |
| 169 | 0.012 | 0.002 | 2.61 | W |
| 172 | 5.91 | 1.12 | 0.01 | A |
| 180 | 7.41 | 0.3 | 0.65 | A |
| 182 | 7.39 | 0.16 | 0.64 | A |
| 186 | 8.35 | 0.31 | 1.07 | A |
| 189 | 5.01 | 0.24 | 0.41 | A |
| 216 | 7.46 | 0.45 | 0.67 | A |
| 219 | 6.52 | 0.33 | 0.26 | A |
| 227 | 5.09 | 0.18 | 0.37 | A |
| 230 | 7.5 | 0.24 | 0.69 | A |
| 233 | 4.8520 | 0.2471 | 0.47 | A |
| 239 | 7.85 | 0.27 | 0.85 | A |
| 240 | 5.9 | 0.2 | 0.01 | A |
| 242 | 5.89 | 0.17 | 0.02 | A |
| 243 | 3.2 | 0.21 | 1.20 | A |
| 251 | 4.63 | 0.13 | 0.57 | A |
| 254 | 7.34 | 0.45 | 0.62 | A |
| 261 | 6.29 | 0.23 | 0.16 | A |
| 268 | 1.5 | 0.1 | 1.95 | A |
| 283 | 8.8 | 0.31 | 1.26 | A |
| 297 | 4.56 | 0.33 | 0.60 | A |
| 310 | 0.00764 | 0.00049 | 2.61 | W |
| 322 | 4.5 | 1.2 | 0.63 | A |
| 323 | 7.89 | 0.08 | 0.86 | A |
| 323 | 7.89 | 0.08 | 0.86 | A |
| 324 | 6.44 | 0.39 | 0.22 | A |

## Reported Results for Am-241 Determination in Sample 5

TABLE 49. Target values

| Parameter | Value |
| --- | --- |
| Robust Mean, Bq/kg | 4.85 |
| Robust SD, Bq/kg | 0.57 |
| MARB, % | 30 |

FIGURE 40. Reported results and evaluation parameters for Am-241 in Sample 5.

TABLE 50. IC Evaluation Results for Am-241 in Sample 5, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 2 | 5.5 | 0.6 | 1.14 | A |
| 3 | 4.6 | 0.5 | 0.44 | A |
| 4 | 2.33 | 0.19 | 4.42 | N |
| 6 | 5.9 | 0.7 | 1.84 | A |
| 7 | 4.24 | 0.37 | 1.07 | A |
| 8 | 4.80 | 0.18 | 0.09 | A |
| 9 | 5.5 | 0.9 | 1.14 | A |
| 10 | 4.23 | 0.41 | 1.09 | A |
| 11 | 4.73 | 0.30 | 0.21 | A |
| 12 | 6.62 | 0.5 | 3.11 | N |
| 13 | 4.66 | 0.36 | 0.33 | A |
| 14 | 5.0 | 0.5 | 0.26 | A |
| 15 | 5.8 | 1.4 | 1.67 | A |
| 16 | 6.2 | 1.4 | 2.37 | W |
| 18 | 5.20 | 0.60 | 0.61 | A |
| 19 | 4.91 | 0.36 | 0.11 | A |
| 20 | 5.59 | 0.28 | 1.30 | A |
| 21 | 4.34 | 0.66 | 0.89 | A |
| 22 | 4.3 | 0.3 | 0.96 | A |
| 25 | 5.1 | 0.4 | 0.44 | A |
| 26 | 3.31 | 1.00 | 2.70 | W |

TABLE 50. IC Evaluation Results for Am-241 in Sample 5, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 31 | 5.57 | 0.61 | 1.26 | A |
| 32 | 5.3 | 0.8 | 0.79 | A |
| 34 | 4.8 | 0.6 | 0.09 | A |
| 36 | 4.9 | 0.5 | 0.09 | A |
| 38 | 4.56 | 0.38 | 0.51 | A |
| 39 | 4.3 | 0.3 | 0.96 | A |
| 41 | 4.7 | 0.2 | 0.26 | A |
| 45 | 4.9 | 0.4 | 0.09 | A |
| 46 | 7.33 | 1.25 | 4.35 | N |
| 50 | 3.49 | 2.91 | 2.39 | W |
| 51 | 4.3 | 0.5 | 0.96 | A |
| 52 | 4.48 | 0.20 | 0.65 | A |
| 53 | 4.45 | 0.25 | 0.70 | A |
| 56 | 4.37 | 0.22 | 0.84 | A |
| 59 | 4.8 | 0.4 | 0.09 | A |
| 64 | 5.3 | 0.5 | 0.79 | A |
| 66 | 4.5 | 0.5 | 0.61 | A |
| 68 | 5.2 | 0.4 | 0.61 | A |
| 69 | 4.99 | 1.05 | 0.25 | A |
| 70 | 5.95 | 0.89 | 1.93 | A |
| 74 | 4.8 | 0.4 | 0.09 | A |
| 75 | 4.2 | 0.9 | 1.14 | A |
| 78 | 4.72 | 0.13 | 0.23 | A |
| 80 | 7.41 | 0.71 | 4.49 | N |
| 81 | 4.1 | 0.7 | 1.32 | A |
| 84 | 4.79 | 0.16 | 0.11 | A |
| 85 | 5.66 | 0.66 | 1.42 | A |
| 87 | 4.9 | 0.4 | 0.09 | A |
| 92 | 5.05 | 0.64 | 0.35 | A |
| 93 | 4.70 | 0.18 | 0.26 | A |
| 96 | 4.9 | 0.3 | 0.09 | A |
| 98 | 5.2 | 0.9 | 0.61 | A |
| 99 | 4.84 | 1.3 | 0.02 | A |
| 101 | 4.64 | 0.32 | 0.37 | A |
| 102 | 5.8 | 0.4 | 1.67 | A |
| 103 | 5.06 | 0.17 | 0.37 | A |
| 105 | 4.36 | 0.14 | 0.86 | A |
| 107 | 3.04 | 0.08 | 3.18 | N |
| 109 | 4.8 | 0.4 | 0.09 | A |
| 111 | 6.1 | 0.4 | 2.19 | W |
| 113 | 5.6 | 0.5 | 1.32 | A |
| 116 | 4.9 | 0.4 | 0.09 | A |
| 120 | 4.4 | 0.3 | 0.79 | A |
| 121 | 4.70 | 0.55 | 0.26 | A |
| 123 | 3.9 | 0.3 | 1.67 | A |
| 124 | 4.0 | 0.4 | 1.49 | A |
| 125 | 4.0 | 0.7 | 1.49 | A |
| 126 | 4.37 | 0.45 | 0.84 | A |
| 129 | 4.8 | 0.2 | 0.09 | A |
| 130 | 4.3 | 0.5 | 0.96 | A |
| 133 | 4.62 | 0.38 | 0.40 | A |
| 134 | 37.25 | 7.25 | 56.84 | N |
| 135 | 4.90 | 0.3 | 0.09 | A |
| 137 | 5.26 | 0.337 | 0.72 | A |

TABLE 50. IC Evaluation Results for Am-241 in Sample 5, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 139 | 4.5711 | 0.67289 | 0.49 | A |
| 144 | 5.16 | 0.97 | 0.54 | A |
| 146 | 4.4 | 0.5 | 0.79 | A |
| 147 | 4.47 | 0.49 | 0.67 | A |
| 148 | 4.91 | 0.586 | 0.11 | A |
| 149 | 4.66 | 0.15 | 0.33 | A |
| 150 | 5.60 | 0.28 | 1.32 | A |
| 152 | 4.87 | 0.39 | 0.04 | A |
| 153 | 4.69 | 0.37 | 0.28 | A |
| 154 | 4.3 | 1.8 | 0.96 | A |
| 156 | 4.62 | 0.79 | 0.40 | A |
| 157 | 4.24 | 0.69 | 1.07 | A |
| 160 | 4.36 | 0.25 | 0.86 | A |
| 161 | 5.6 | 1.0 | 1.32 | A |
| 162 | 4.62 | 0.42 | 0.40 | A |
| 163 | 4.2 | 0.5 | 1.14 | A |
| 164 | 4.9 | 0.3 | 0.09 | A |
| 165 | 5.3 | 0.6 | 0.79 | A |
| 166 | 5.6 | 0.9 | 1.32 | A |
| 169 | 4.96 | 0.22 | 0.19 | A |
| 171 | 4.11 | 0.50 | 1.30 | A |
| 172 | 3.5 | 0.7 | 2.37 | W |
| 177 | 35.7 | 5 | 54.12 | N |
| 179 | 4.7 | 0.2 | 0.26 | A |
| 180 | 4.94 | 0.35 | 0.16 | A |
| 181 | 4.611 | 0.303 | 0.42 | A |
| 182 | 6.67 | 0.63 | 3.19 | N |
| 183 | 5.51 | 0.2 | 1.16 | A |
| 184 | 5.39 | 0.20 | 0.95 | A |
| 185 | 4.57 | 0.42 | 0.49 | A |
| 186 | 5.73 | 0.63 | 1.54 | A |
| 187 | 124.4 | 0.3 | 209.74 | N |
| 189 | 3.2 | 0.9 | 2.89 | W |
| 190 | 4.9 | 0.5 | 0.09 | A |
| 191 | 4.95 | 0.24 | 0.18 | A |
| 193 | 4.3 | 0.77 | 0.96 | A |
| 195 | 6.25 | 0.50 | 2.46 | W |
| 196 | 4.85 | 0.37 | 0.00 | A |
| 197 | 4.57 | 0.27 | 0.49 | A |
| 198 | 5.52 | 0.87 | 1.18 | A |
| 199 | 7.07 | 0.30 | 3.89 | N |
| 205 | 4.9 | 0.3 | 0.09 | A |
| 207 | 4.33 | 0.38 | 0.91 | A |
| 209 | 4.9 | 0.49 | 0.09 | A |
| 211 | 4.82 | 0.44 | 0.05 | A |
| 212 | 5.52 | 0.27 | 1.18 | A |
| 213 | 5.2 | 0.4 | 0.61 | A |
| 219 | 4.58 | 2.8 | 0.47 | A |
| 223 | 4.65 | 0.45 | 0.35 | A |
| 224 | 4.8 | 0.4 | 0.09 | A |
| 225 | 5.1 | 0.3 | 0.44 | A |
| 227 | 4.73 | 0.15 | 0.21 | A |
| 228 | 4.6 | 0.3 | 0.44 | A |
| 230 | 4.995 | 0.28 | 0.25 | A |

TABLE 50. IC Evaluation Results for Am-241 in Sample 5, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 231 | 5.651 | 1.568 | 1.41 | A |
| 232 | 5.4349 | 0.3340 | 1.03 | A |
| 233 | 4.78 | 0.13 | 0.12 | A |
| 234 | 5.41 | 0.43 | 0.98 | A |
| 237 | 4.6 | 0.5 | 0.44 | A |
| 238 | 3.5 | 0.45 | 2.37 | W |
| 239 | 4.37 | 0.32 | 0.84 | A |
| 240 | 4.7 | 0.3 | 0.26 | A |
| 241 | 6.12 | 0.62 | 2.23 | W |
| 242 | 5.1 | 0.3 | 0.44 | A |
| 246 | 2.58 | 0.338 | 3.98 | N |
| 248 | 4.71 | 0.36 | 0.25 | A |
| 249 | 4.21 | 0.36 | 1.12 | A |
| 250 | 4.4 | 0.2 | 0.79 | A |
| 251 | 4.3 | 1.1 | 0.96 | A |
| 254 | 5.07 | 0.45 | 0.39 | A |
| 255 | 4.2 | 0.5 | 1.14 | A |
| 256 | 4.61 | 0.19 | 0.42 | A |
| 257 | 6.6 | 0.4 | 3.07 | N |
| 261 | 4.79 | 0.69 | 0.11 | A |
| 262 | 6.1 | 0.6 | 2.19 | W |
| 263 | 4.88 | 0.63 | 0.05 | A |
| 266 | 5.1 | 0.6 | 0.44 | A |
| 267 | 5.37 | 0.52 | 0.91 | A |
| 268 | 4.59 | 0.23 | 0.46 | A |
| 270 | 4.90 | 0.449 | 0.09 | A |
| 274 | 4.5 | 0.4 | 0.61 | A |
| 275 | 4.44 | 0.35 | 0.72 | A |
| 279 | 5.2 | 0.6 | 0.61 | A |
| 281 | 4.3 | 0.2 | 0.96 | A |
| 283 | 5.48 | 0.38 | 1.11 | A |
| 287 | 5.42 | 0.2 | 1.00 | A |
| 289 | 5.0 | 0.5 | 0.26 | A |
| 290 | 4.36 | 0.25 | 0.86 | A |
| 291 | 5.43 | 0.29 | 1.02 | A |
| 299 | 2.9 | 0.2 | 3.42 | N |
| 305 | 4.45 | 0.49 | 0.70 | A |
| 307 | 4.68 | 0.5 | 0.30 | A |
| 310 | 5.06 | 0.3 | 0.37 | A |
| 312 | 4.25 | 0.15 | 1.05 | A |
| 313 | 5.1 | 0.28 | 0.44 | A |
| 314 | 5.19 | 0.33 | 0.60 | A |
| 317 | 5.05 |  | 0.35 | A |
| 320 | 3.82 | 1.24 | 1.81 | A |
| 321 | 5.2 | 0.36 | 0.61 | A |
| 322 | 3.9 | 0.5 | 1.67 | A |
| 323 | 5.50 | 0.87 | 1.14 | A |
| 324 | 5.16 | 0.09 | 0.54 | A |
| 325 | 5.1 | 0.5 | 0.44 | A |

## Reported Results for Cm-244 Determination in Sample 5

TABLE 51. Target values

| Parameter | Value |
| --- | --- |
| Robust Mean, Bq/kg | 7.02 |
| Robust SD, Bq/kg | 2 |
| MARB, % | 25 |

FIGURE 41. Reported results and evaluation parameters for Cm-244 in Sample 5.

TABLE 52. IC Evaluation Results for Cm-244 in Sample 5, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 5 | 8.59 | 0.9 | 0.79 | A |
| 9 | 6.5 | 1.1 | 0.26 | A |
| 18 | 7.77 | 0.61 | 0.38 | A |
| 32 | 6.5 | 1.1 | 0.26 | A |
| 41 | 8.1 | 0.5 | 0.54 | A |
| 50 | 8.635 | 1.822 | 0.81 | A |
| 76 | 6.2496 | 0.1717 | 0.39 | A |
| 81 | 5.2 | 0.5 | 0.91 | A |
| 84 | 5.15 | 0.7 | 0.93 | A |
| 86 | 8.1 | 1 | 0.54 | A |
| 91 | 3.6 | 0.7 | 1.71 | A |
| 99 | 7.96 | 0.5 | 0.47 | A |
| 125 | 4.58 | 0.27 | 1.22 | A |
| 132 | 4.06 | 0.14 | 1.48 | A |
| 135 | 9.2 | 0.3 | 1.09 | A |
| 148 | 5.10 | 0.573 | 0.96 | A |
| 156 | 9.3 | 1.0 | 1.14 | A |
| 161 | 11.4 | 0.6 | 2.19 | W |
| 180 | 1.53 | 0.16 | 2.74 | W |
| 182 | 1.56 | 0.08 | 2.73 | W |
| 186 | 8.77 | 0.34 | 0.88 | A |

TABLE 52. IC Evaluation Results for Cm-244 in Sample 5, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 189 | 5.07 | 0.24 | 0.97 | A |
| 227 | 7.19 | 0.34 | 0.09 | A |
| 239 | 7.09 | 0.21 | 0.04 | A |
| 240 | 7.2 | 0.2 | 0.09 | A |
| 242 | 7.46 | 0.14 | 0.22 | A |
| 251 | 4.35 | 0.12 | 1.34 | A |
| 254 | 6.49 | 0.46 | 0.26 | A |
| 261 | 8.09 | 0.34 | 0.54 | A |
| 268 | 6.4 | 0.3 | 0.31 | A |
| 283 | 4.01 | 0.2 | 1.50 | A |
| 297 | 7.77 | 0.69 | 0.38 | A |
| 322 | 5.3 | 1.5 | 0.86 | A |

## Reported Results for Gross Alpha Determination in Sample 7

TABLE 53. Target values

| Parameter | Value |
| --- | --- |
| Robust Mean, Bq/sample | 7.64 |
| Robust SD, Bq/sample | 1.34 |
| MARB, % | 40 |

FIGURE 42. Reported results and evaluation parameters for Gross Alpha in Sample 7.

TABLE 54. IC Evaluation Results for Gross Alpha in Sample 7, Rep. Value and Rep. Unc. in Bq/sample

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 1 | 10.41 | 0.31 | 2.07 | W |
| 5 | 8.02 | 0.8 | 0.28 | A |
| 6 | 7.5 | 1.7 | 0.10 | A |
| 8 | 6.951 | 0.4733 | 0.51 | A |
| 9 | 6.9 | 0.7 | 0.55 | A |
| 10 | 6.36 | 0.72 | 0.96 | A |
| 14 | 6.5 | 0.3 | 0.85 | A |
| 18 | 6.56 | 0.46 | 0.81 | A |
| 19 | 7.02 | 0.56 | 0.46 | A |
| 25 | 7.1 | 0.7 | 0.40 | A |
| 26 | 7.84 | 0.05 | 0.15 | A |
| 27 | 6.2 | 0.8 | 1.07 | A |
| 31 | 15.22 | 0.027 | 5.66 | N |
| 34 | 7.0 | 0.7 | 0.48 | A |
| 35 | 9.24 | 1.0 | 1.19 | A |
| 36 | 6.9 | 0.3 | 0.55 | A |
| 41 | 7.3 | 1.5 | 0.25 | A |
| 43 | 6.83 | 0.81 | 0.60 | A |
| 48 | 7.1 | 0.1 | 0.40 | A |
| 50 | 7.36 | 0.62 | 0.21 | A |
| 52 | 23.6 | 2.0 | 11.91 | N |

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 53 | 6.5 | 1.6 | 0.85 | A |
| 58 | 8.37 | 1.26 | 0.54 | A |
| 60 | 10.2 | 2.8 | 1.91 | A |
| 61 | 7.80 | 0.24 | 0.12 | A |
| 63 | 37.42 | 0.58 | 22.22 | N |
| 69 | 5.09 | 0.50 | 1.90 | A |
| 78 | 8.23 | 0.048 | 0.44 | A |
| 84 | 7 | 0.2 | 0.48 | A |
| 88 | 7.31 | 0.36 | 0.25 | A |
| 90 | 6.15 | 1.2 | 1.11 | A |
| 96 | 6.1 | 0.6 | 1.15 | A |
| 98 | 6.74 | 0.40 | 0.67 | A |
| 99 | 7.12 | 0.8 | 0.39 | A |
| 102 | 8.6 | 0.1 | 0.72 | A |
| 103 | 7.04 | 0.57 | 0.45 | A |
| 106 | 6.1 | 0.2 | 1.15 | A |
| 108 | 7.4469 | 0.06702 | 0.14 | A |
| 111 | 9.1 | 1.4 | 1.09 | A |
| 112 | 14.2924 | 0.0212 | 4.96 | N |
| 116 | 7.09 | 0.72 | 0.41 | A |
| 117 | 16.19 | 1.52 | 6.38 | N |
| 120 | 7.2 | 0.4 | 0.33 | A |
| 123 | 16.3 | 0.2 | 6.46 | N |
| 126 | 12.1 | 0.7 | 3.33 | N |
| 128 | 6.1 | 2.7 | 1.15 | A |
| 136 | 9.38 | 0.70 | 1.30 | A |
| 137 | 8.62 | 0.17 | 0.73 | A |
| 146 | 7.1 | 1.4 | 0.40 | A |
| 147 | 6.44 | 0.95 | 0.90 | A |
| 148 | 7.97 | 0.0828 | 0.25 | A |
| 150 | 8.3 | 0.1 | 0.49 | A |
| 154 | 9.4 | 1.4 | 1.31 | A |
| 156 | 7.5 | 1.1 | 0.10 | A |
| 157 | 6.43 | 0.12 | 0.90 | A |
| 162 | 15.43 | 0.13 | 5.81 | N |
| 169 | 7.22 | 0.30 | 0.31 | A |
| 171 | 6.56 | 0.92 | 0.81 | A |
| 172 | 6.9 | 1.3 | 0.55 | A |
| 178 | 6.38 | 0.32 | 0.94 | A |
| 180 | 14.68 | 0.65 | 5.25 | N |
| 181 | 8.306 | 0.043 | 0.50 | A |
| 182 | 14.68 | 0.009 | 5.25 | N |
| 183 | 17.64 | 0.8 | 7.46 | N |
| 186 | 6.95 | 0.45 | 0.51 | A |
| 189 | 7.56 | 0.32 | 0.06 | A |
| 190 | 9.40 | 2.10 | 1.31 | A |
| 193 | 11.4 | 2 | 2.81 | W |
| 195 | 7.24 | 0.26 | 0.30 | A |
| 197 | 7.57 | 0.42 | 0.05 | A |
| 206 | 6.6 | 2.0 | 0.78 | A |
| 208 | 8.71 | 0.47 | 0.80 | A |
| 210 | 7.26 | 0.62 | 0.28 | A |
| 212 | 7.42 | 0.12 | 0.16 | A |
| 213 | 6.97 | 0.51 | 0.50 | A |

TABLE 54. IC Evaluation Results for Gross Alpha in Sample 7, Rep. Value and Rep. Unc. in Bq/sample

TABLE 54. IC Evaluation Results for Gross Alpha in Sample 7, Rep. Value and Rep. Unc. in Bq/sample

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 215 | 7.26 | 0.16 | 0.28 | A |
| 219 | 7.4 | 0.30 | 0.18 | A |
| 221 | 12.7 | 1.05 | 3.78 | N |
| 227 | 6.76 | 0.05 | 0.66 | A |
| 228 | 6.93 | 0.24 | 0.53 | A |
| 233 | 8.14 | 0.21 | 0.37 | A |
| 234 | 7.4 | 0.8 | 0.18 | A |
| 238 | 6.64 | 0.996 | 0.75 | A |
| 240 | 8.5 | 0.3 | 0.64 | A |
| 243 | 5.7 | 0.5 | 1.45 | A |
| 248 | 6.94 | 0.13 | 0.52 | A |
| 249 | 6 | 1 | 1.22 | A |
| 251 | 6.69 | 0.21 | 0.71 | A |
| 254 | 9.3 | 1.9 | 1.24 | A |
| 255 | 8.2 | 0.1 | 0.42 | A |
| 256 | 9.500 | 0.053 | 1.39 | A |
| 257 | 7.5 | 0.4 | 0.10 | A |
| 259 | 7.0 | 0.14 | 0.48 | A |
| 261 | 7.09 | 0.36 | 0.41 | A |
| 267 | 7.28 | 0.37 | 0.27 | A |
| 268 | 7.1 | 0.5 | 0.40 | A |
| 283 | 5.89 | 0.31 | 1.31 | A |
| 287 | 135.45 | 3.66 | 95.38 | N |
| 291 | 7.8 | 0.4 | 0.12 | A |
| 297 | 7.15 | 0.5 | 0.37 | A |
| 306 | 7.716 | 0.668 | 0.06 | A |
| 310 | 8.19 | 0.39 | 0.41 | A |
| 320 | 15.08 | 2.1 | 5.55 | N |
| 322 | 6.6 | 1.1 | 0.78 | A |
| 326 | 7.45 | 0.75 | 0.14 | A |

## Reported Results for Gross Beta Determination in Sample 7

TABLE 55. Target values

| Parameter | Value |
| --- | --- |
| Robust Mean, Bq/sample | 10.18 |
| Robust SD, Bq/sample | 2.69 |
| MARB, % | 40 |

FIGURE 43. Reported results and evaluation parameters for Gross Beta in Sample 7.

TABLE 56. IC Evaluation Results for Gross Beta in Sample 7, Rep. Value and Rep. Unc. in Bq/sample

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 1 | 9.60 | 0.16 | 0.22 | A |
| 5 | 11.89 | 1 | 0.64 | A |
| 6 | 9.5 | 1 | 0.25 | A |
| 8 | 6.298 | 0.685 | 1.44 | A |
| 9 | 14 | 1.1 | 1.42 | A |
| 10 | 5.93 | 0.26 | 1.58 | A |
| 14 | 10.9 | 0.4 | 0.27 | A |
| 18 | 11.67 | 1.29 | 0.55 | A |
| 19 | 10.93 | 0.82 | 0.28 | A |
| 25 | 11.5 | 1.1 | 0.49 | A |
| 26 | 11.20 | 0.05 | 0.38 | A |
| 27 | 10.7 | 1 | 0.19 | A |
| 31 | 16.36 | 0.02 | 2.30 | W |
| 34 | 11.5 | 0.7 | 0.49 | A |
| 35 | 8.93 | 0.63 | 0.46 | A |
| 36 | 11.5 | 0.4 | 0.49 | A |
| 43 | 11.34 | 1.35 | 0.43 | A |
| 48 | 10.8 | 0.12 | 0.23 | A |
| 50 | 11.01 | 0.84 | 0.31 | A |
| 51 | 3.0 | 0.3 | 2.67 | W |
| 52 | 6.84 | 0.82 | 1.24 | A |

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 53 | 10.4 | 1.9 | 0.08 | A |
| 58 | 8.20 | 1.23 | 0.74 | A |
| 60 | 12.7 | 3.5 | 0.94 | A |
| 61 | 4.98 | 0.15 | 1.93 | A |
| 63 | 16.43 | 0.23 | 2.32 | W |
| 69 | 4.96 | 0.50 | 1.94 | A |
| 78 | 9.51 | 0.038 | 0.25 | A |
| 84 | 10.23 | 0.13 | 0.02 | A |
| 88 | 10.64 | 0.75 | 0.17 | A |
| 90 | 6.097 | 0.85 | 1.52 | A |
| 96 | 10.8 | 1.1 | 0.23 | A |
| 98 | 10.5 | 0.7 | 0.12 | A |
| 99 | 7.04 | 0.5 | 1.17 | A |
| 102 | 12.8 | 0.1 | 0.97 | A |
| 103 | 6.06 | 0.89 | 1.53 | A |
| 105 | 10.2 | 1.4 | 0.01 | A |
| 106 | 8.0 | 0.2 | 0.81 | A |
| 108 | 13.90406 | 0.3124 | 1.38 | A |
| 111 | 10.8 | 1.8 | 0.23 | A |
| 112 | 42.4124 | 0.1718 | 11.98 | N |
| 116 | 10.4 | 1.0 | 0.08 | A |
| 117 | 12.87 | 1.46 | 1.00 | A |
| 120 | 14.1 | 0.3 | 1.46 | A |
| 123 | 7.64 | 0.10 | 0.94 | A |
| 126 | 21 | 3.1 | 4.02 | N |
| 128 | 11.0 | 3.8 | 0.30 | A |
| 136 | 6.25 | 0.093 | 1.46 | A |
| 137 | 8.78 | 0.12 | 0.52 | A |
| 146 | 10.8 | 2.1 | 0.23 | A |
| 147 | 11.3 | 1.7 | 0.42 | A |
| 148 | 9.52 | 0.0719 | 0.25 | A |
| 150 | 12.3 | 0.1 | 0.79 | A |
| 154 | 7.4 | 0.2 | 1.03 | A |
| 156 | 8.6 | 1.2 | 0.59 | A |
| 157 | 6.44 | 0.11 | 1.39 | A |
| 160 | 11.6 | 0.3 | 0.53 | A |
| 162 | 8.16 | 0.10 | 0.75 | A |
| 169 | 10.59 | 0.41 | 0.15 | A |
| 171 | 9.70 | 1.30 | 0.18 | A |
| 172 | 11.6 | 1.6 | 0.53 | A |
| 178 | 9.71 | 0.45 | 0.17 | A |
| 180 | 26.0 | 1.1 | 5.88 | N |
| 181 | 6.767 | 0.054 | 1.27 | A |
| 182 | 26.0 | 0.40 | 5.88 | N |
| 183 | 27.6 | 0.8 | 6.48 | N |
| 185 | 10.6 | 0.9 | 0.16 | A |
| 186 | 9.08 | 0.60 | 0.41 | A |
| 189 | 11.59 | 0.46 | 0.52 | A |
| 190 | 14.4 | 3.2 | 1.57 | A |
| 193 | 6.5 | 0.7 | 1.37 | A |
| 195 | 9.32 | 0.24 | 0.32 | A |
| 197 | 10.8 | 0.5 | 0.23 | A |
| 201 | 9.37 | 0.94 | 0.30 | A |
| 202 | 21.61 | 0.60 | 4.25 | N |

TABLE 56. IC Evaluation Results for Gross Beta in Sample 7, Rep. Value and Rep. Unc. in Bq/sample

TABLE 56. IC Evaluation Results for Gross Beta in Sample 7, Rep. Value and Rep. Unc. in Bq/sample

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 206 | 11.4 | 2.1 | 0.45 | A |
| 208 | 9.75 | 0.39 | 0.16 | A |
| 210 | 11.1 | 0.8 | 0.34 | A |
| 212 | 11.0 | 0.2 | 0.30 | A |
| 213 | 10.6 | 0.8 | 0.16 | A |
| 215 | 9.87 | 0.17 | 0.12 | A |
| 216 | 20.38 | 0.58 | 3.79 | N |
| 219 | 11.2 | 0.40 | 0.38 | A |
| 221 | 17.17 | 1.61 | 2.60 | W |
| 223 | 8.42 | 0.43 | 0.65 | A |
| 227 | 6.25 | 0.03 | 1.46 | A |
| 228 | 10.89 | 0.38 | 0.26 | A |
| 233 | 7.69 | 0.20 | 0.93 | A |
| 234 | 11.5 | 1.1 | 0.49 | A |
| 238 | 13.65 | 2.05 | 1.29 | A |
| 240 | 7.1 | 0.2 | 1.14 | A |
| 243 | 6.9 | 0.5 | 1.22 | A |
| 248 | 11.75 | 0.34 | 0.58 | A |
| 249 | 13 | 2 | 1.05 | A |
| 251 | 9.11 | 0.15 | 0.40 | A |
| 254 | 14.8 | 3.1 | 1.72 | A |
| 255 | 8.3 | 0.1 | 0.70 | A |
| 256 | 11.007 | 0.048 | 0.31 | A |
| 257 | 9.6 | 0.5 | 0.22 | A |
| 259 | 9.74 | 0.19 | 0.16 | A |
| 261 | 18.3 | 0.58 | 3.02 | N |
| 267 | 9.89 | 0.64 | 0.11 | A |
| 268 | 11.3 | 0.7 | 0.42 | A |
| 283 | 8.36 | 0.43 | 0.68 | A |
| 287 | 256.6 | 5.11 | 91.61 | N |
| 291 | 11.2 | 0.7 | 0.38 | A |
| 297 | 10.85 | 0.7 | 0.25 | A |
| 310 | 10.59 | 0.54 | 0.15 | A |
| 320 | 18.79 | 2.62 | 3.20 | N |
| 322 | 12.1 | 1.7 | 0.71 | A |
| 326 | 10.6 | 1.1 | 0.16 | A |

# References References

[1] International Organization for Standardization (ISO). (2010). Conformity assessment - General requirements for

proficiency testing, ISO/IEC 17043:2010. Geneva: Switzerland.

[2] International Organization for Standardization (ISO). (2015). Statistical methods for use in proficiency testing by interlaboratory comparison, ISO 13528:2015. Geneva: Switzerland.

[3] OSVATH, I., TARJAN, S., PITOIS, A., GROENING, M., OSBORN, D., IAEA’s ALMERA network: Support-

ing the quality of environmental radioactivity measurements, Applied Radiation and Isotopes (2015) DOI: 10.1016/j.arpasido.2015.12.062

# Contributors to drafting and review

S. Patterson, B. Seslak, A. Mauring, S. Tarjan, A. Trinkl: Terrestrial Environment Laboratory, NA Environment Labora- tories