IAEA-TERC-2023-01 World Wide Open Proficiency Test Exercise, Pie-charts, S-Shapes and Reported Results with Scores

Terrestrial Environmental Radiochemistry Laboratory, International Atomic Energy Agency

November 23, 2023

Contents

- Description of Samples	1

  - Sample 1, 2 and 3, Spiked Water	2

  - Sample 4, Japanese Soil	2

  - Samples 5, 6 and 7, Simulated contaminated surface samples	3

- Evaluation of results	4

  - Relative bias (Samples 1, 2, 4, 5, 6 and 7)	4

  - Evaluation approach for radionuclide activity concentration values (Samples 1, 2 and 4) and activity values per sample (Samples 5, 6 and 7)	4

  - Evaluation parameters calculated for intercomparison analyte values	4

2.3.1 Z-Score	4

- Data Evaluation	5

  - Evaluation of activity concentrations of radionuclides	5

  - Display of intercomparison values	121

- References	142

- Contributors to evaluation and report	142

- Disclaimer	142

Abstract

The IAEA Terrestrial Environmental Radiochemistry Laboratory (TERC) provided a Proficiency Test (PT) on determination of anthropogenic and natural radionuclides in water, soil and simulated contaminated surface samples. This year’s PT was designed to monitor and demonstrate the performance and analytical capability of the participating laboratories and to identify gaps and problem areas where further development is needed. The sample set was formulated to support these goals.

# 1	Description of Samples

- Sample 1: Spiked Water, anthropogenic and natural radionuclides, gross alpha and beta (GAB)

- Sample 2: Spiked Water, anthropogenic and natural radionuclides and GAB (environmental levels)

- Sample 3: Spiked Water, quality control sample, no data reporting and evaluation

- Sample 4: Japanese Soil, anthropogenic and natural radionuclides

- Sample 5: Simulated contaminated surface sample, gamma and beta emitting radionuclides

- Sample 6: Simulated contaminated surface sample, gamma emitting radionuclide (Cs-134)

- Sample 7: Simulated contaminated surface sample, beta emitting radionuclide (Sr-90)

## Sample 1, 2 and 3, Spiked Water

Matrix origin:

Drinking water sourced from Seibersdorf, Austria. Sample preparation:

The raw water was gravimetrically spiked with known amounts of a prepared standard solution, containing a mixture of certified radionuclides and acidified to < pH 2 ( 0.05 M HNO3) for stability. The stability of the samples has been tested and guaranteed during the PT reporting period only. Sample 01 and Sample 02 were spiked with several radionuclides. This included H-3 and U isotopes for Sample 01 and Sr-90, Ra-226 and progeny for Sample 02. The identification of gamma emitting radionuclides is one of the tasks of this proficiency test, therefore they are not specified in advance.

The activity concentration and uncertainty were assigned by formulation, and control measurements by gamma-ray spectrom- etry, alpha spectrometry and liquid scintillation counting have been performed prior to shipment.

## Sample 4, Japanese Soil

The soil was collected on 25 September 2020 in Japan by the University of Tsukuba. It was milled and sieved, and the grain size fraction used for this exercise was 90-250 µm. The bottled samples were sterilised (25 kGy gamma dose) and the related certificate is available on the PT website. It contains caesium and naturally occurring radionuclides typically found in soils.

Approximately 100 g has been provided for gamma-ray spectrometry measurement.

The determination of target values and associated uncertainties of the anthropogenic and natural radionuclides has been carried-out by measurements in one laboratory and confirmed by measurements in another laboratory. The samples were characterized for gamma emitting radionuclides only.

## Samples 5, 6 and 7, Simulated contaminated surface samples

Three circles for surface contamination monitoring have been included in this proficiency test (Figure 1). They were prepared using our in-house printing technique. One blank sample of the same size, outlined with black ink (not-spiked) was also added for correction (Sample 08).

The blue ink was spiked with a gamma-emitting radionuclide, Cs-134. The exercise asked for it’s identification and activity to be reported (Sample 06).

The magenta ink was spiked with Sr-90 and the activity should be reported (Sample 07), with correction for Y-90 activity.

The mixed ink sample (blue and magenta) has been printed at a lower activity compared to the single ink circles for both ra- dionuclides (Sample 05). Both Cs-134 and Sr-90 activity should be reported for this sample.

The reference date 2023-01-01 should be used for decay correction.

# 2	Evaluation of results

Results reported by participants were evaluated using the following stepwise approach.

The evaluation follows the methodology applied for the annual IAEA proficiency testing schemes in the areas of radionuclide measurements and trace element analysis.

## 2.1	Relative bias (Samples 1, 2, 4, 5, 6 and 7)

The relative bias is the relative difference between the reported and the target value (the best estimate of the true value)

V a l u e r e p o r t e d − V a l u e t a r g e t

Bi asrelative =

V alue

tar get

∗ 100%

The relative bias is compared to the Maximum Acceptable Relative Bias (MARB) which has been determined for each prop- erty value, considering the analytical methods, the analyte level in the sample and the complexity of the analysis.

When |Bi asrelative | ≤ M ARB , the result will be rated "Accepted (A)" for trueness.

## Evaluation approach for radionuclide activity concentration values (Samples 1, 2 and 4) and activ- ity values per sample (Samples 5, 6 and 7)

Based on fit-for-purpose and good laboratory practice principles, the expanded relative uncertainty should cover the relative bias (as defined in 2.1):

The P value corresponds to the relative combined uncertainty of the relative bias

「

， utar get ‘2

P =

， urepor ted ‘2

+

∗ 100%

I Atar get

Arepor ted

The relative bias is then compared to the expanded uncertainty of the relative bias:

|Bi asrelative | ≤ k ∗ P

where k is the coverage factor: k = 2.58 for a level of confidence of 99%.

When the above criterion is fulfilled, the reported result is not significantly different from the target values considering the uncertainties associated with both values. The reported uncertainty of measurement is large enough to cover the bias of the method.

In addition, the P value is compared to the MARB.

P ≤ M ARB

When this criterion is fulfilled, the measurement uncertainty is not overestimated and fit-for-purpose in relation to the MARB criterion of this PT exercise. When both criteria related to the measurement uncertainty are fulfilled, the reported result is rated "accepted (A)” for precision (measurement uncertainty). The result is rated "Not Accepted (N)" for precision if either of the two conditions are not fulfilled.

The final score is assigned according to the detailed evaluation described above. The possible scores are listed below:

    - "Accepted (A)" when both, trueness and precision were rated "Accepted"

    - "Not Accepted (N)" when the trueness rating is "Not Accepted"

    - "Warning (W)" when the trueness rating is “Accepted” but the precision rating is "Not Accepted"

## Evaluation parameters calculated for intercomparison analyte values

    - Z-Score

The z score provides information on the trueness of the reported value in relation to the standard deviation for proficiency assessment σPT , which is assigned based on fitness for purpose.

V a l u e r e p o r t e d − V a l u e t a r g e t

z =

σPT

The following criteria apply for performance ratings based on obtained z scores

      - |z| ≤ 2 ... accepted (A)

      - 2 < |z| < 3 ... questionable (Q)

      - |z| ≥ 3 ... not accepted (N)

# Data Evaluation

A total of 316 laboratories successfully reported data in the IAEA-TERC-2023-01 proficiency test exercise.

For each of the analytes evaluated, the target value, target uncertainty and the MARB are listed. A pie chart showing the distribution of accepted, warning, not acceptable and not reported results, the s-shape graph, and a detailed table listing all evaluation results for the analyte (sorted by the laboratory code) are provided. The data points are coloured in blue, yellow and red, corresponding to a result of "Accepted", "Warning", and "Not Accepted", respectively.

  - Evaluation of activity concentrations of radionuclides Reported Results for H-3 determination in Sample 1

TABLE 1. Target values

| Parameter | Value |
| --- | --- |
| Activity Concentration [ Bq/kg ] | 29.0 |
| Uncertainty [ Bq/kg ] | 1.5 |
| Maximum Acceptable Relative Bias, [%] | 30 |

Not Accepted 5%

Warning 2%

Accepted 32%

Not Reported 61%

FIGURE 1. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

Activity concentration  [ Bq/kg ]

30

35

40

45

15

20

25

224

115

89

83

42

304

58

97

97

307

116

189

164

299

205

293

19

133

72

146

272

201

26

193

34

295

62

47

157

25

75

175

100

149

141

158

57

266

67

298

1

11

328

196

203

195

14

43

329

126

296

105

277

186

290

87

35

267

185

291

166

106

268

289

217

300

44

248

257

114

95

223

209

79

255

145

235

333

341

151

270

46

27

188

150

245

69

241

129

264

342

262

54

107

137

94

218

130

173

120

199

6

40

281

119

279

317

172

85

232

63

96

112

102

123

155

45

160

104

247

202

167

20

98

263

Laboratory code

FIGURE 2. Reported results and evaluation parameters for H-3. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

TABLE 2. Evaluation Results for H-3 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg by beta

| Reported value [Bq/kg ] 27.6 | Reported uncertainty (k = 1) [Bq/kg ] 1.9 | Relative bias [%] -4.83 | P-Test [%] 8.61 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 31.6 | 0.6 | 8.97 | 5.51 | A | A | A |
| 27.6 | 2.0 | -4.83 | 8.90 | A | A | A |
| 27.9 | 1.0 | -3.79 | 6.29 | A | A | A |
| 24.9 | 1.0 | -14.14 | 6.55 | A | A | A |
| 52.8 | 3.696 | 82.07 | 8.70 | N | N | N |
| 26.7 | 4.0 | -7.93 | 15.85 | A | A | A |
| 25.6 | 2.02 | -11.72 | 9.43 | A | A | A |
| 29.48 | 0.93 | 1.66 | 6.06 | A | A | A |
| 26 | 1.4 | -10.34 | 7.47 | A | A | A |
| 28.2 | 1.8 | -2.76 | 8.22 | A | A | A |
| 31.81 | 2.23 | 9.69 | 8.71 | A | A | A |
| 19.4 | 2.8 | -33.10 | 15.33 | N | A | N |
| 27.9 | 1.00 | -3.79 | 6.29 | A | A | A |
| 28.7 | 1.2 | -1.03 | 6.65 | A | A | A |
| 38.6 | 5.45 | 33.10 | 15.04 | N | A | N |
| 29.3 | 2.3 | 1.03 | 9.40 | A | A | A |
| 26.3 | 1.9 | -9.31 | 8.89 | A | A | A |
| 30.5 | 5.7 | 5.17 | 19.39 | A | A | A |
| 27.2 | 1.7 | -6.21 | 8.11 | A | A | A |
| 20.08 | 3.64 | -30.76 | 18.85 | N | A | N |
| 26.3 | 5.4 | -9.31 | 21.17 | A | A | A |
| 33.3 | 3.0 | 14.83 | 10.39 | A | A | A |
| 27.45 | 2.75 | -5.34 | 11.27 | A | A | A |
| 29.6 | 2 | 2.07 | 8.51 | A | A | A |
| 25.2 | 1.0 | -13.10 | 6.52 | A | A | A |
| 26.74 | 0.77 | -7.79 | 5.92 | A | A | A |
| 29 | 5.95 | 0.00 | 21.16 | A | A | A |
| 19.01 | 2.66 | -34.45 | 14.92 | N | A | N |
| 33.00 | 1.24 | 13.79 | 6.39 | A | A | A |
| 28.20 | 6.96 | -2.76 | 25.22 | A | A | A |
| 17.92 | 1.6 | -38.21 | 10.32 | N | N | N |
| 30.72 | 0.95 | 5.93 | 6.03 | A | A | A |
| 28.8 | 1.6 | -0.69 | 7.59 | A | A | A |

Labcode

1

6

11

14

19

20

25

26

27

34

35

40

42

43

44

45

46

47

54

57

58

62

63

67

69

72

75

79

83

85

87

89

94

95

| Reported value [Bq/kg ] 33.698 | Reported uncertainty (k = 1) [Bq/kg ] 1.046 | Relative bias [%] 16.20 | P-Test [%] 6.03 | Trueness evaluation A | Precision evaluation N | Final Score W |
| --- | --- | --- | --- | --- | --- | --- |
| 21.25 | 4.72 | -26.72 | 22.81 | A | A | A |
| 54.8 | 3.84 | 88.97 | 8.71 | N | N | N |
| 27.0 | 1.8 | -6.90 | 8.44 | A | A | A |
| 34.10 | 4.37 | 17.59 | 13.82 | A | A | A |
| 43.31 | 12.2 | 49.34 | 28.64 | N | A | N |
| 28.09 | 1.55 | -3.14 | 7.56 | A | A | A |
| 28.5 | 8.6 | -1.72 | 30.62 | A | N | W |
| 30.7 | 4.1 | 5.86 | 14.32 | A | A | A |
| 34 | 5 | 17.24 | 15.59 | A | A | A |
| 28.8 | 2.0 | -0.69 | 8.66 | A | A | A |
| 15.09 | 0.95 | -47.97 | 8.15 | N | N | N |
| 21.86 | 2.12 | -24.62 | 10.99 | A | A | A |
| 32.3 | 5.6 | 11.38 | 18.09 | A | A | A |
| 31.34 | 2.89 | 8.07 | 10.57 | A | A | A |
| 34.7 | 1.3 | 19.66 | 6.39 | A | N | W |
| 28 | 3 | -3.45 | 11.90 | A | A | A |
| 30.09 | 0.57 | 3.76 | 5.51 | A | A | A |
| 30.9 | 4.6 | 6.55 | 15.76 | A | A | A |
| 25.0 | 1.8 | -13.79 | 8.87 | A | A | A |
| 30.72 | 0.95 | 5.93 | 6.03 | A | A | A |
| 27.1 | 0.99 | -6.55 | 6.33 | A | A | A |
| 29.1 | 2.7 | 0.34 | 10.62 | A | A | A |
| 25.2 | 2.35 | -13.10 | 10.66 | A | A | A |
| 27.05 | 0.82 | -6.72 | 6.00 | A | A | A |
| 29.5 | 2.8 | 1.72 | 10.81 | A | A | A |
| 29.3 | 1.7 | 1.03 | 7.77 | A | A | A |
| 38 | 4 | 31.03 | 11.73 | N | N | N |
| 26.5 | 2.5 | -8.62 | 10.76 | A | A | A |
| 27.12 | 2.71 | -6.48 | 11.25 | A | A | A |
| 42.6 | 1.7 | 46.90 | 6.53 | N | N | N |
| 23.3141 | 1.9203 | -19.61 | 9.73 | A | A | A |
| 28.3 | 5.1 | -2.41 | 18.75 | A | A | A |
| 50.10 | 3.09 | 72.76 | 8.05 | N | N | N |
| 32.9 | 0.82 | 13.45 | 5.74 | A | A | A |
| 30.97 | 2.22 | 6.79 | 8.84 | A | A | A |
| 26.9 | 1.8 | -7.24 | 8.46 | A | A | A |
| 28.29 | 5.09 | -2.45 | 18.72 | A | A | A |
| 28.1 | 0.7 | -3.10 | 5.74 | A | A | A |
| 29.5 | 3.1 | 1.72 | 11.71 | A | A | A |
| 23 | 3 | -20.69 | 14.03 | A | A | A |
| 25.61 | 1.28 | -11.69 | 7.19 | A | A | A |
| 27.8 | 2.1 | -4.14 | 9.16 | A | A | A |
| 27.7 | 3.2 | -4.48 | 12.66 | A | A | A |
| 31.59 | 1.42 | 8.93 | 6.85 | A | A | A |
| 25.4 | 1.4 | -12.41 | 7.56 | A | A | A |
| 44.51 | 0.10 | 53.48 | 5.18 | N | N | N |
| 27.78 | 1.61 | -4.21 | 7.77 | A | A | A |
| 24.42 | 0.19 | -15.79 | 5.23 | A | N | W |
| 28.9 | 2.3 | -0.34 | 9.49 | A | A | A |
| 28.6 | 6.8 | -1.38 | 24.33 | A | A | A |
| 30.8 | 4.0 | 6.21 | 13.98 | A | A | A |
| 28.84 | 2.93 | -0.55 | 11.40 | A | A | A |
| 5.497 | 2.4 | -81.04 | 43.97 | N | N | N |
| 33.1 | 2.3 | 14.14 | 8.66 | A | A | A |
| 29.2 | 1.2 | 0.69 | 6.61 | A | A | A |
| 30 | 7 | 3.45 | 23.90 | A | A | A |
| 29.5 | 1.4 | 1.72 | 7.02 | A | A | A |
| 43.7 | 1.3 | 50.69 | 5.97 | N | N | N |
| 28.8 | 3.7 | -0.69 | 13.85 | A | A | A |

TABLE 2. Evaluation Results for H-3 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg by beta

Labcode

96

97

98

100

102

104

105

106

107

112

114

115

116

119

120

123

126

129

130

133

137

141

145

146

149

150

151

155

157

158

160

164

166

167

172

173

175

185

186

188

189

193

195

196

199

201

202

203

205

209

217

218

223

224

232

235

241

245

247

248

TABLE 2. Evaluation Results for H-3 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg by beta

| Reported value [Bq/kg ] 29.1 | Reported uncertainty (k = 1) [Bq/kg ] 1.5 | Relative bias [%] 0.34 | P-Test [%] 7.30 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 28.8 | 2.0 | -0.69 | 8.66 | A | A | A |
| 30.4 | 2.43 | 4.83 | 9.52 | A | A | A |
| 73 | 8 | 151.72 | 12.12 | N | N | N |
| 30.15 | 2.72 | 3.97 | 10.40 | A | A | A |
| 27.4 | 4.65 | -5.52 | 17.74 | A | A | A |
| 28.2 | 6.4 | -2.76 | 23.28 | A | A | A |
| 28.5 | 0.97 | -1.72 | 6.19 | A | A | A |
| 29.3 | 2.1 | 1.03 | 8.84 | A | A | A |
| 25.3 | 3.22 | -12.76 | 13.74 | A | A | A |
| 28.10 | 1.00 | -3.10 | 6.28 | A | A | A |
| 32.39 | 1.74 | 11.69 | 7.46 | A | A | A |
| 31.9 | 2.2 | 10.00 | 8.62 | A | A | A |
| 28.55 | 0.95 | -1.55 | 6.15 | A | A | A |
| 28.1 | 0.81 | -3.10 | 5.92 | A | A | A |
| 28.3 | 1.3 | -2.41 | 6.92 | A | A | A |
| 24.8 | 1.4 | -14.48 | 7.66 | A | A | A |
| 26.1 | 3.2 | -10.00 | 13.31 | A | A | A |
| 28.06 | 1.07 | -3.24 | 6.43 | A | A | A |
| 27.5 | 0.8 | -5.17 | 5.93 | A | A | A |
| 23.9 | 0.9 | -17.59 | 6.40 | A | N | W |
| 28.7 | 2 | -1.03 | 8.68 | A | A | A |
| 19.9 | 1.6 | -31.38 | 9.56 | N | N | N |
| 21.30 | 1.98 | -26.55 | 10.64 | A | A | A |
| 32.42 | 1.05 | 11.79 | 6.10 | A | A | A |
| 27.7 | 1.1 | -4.48 | 6.52 | A | A | A |
| 27.9 | 0.8 | -3.79 | 5.91 | A | A | A |
| 29.26 | 0.93 | 0.90 | 6.07 | A | A | A |
| 29.26 | 12.49 | 0.90 | 43.00 | A | N | W |
| 30.20 | 2.11 | 4.14 | 8.69 | A | A | A |

Labcode

255

257

262

263

264

266

267

268

270

272

277

279

281

289

290

291

293

295

296

298

299

300

304

307

317

328

329

333

341

342

Activity concentration  [ Bq/kg ]

14

16

Reported Results for Co-60 determination in Sample 1

TABLE 3. Target values

| Parameter | Value |
| --- | --- |
| Activity Concentration [ Bq/kg ] | 14.7 |
| Uncertainty [ Bq/kg ] | 0.7 |
| Maximum Acceptable Relative Bias, [%] | 20 |

Accepted 83%

Not Reported 11%

Warning 3%

Not Accepted 3%

18

20

FIGURE 3. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

10

12

336

282

154

144

56

13

210

89

93

172

108

170

47

142

80

46

153

132

57

222

102

70

169

241

306

147

309

284

305

75

258

21

207

42

224

52

277

38

199

177

19

148

227

95

155

179

202

188

287

299

107

273

292

301

27

276

300

73

63

248

134

260

120

182

291

307

240

23

122

166

35

316

8

106

146

91

94

290

16

81

186

130

218

173

69

205

141

251

167

123

118

59

334

162

11

5

296

234

228

136

204

139

158

32

270

264

209

36

293

272

184

245

232

303

314

140

77

328

335

190

261

2

115

181

295

193

15

275

1

33

98

34

318

112

249

124

116

214

255

97

129

322

62

Laboratory code

FIGURE 4. Reported results and evaluation parameters for Co-60. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

| Reported value [Bq/kg ] 15.6 | Reported uncertainty (k = 1) [Bq/kg ] 0.5 | Relative bias [%] 6.12 | P-Test [%] 5.74 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 15.39 | 0.67 | 4.69 | 6.45 | A | A | A |
| 14.4 | 0.44 | -2.04 | 5.66 | A | A | A |
| 14.92 | 0.95 | 1.50 | 7.95 | A | A | A |
| 14.3 | 0.5 | -2.72 | 5.91 | A | A | A |
| 13.64 | 0.57 | -7.21 | 6.34 | A | A | A |
| 14.71 | 0.47 | 0.07 | 5.73 | A | A | A |
| 14.6 | 0.6 | -0.68 | 6.29 | A | A | A |
| 14.9 | 0.7 | 1.36 | 6.69 | A | A | A |
| 13.9 | 0.6 | -5.44 | 6.43 | A | A | A |
| 12.5 | 0.3 | -14.97 | 5.33 | A | N | W |
| 14.85 | 0.97 | 1.02 | 8.08 | A | A | A |
| 15.5 | 0.7 | 5.44 | 6.56 | A | A | A |
| 14.80 | 0.39 | 0.68 | 5.44 | A | A | A |
| 14.6 | 0.8 | -0.68 | 7.26 | A | A | A |
| 14.40 | 0.81 | -2.04 | 7.37 | A | A | A |
| 14.4 | 1.0 | -2.04 | 8.42 | A | A | A |
| 13.2 | 1.02 | -10.20 | 9.08 | A | A | A |
| 14.2 | 0.64 | -3.40 | 6.56 | A | A | A |
| 14.82 | 0.52 | 0.82 | 5.91 | A | A | A |
| 14.7 | 0.3 | 0.00 | 5.18 | A | A | A |
| 15.6 | 0.695 | 6.12 | 6.52 | A | A | A |
| 14.99 | 1.5 | 1.97 | 11.08 | A | A | A |
| 14.5 | 0.58 | -1.36 | 6.22 | A | A | A |
| 14.59 | 0.39 | -0.75 | 5.46 | A | A | A |
| 15.8 | 1.6 | 7.48 | 11.19 | A | A | A |
| 15.2 | 0.7 | 3.40 | 6.62 | A | A | A |
| 14.7 | 0.7 | 0.00 | 6.73 | A | A | A |
| 14.8 | 0.9 | 0.68 | 7.72 | A | A | A |
| 15.01 | 0.26 | 2.11 | 5.07 | A | A | A |
| 15.6 | 2.1 | 6.12 | 14.28 | A | A | A |
| 15.8 | 2.2 | 7.48 | 14.72 | A | A | A |
| 14.7 | 1.2 | 0.00 | 9.45 | A | A | A |
| 15.1 | 1.2 | 2.72 | 9.26 | A | A | A |
| 14.89 | 0.97 | 1.29 | 8.07 | A | A | A |
| 14.3 | 0.9 | -2.72 | 7.89 | A | A | A |
| 14.1 | 0.6 | -4.08 | 6.39 | A | A | A |
| 14.7 | 0.614 | 0.00 | 6.33 | A | A | A |
| 14.2 | 1.98 | -3.40 | 14.73 | A | A | A |
| 14.5 | 1.3 | -1.36 | 10.15 | A | A | A |
| 14.7 | 0.6 | 0.00 | 6.27 | A | A | A |
| 15.3 | 0.92 | 4.08 | 7.67 | A | A | A |
| 13.8 | 0.14 | -6.12 | 4.87 | A | A | A |
| 13.6 | 0.4 | -7.48 | 5.60 | A | A | A |
| 15.1 | 0.7 | 2.72 | 6.65 | A | A | A |
| 14.71 | 0.25 | 0.07 | 5.06 | A | A | A |
| 14.81 | 0.24 | 0.75 | 5.03 | A | A | A |
| 16.0 | 1.0 | 8.84 | 7.86 | A | A | A |
| 14.3 | 0.6 | -2.72 | 6.35 | A | A | A |
| 13.9 | 0.2 | -5.44 | 4.97 | A | A | A |
| 15.3 | 0.8 | 4.08 | 7.07 | A | A | A |
| 12.12 | 0.61 | -17.55 | 6.93 | A | A | A |
| 13.9 | 1.0 | -5.44 | 8.63 | A | A | A |
| 14.80 | 2.37 | 0.68 | 16.71 | A | A | A |
| 14.9 | 0.866 | 1.36 | 7.51 | A | A | A |
| 14.42 | 0.64 | -1.90 | 6.51 | A | A | A |
| 34.7 | 1.6 | 136.05 | 6.63 | N | N | N |
| 14.6 | 1.9 | -0.68 | 13.86 | A | A | A |
| 14.2 | 1.28 | -3.40 | 10.19 | A | A | A |
| 15.33 | 0.3 | 4.29 | 5.15 | A | A | A |

TABLE 4. Evaluation Results for Co-60 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

1

2

3

5

6

7

8

10

11

12

13

14

15

16

17

18

19

20

21

22

23

24

25

26

27

28

29

30

31

32

33

34

35

36

37

38

39

40

42

43

44

45

46

47

48

49

50

51

52

54

55

56

57

58

59

60

62

63

64

66

| Reported value [Bq/kg ] 14.53 | Reported uncertainty (k = 1) [Bq/kg ] 0.3 | Relative bias [%] -1.16 | P-Test [%] 5.19 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 15.00 | 0.56 | 2.04 | 6.05 | A | A | A |
| 14.82 | 0.28 | 0.82 | 5.12 | A | A | A |
| 14.0 | 0.1 | -4.76 | 4.82 | A | A | A |
| 14.6 | 0.2 | -0.68 | 4.96 | A | A | A |
| 14.30 | 0.23 | -2.72 | 5.03 | A | A | A |
| 14.6 | 0.3 | -0.68 | 5.19 | A | A | A |
| 14.76 | 0.58 | 0.41 | 6.17 | A | A | A |
| 14.14 | 0.16 | -3.81 | 4.89 | A | A | A |
| 31.99 | 5.47 | 117.62 | 17.75 | N | N | N |
| 15.2 | 0.6 | 3.40 | 6.19 | A | A | A |
| 14.8 | 0.5 | 0.68 | 5.84 | A | A | A |
| 13.7 | 0.2 | -6.80 | 4.98 | A | A | A |
| 14.8 | 0.5 | 0.68 | 5.84 | A | A | A |
| 15.04 | 0.7 | 2.31 | 6.66 | A | A | A |
| 15.66 | 0.68 | 6.53 | 6.44 | A | A | A |
| 15.4 | 0.69 | 4.76 | 6.54 | A | A | A |
| 14.96 | 0.89 | 1.77 | 7.62 | A | A | A |
| 15.0 | 0.6 | 2.04 | 6.22 | A | A | A |
| 15.09 | 0.40 | 2.65 | 5.45 | A | A | A |
| 13.1 | 0.8 | -10.88 | 7.74 | A | A | A |
| 14.49 | 0.52 | -1.43 | 5.96 | A | A | A |
| 14.76 | 0.37 | 0.41 | 5.38 | A | A | A |
| 2.43 | 0.061 | -83.47 | 5.38 | N | N | N |
| 13.20 | 0.44 | -10.20 | 5.81 | A | A | A |
| 14.77 | 0.61 | 0.48 | 6.30 | A | A | A |
| 14.4 | 1.1 | -2.04 | 9.00 | A | A | A |
| 15.39 | 2.10 | 4.69 | 14.45 | A | A | A |
| 17.15 | 0.23 | 16.67 | 4.95 | A | N | W |
| 15.68 | 0.83 | 6.67 | 7.12 | A | A | A |
| 15.03 | 0.40 | 2.24 | 5.46 | A | A | A |
| 10.18 |  | -30.75 |  | N |  | N |
| 13.95 | 0.48 | -5.10 | 5.87 | A | A | A |
| 14.03 | 0.58 | -4.56 | 6.31 | A | A | A |
| 14.58 | 0.87 | -0.82 | 7.63 | A | A | A |
| 14.74 | 0.76 | 0.27 | 7.02 | A | A | A |
| 14.5 | 0.8 | -1.36 | 7.29 | A | A | A |
| 13.5 | 2.5 | -8.16 | 19.12 | A | A | A |
| 12.29 | 0.53 | -16.39 | 6.42 | A | A | A |
| 16 | 1 | 8.84 | 7.86 | A | A | A |
| 14.7 | 1.1 | 0.00 | 8.87 | A | A | A |
| 15.39 | 0.15 | 4.69 | 4.86 | A | A | A |
| 16.47 | 0.84 | 12.04 | 6.98 | A | A | A |
| 14.9 | 0.4 | 1.36 | 5.47 | A | A | A |
| 14.6 | 0.5 | -0.68 | 5.87 | A | A | A |
| 14.66 | 0.51 | -0.27 | 5.90 | A | A | A |
| 14.7 | 0.8 | 0.00 | 7.23 | A | A | A |
| 14.9 | 0.7 | 1.36 | 6.69 | A | A | A |
| 16.31 | 0.47 | 10.95 | 5.57 | A | A | A |
| 14.91 | 1.19 | 1.43 | 9.29 | A | A | A |
| 12.7 | 0.2 | -13.61 | 5.02 | A | N | W |
| 17.64 | 0.82 | 20.00 | 6.65 | A | N | W |
| 14.8 | 0.9 | 0.68 | 7.72 | A | A | A |
| 14.66 | 0.51 | -0.27 | 5.90 | A | A | A |
| 13.9 | 1.3 | -5.44 | 10.50 | A | A | A |
| 15.9 | 1.3 | 8.16 | 9.46 | A | A | A |
| 14.6 | 0.6 | -0.68 | 6.29 | A | A | A |
| 14.6 | 1.2 | -0.68 | 9.50 | A | A | A |
| 15 | 1 | 2.04 | 8.19 | A | A | A |
| 14.77 | 0.61 | 0.48 | 6.30 | A | A | A |

TABLE 4. Evaluation Results for Co-60 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

67

68

69

70

71

72

73

74

75

76

77

79

80

81

82

83

84

85

86

87

89

90

91

92

93

94

95

96

97

98

99

101

102

104

105

106

107

108

110

112

114

115

116

118

119

120

122

123

124

125

126

129

130

131

132

133

134

135

136

137

| Reported value [Bq/kg ] 15.0 | Reported uncertainty (k = 1) [Bq/kg ] 0.54 | Relative bias [%] 2.04 | P-Test [%] 5.97 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 15.2 | 0.9 | 3.40 | 7.60 | A | A | A |
| 14.85 | 0.75 | 1.02 | 6.94 | A | A | A |
| 13.625 | 1.074 | -7.31 | 9.21 | A | A | A |
| 11.79 | 0.67 | -19.80 | 7.41 | A | N | W |
| 14.0 | 0.9 | -4.76 | 8.00 | A | A | A |
| 14.74 | 0.91 | 0.27 | 7.80 | A | A | A |
| 14.04 | 0.17 | -4.49 | 4.91 | A | A | A |
| 14.4 | 0.7 | -2.04 | 6.80 | A | A | A |
| 16.6 | 0.535 | 12.93 | 5.75 | A | A | A |
| 15.0 | 0.6 | 2.04 | 6.22 | A | A | A |
| 14.9 | 0.9 | 1.36 | 7.69 | A | A | A |
| 15.4 | 0.67 | 4.76 | 6.45 | A | A | A |
| 13.846 | 0.651 | -5.81 | 6.69 | A | A | A |
| 11.16 | 0.33 | -24.08 | 5.61 | N | N | N |
| 14.41 | 0.53 | -1.97 | 6.02 | A | A | A |
| 14.8 | 1 | 0.68 | 8.27 | A | A | A |
| 15.1 | 0.3 | 2.72 | 5.16 | A | A | A |
| 15.01 | 0.33 | 2.11 | 5.24 | A | A | A |
| 13.58 | 0.40 | -7.62 | 5.60 | A | A | A |
| 14.8 | 0.8 | 0.68 | 7.20 | A | A | A |
| 15.6 | 1.3 | 6.12 | 9.60 | A | A | A |
| 14.9 | 1.0 | 1.36 | 8.23 | A | A | A |
| 14.5 | 0.76 | -1.36 | 7.08 | A | A | A |
| 14.83 | 0.26 | 0.88 | 5.07 | A | A | A |
| 14.7 | 1.5 | 0.00 | 11.26 | A | A | A |
| 14.90 | 0.72 | 1.36 | 6.78 | A | A | A |
| 15.1 | 0.8 | 2.72 | 7.12 | A | A | A |
| 14.0 | 1.2 | -4.76 | 9.81 | A | A | A |
| 13.6 | 0.6 | -7.48 | 6.49 | A | A | A |
| 15.4 | 0.45 | 4.76 | 5.59 | A | A | A |
| 13.3 | 0.835 | -9.52 | 7.88 | A | A | A |
| 14.814 | 0.5 | 0.78 | 5.84 | A | A | A |
| 15.7 | 0.6 | 6.80 | 6.11 | A | A | A |
| 14.2 | 1.3 | -3.40 | 10.32 | A | A | A |
| 14.38 | 0.26 | -2.18 | 5.09 | A | A | A |
| 14.14 | 0.43 | -3.81 | 5.65 | A | A | A |
| 14.46 | 0.42 | -1.63 | 5.58 | A | A | A |
| 14.35 | 0.31 | -2.38 | 5.23 | A | A | A |
| 15.4 | 0.8 | 4.76 | 7.05 | A | A | A |
| 14.68 | 1.18 | -0.14 | 9.34 | A | A | A |
| 14.6 | 0.7 | -0.68 | 6.76 | A | A | A |
| 15.1 | 1.7 | 2.72 | 12.22 | A | A | A |
| 14.67 | 1.54 | -0.20 | 11.53 | A | A | A |
| 14.8 | 0.9 | 0.68 | 7.72 | A | A | A |
| 14.5 | 0.9 | -1.36 | 7.82 | A | A | A |
| 14.6 | 1.2 | -0.68 | 9.50 | A | A | A |
| 15.3 | 0.5 | 4.08 | 5.78 | A | A | A |
| 16.2 | 1.1 | 10.20 | 8.29 | A | A | A |
| 15.45 | 0.66 | 5.10 | 6.40 | A | A | A |
| 15.2 | 1.3 | 3.40 | 9.79 | A | A | A |
| 14.9213 | 0.6372 | 1.51 | 6.40 | A | A | A |
| 16.4 | 0.8 | 11.56 | 6.82 | A | A | A |
| 14.37 | 0.52 | -2.24 | 5.98 | A | A | A |
| 14.74 | 0.60 | 0.27 | 6.26 | A | A | A |
| 14.70 | 0.95 | 0.00 | 8.03 | A | A | A |
| 14.50 | 1.55 | -1.36 | 11.70 | A | A | A |
| 15.0 | 0.8 | 2.04 | 7.15 | A | A | A |
| 14.85 | 0.48 | 1.02 | 5.76 | A | A | A |
| 14.404 | 0.400 | -2.01 | 5.51 | A | A | A |

TABLE 4. Evaluation Results for Co-60 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

139

140

141

142

144

145

146

147

148

149

150

151

152

153

154

155

156

157

158

159

160

161

162

163

165

166

167

168

169

170

171

172

173

174

176

177

178

179

180

181

182

183

184

185

186

188

189

190

192

193

195

197

198

199

200

201

202

204

205

206

| Reported value [Bq/kg ] 14.20 | Reported uncertainty (k = 1) [Bq/kg ] 0.50 | Relative bias [%] -3.40 | P-Test [%] 5.92 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 15.07 | 1.24 | 2.52 | 9.51 | A | A | A |
| 12.74 | 0.87 | -13.33 | 8.33 | A | A | A |
| 15.2 | 0.9 | 3.40 | 7.60 | A | A | A |
| 16.6 | 1.1 | 12.93 | 8.16 | A | A | A |
| 14.10 | 0.60 | -4.08 | 6.39 | A | A | A |
| 14.8 | 0.5 | 0.68 | 5.84 | A | A | A |
| 12.02 | 0.60 | -18.23 | 6.90 | A | N | W |
| 13.90 | 0.40 | -5.44 | 5.56 | A | A | A |
| 13.80 | 0.46 | -6.12 | 5.81 | A | A | A |
| 14.3 | 0.2 | -2.72 | 4.96 | A | A | A |
| 15.1 | 1.2 | 2.72 | 9.26 | A | A | A |
| 17.57 | 0.2 | 19.52 | 4.90 | A | N | W |
| 14.4 | 0.02 | -2.04 | 4.76 | A | A | A |
| 15.0 | 0.4 | 2.04 | 5.46 | A | A | A |
| 15.1 | 1.0 | 2.72 | 8.16 | A | A | A |
| 15.02 | 0.9 | 2.18 | 7.65 | A | A | A |
| 14.99 | 0.38 | 1.97 | 5.39 | A | A | A |
| 14.9 | 0.8 | 1.36 | 7.18 | A | A | A |
| 12.9 | 0.55 | -12.24 | 6.39 | A | A | A |
| 14.90 | 0.16 | 1.36 | 4.88 | A | A | A |
| 15.00 | 0.44 | 2.04 | 5.59 | A | A | A |
| 14.7 | 0.8 | 0.00 | 7.23 | A | A | A |
| 14.0 | 1.0 | -4.76 | 8.58 | A | A | A |
| 14 | 1.4 | -4.76 | 11.08 | A | A | A |
| 15.2 | 1.13 | 3.40 | 8.83 | A | A | A |
| 15.1 | 1.1 | 2.72 | 8.70 | A | A | A |
| 15.14 | 0.82 | 2.99 | 7.21 | A | A | A |
| 14.26 | 0.73 | -2.99 | 6.99 | A | A | A |
| 14.6 | 0.8 | -0.68 | 7.26 | A | A | A |
| 16.1 | 0.5 | 9.52 | 5.69 | A | A | A |
| 14.88 | 0.18 | 1.22 | 4.91 | A | A | A |
| 14.4 | 1.4 | -2.04 | 10.83 | A | A | A |
| 11.4 | 0.21 | -22.45 | 5.11 | N | N | N |
| 16.730 | 1.1653 | 13.81 | 8.44 | A | A | A |
| 15.6 | 0.784 | 6.12 | 6.92 | A | A | A |
| 14.9 | 0.6 | 1.36 | 6.24 | A | A | A |
| 14.17 | 0.52 | -3.61 | 6.01 | A | A | A |
| 14.5 | 1.5 | -1.36 | 11.39 | A | A | A |
| 14.63 | 1.10 | -0.48 | 8.90 | A | A | A |
| 15.33 | 0.65 | 4.29 | 6.38 | A | A | A |
| 15.5 | 0.882 | 5.44 | 7.42 | A | A | A |
| 16.6 | 0.3 | 12.93 | 5.09 | A | A | A |
| 15.04 | 1.343 | 2.31 | 10.12 | A | A | A |
| 15.2 | 1.0 | 3.40 | 8.12 | A | A | A |
| 14.06 | 0.19 | -4.35 | 4.95 | A | A | A |
| 15.02 | 0.87 | 2.18 | 7.50 | A | A | A |
| 14.7 | 0.9 | 0.00 | 7.76 | A | A | A |
| 15.1 | 0.7 | 2.72 | 6.65 | A | A | A |
| 14.5 | 0.6 | -1.36 | 6.31 | A | A | A |
| 14.5 | 0.882 | -1.36 | 7.73 | A | A | A |
| 15.6 | 0.8 | 6.12 | 7.00 | A | A | A |
| 14.6 | 1.0 | -0.68 | 8.34 | A | A | A |
| 14.30 | 0.19 | -2.72 | 4.94 | A | A | A |
| 15.01 | 0.466 | 2.11 | 5.68 | A | A | A |
| 14.39 | 1.73 | -2.11 | 12.93 | A | A | A |
| 7.5 | 0.17 | -48.98 | 5.27 | N | N | N |
| 14.55 | 2.207 | -1.02 | 15.90 | A | A | A |
| 14.1 | 1.1 | -4.08 | 9.14 | A | A | A |
| 13.983 | 0.58804 | -4.88 | 6.35 | A | A | A |

TABLE 4. Evaluation Results for Co-60 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

207

209

210

213

214

217

218

221

222

223

224

225

226

227

228

232

233

234

235

236

237

239

240

241

243

244

245

246

247

248

249

251

252

254

255

256

257

258

259

260

261

262

263

264

268

269

270

271

272

273

274

275

276

277

279

281

282

283

284

285

TABLE 4. Evaluation Results for Co-60 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg by gamma

| Reported value [Bq/kg ] 13.7 | Reported uncertainty (k = 1) [Bq/kg ] 0.7 | Relative bias [%] -6.80 | P-Test [%] 6.98 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 14.5 | 0.79 | -1.36 | 7.24 | A | A | A |
| 13.6 | 0.23 | -7.48 | 5.05 | A | A | A |
| 14.8 | 0.24 | 0.68 | 5.03 | A | A | A |
| 14.7 | 1.01 | 0.00 | 8.36 | A | A | A |
| 14.53 | 0.84 | -1.16 | 7.49 | A | A | A |
| 15.1 | 0.7 | 2.72 | 6.65 | A | A | A |
| 13.6 | 0.7 | -7.48 | 7.01 | A | A | A |
| 15.4 | 1.1 | 4.76 | 8.58 | A | A | A |
| 14.94 | 0.56 | 1.63 | 6.06 | A | A | A |
| 14.0 | 0.8 | -4.76 | 7.44 | A | A | A |
| 14.5 | 0.3 | -1.36 | 5.19 | A | A | A |
| 14.6 | 0.9 | -0.68 | 7.79 | A | A | A |
| 14.55 | 0.53 | -1.02 | 6.00 | A | A | A |
| 14.8 | 0.7 | 0.68 | 6.71 | A | A | A |
| 15.15 | 1.68 | 3.06 | 12.07 | A | A | A |
| 14.7 | 1.2 | 0.00 | 9.45 | A | A | A |
| 14.1 | 0.4 | -4.08 | 5.54 | A | A | A |
| 14.01 | 1.90 | -4.69 | 14.37 | A | A | A |
| 14.70 | 0.29 | 0.00 | 5.15 | A | A | A |
| 13.89 | 0.23 | -5.51 | 5.04 | A | A | A |
| 14.06 | 0.80 | -4.35 | 7.42 | A | A | A |
| 18.6 | 1.1 | 26.53 | 7.59 | N | N | N |
| 15.2 | 1.5 | 3.40 | 10.96 | A | A | A |
| 14.2 | 1.2 | -3.40 | 9.70 | A | A | A |
| 14.7 | 0.6 | 0.00 | 6.27 | A | A | A |
| 13.35 | 0.58 | -9.18 | 6.45 | A | A | A |
| 15.9 | 0.6 | 8.16 | 6.08 | A | A | A |
| 14.064 | 3.375 | -4.33 | 24.47 | A | N | W |
| 28.84 | 3.06 | 96.19 | 11.63 | N | N | N |
| 14.9 | 0.3 | 1.36 | 5.17 | A | A | A |
| 15.2 | 0.9 | 3.40 | 7.60 | A | A | A |
| 15.1 | 1.0 | 2.72 | 8.16 | A | A | A |
| 14.7 | 1.4 | 0.00 | 10.65 | A | A | A |
| 13.94 | 0.34 | -5.17 | 5.35 | A | A | A |
| 14.3 | 0.7 | -2.72 | 6.83 | A | A | A |
| 14.9 | 0.7 | 1.36 | 6.69 | A | A | A |
| 15.23 | 1.05 | 3.61 | 8.38 | A | A | A |
| 1.88 | 0.12 | -87.21 | 7.96 | N | N | N |
| 14.9 | 0.7 | 1.36 | 6.69 | A | A | A |
| 14.38 | 0.26 | -2.18 | 5.09 | A | A | A |
| 14.72 | 0.60 | 0.14 | 6.27 | A | A | A |
| 14.85 | 0.27 | 1.02 | 5.10 | A | A | A |
| 13.1779 |  | -10.35 |  | A |  |  |
| 16.8 | 2.9 | 14.29 | 17.91 | A | A | A |

Labcode

286

287

288

290

291

292

293

294

295

296

298

299

300

301

302

303

304

305

306

307

308

309

311

314

315

316

317

318

319

322

327

328

329

330

331

333

334

335

336

337

338

341

342

343

344

Reported Results for U-234 determination in Sample 1

TABLE 5. Target values

| Parameter | Value |
| --- | --- |
| Activity Concentration [ Bq/kg ] | 5.44 |
| Uncertainty [ Bq/kg ] | 0.27 |
| Maximum Acceptable Relative Bias, [%] | 30 |

Warning 1%

Not Accepted 3%

Activity concentration  [ Bq/kg ]

5

6

7

Accepted 22%

Not Reported 74%

8

FIGURE 5. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

2

3

4

126

75

76

166

185

223

104

140

68

57

123

210

210

89

232

333

173

129

266

300

102

135

112

172

288

130

279

54

268

270

215

215

203

31

178

63

272

79

248

81

1

119

180

186

245

85

235

72

44

258

22

67

105

145

262

19

202

21

167

290

27

188

69

43

35

189

25

39

225

120

330

315

184

199

263

106

198

127

247

65

Laboratory code

FIGURE 6. Reported results and evaluation parameters for U-234. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

| Reported value [Bq/kg ] 5.44 | Reported uncertainty (k = 1) [Bq/kg ] 0.28 | Relative bias [%] 0.00 | P-Test [%] 7.15 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 5.66 | 0.40 | 4.04 | 8.64 | A | A | A |
| 5.68 | 0.43 | 4.41 | 9.05 | A | A | A |
| 5.62 | 0.52 | 3.31 | 10.50 | A | A | A |
| 5.8 | 0.9 | 6.62 | 16.29 | A | A | A |
| 5.72 | 0.13 | 5.15 | 5.46 | A | A | A |
| 5.3 |  | -2.57 |  | A |  |  |
| 5.79 | 0.37 | 6.43 | 8.09 | A | A | A |
| 5.82 | 0.31 | 6.99 | 7.28 | A | A | A |
| 5.78 | 0.23 | 6.25 | 6.36 | A | A | A |
| 5.58 | 0.34 | 2.57 | 7.86 | A | A | A |
| 5.23 | 0.28 | -3.86 | 7.30 | A | A | A |
| 4.58 | 0.25 | -15.81 | 7.38 | A | A | A |
| 5.33 | 0.59 | -2.02 | 12.13 | A | A | A |
| 76.62 | 19.16 | 1308.46 | 25.49 | N | N | N |
| 5.64 | 0.6 | 3.68 | 11.74 | A | A | A |
| 4.54 | 0.31 | -16.54 | 8.44 | A | A | A |
| 5.764 | 0.132 | 5.96 | 5.47 | A | A | A |
| 5.57 | 0.13 | 2.39 | 5.48 | A | A | A |
| 0.162 | 0.051 | -97.02 | 31.87 | N | N | N |
| 0.378 | 0.016 | -93.05 | 6.52 | N | N | N |
| 5.357 | 0.179 | -1.53 | 5.98 | A | A | A |
| 5.43 | 0.11 | -0.18 | 5.36 | A | A | A |
| 5.52 | 0.18 | 1.47 | 5.94 | A | A | A |
| 4.64 | 0.23 | -14.71 | 7.01 | A | A | A |
| 5.09 | 0.78 | -6.43 | 16.11 | A | A | A |
| 4.22 | 1.12 | -22.43 | 27.00 | A | A | A |
| 5.64 | 0.38 | 3.68 | 8.37 | A | A | A |
| 6.19 | 0.69 | 13.79 | 12.20 | A | A | A |
| 5.1 | 0.3 | -6.25 | 7.70 | A | A | A |
| 5.45 | 0.3 | 0.18 | 7.41 | A | A | A |
| 5.88 | 1.00 | 8.09 | 17.72 | A | A | A |
| 4.60 | 0.43 | -15.44 | 10.58 | A | A | A |
| 0.000311 | 0.000003 | -99.99 | 5.06 | N | N | N |
| 7.51 | 0.851 | 38.05 | 12.37 | N | N | N |
| 4.97 | 0.50 | -8.64 | 11.22 | A | A | A |
| 5.13 | 0.64 | -5.70 | 13.43 | A | A | A |
| 5.1 | 0.7 | -6.25 | 14.60 | A | A | A |
| 4.32 | 0.22 | -20.59 | 7.11 | A | N | W |
| 5.65 | 0.23 | 3.86 | 6.42 | A | A | A |
| 3.4 | 0.6 | -37.50 | 18.33 | N | A | N |
| 5.69 | 1.62 | 4.60 | 28.90 | A | A | A |
| 5.11 | 0.053 | -6.07 | 5.07 | A | A | A |
| 4.9 | 0.22 | -9.93 | 6.69 | A | A | A |
| 5.30 | 0.17 | -2.57 | 5.91 | A | A | A |
| 5.45 | 0.21 | 0.18 | 6.28 | A | A | A |
| 6.01 | 0.61 | 10.48 | 11.30 | A | A | A |
| 3.41 | 0.58 | -37.32 | 17.72 | N | A | N |
| 5.47 | 0.15 | 0.55 | 5.67 | A | A | A |
| 5.74 | 0.22 | 5.51 | 6.27 | A | A | A |
| 5.8 | 0.6 | 6.62 | 11.47 | A | A | A |
| 6.6 | 0.4 | 21.32 | 7.83 | A | N | W |
| 6.03 | 0.17 | 10.85 | 5.71 | A | A | A |
| 5.68 | 0.20 | 4.41 | 6.09 | A | A | A |
| 5.29 | 0.15 | -2.76 | 5.72 | A | A | A |
| 4.60 | 0.99 | -15.44 | 22.09 | A | A | A |
| 5.28 | 0.12 | -2.94 | 5.46 | A | A | A |
| 3.52 | 0.13 | -35.29 | 6.19 | N | N | N |
| 5.85 | 0.38 | 7.54 | 8.17 | A | A | A |
| 4.7 | 0.4 | -13.60 | 9.85 | A | A | A |

TABLE 6. Evaluation Results for U-234 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg by alpha

Labcode

1

19

21

22

25

27

31

35

39

43

44

54

57

63

65

67

68

69

72

75

76

79

81

85

89

102

104

105

106

112

119

120

123

126

127

129

130

135

140

145

166

167

172

173

178

180

184

185

186

188

189

198

199

202

203

210

215

223

225

232

TABLE 6. Evaluation Results for U-234 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg by alpha

| Reported value [Bq/kg ] 5.55 | Reported uncertainty (k = 1) [Bq/kg ] 0.20 | Relative bias [%] 2.02 | P-Test [%] 6.13 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 5.47 | 0.74 | 0.55 | 14.41 | A | A | A |
| 9.95 | 0.2 | 82.90 | 5.35 | N | N | N |
| 5.36 | 0.79 | -1.47 | 15.55 | A | A | A |
| 5.58 | 0.10 | 2.57 | 5.28 | A | A | A |
| 5.66 | 0.258 | 4.04 | 6.74 | A | A | A |
| 6.11 | 0.24 | 12.32 | 6.33 | A | A | A |
| 5.00 | 0.365 | -8.09 | 8.83 | A | A | A |
| 5.23 | 0.14 | -3.86 | 5.64 | A | A | A |
| 5.25 | 0.56 | -3.49 | 11.76 | A | A | A |
| 5.343 | 0.333 | -1.78 | 7.97 | A | A | A |
| 5.15 | 0.363 | -5.33 | 8.62 | A | A | A |
| 5.12 | 0.06 | -5.88 | 5.10 | A | A | A |
| 5.69 | 0.28 | 4.60 | 6.99 | A | A | A |
| 5.05 | 0.5 | -7.17 | 11.08 | A | A | A |
| 5.94 | 0.30 | 9.19 | 7.08 | A | A | A |
| 5.9 | 0.9 | 8.46 | 16.04 | A | A | A |
| 4.89 | 0.30 | -10.11 | 7.89 | A | A | A |

Labcode

235

245

247

248

258

262

263

266

268

270

272

279

288

290

300

315

330

333

Reported Results for U-235 determination in Sample 1

TABLE 7. Target values

Parameter	Value

Activity Concentration [ Bq/kg ]	0.26

Uncertainty [ Bq/kg ]	0.01

Maximum Acceptable Relative Bias, [%]	30

Warning 1%

Not Accepted 6%

Activity concentration  [ Bq/kg ]

0.25

0.30

Accepted 19%

Not Reported 74%

0.35

0.40

FIGURE 7. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

0.10

0.15

0.20

126

140

166

185

223

245

130

186

85

266

127

57

270

225

129

232

188

1

263

89

68

102

123

189

300

145

180

272

268

21

104

135

178

235

215

215

203

333

279

290

119

43

54

44

112

81

288

19

35

172

40

67

248

63

105

27

199

258

79

72

315

330

198

184

183

171

131

85

106

247

69

262

202

299

167

247

197

136

285

210

210

75

322

Laboratory code

FIGURE 8. Reported results and evaluation parameters for U-235. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

| Reported value [Bq/kg ] 0.22 | Reported uncertainty (k = 1) [Bq/kg ] 0.01 | Relative bias [%] -15.38 | P-Test [%] 5.95 | Trueness evaluation A | Precision evaluation N | Final Score W |
| --- | --- | --- | --- | --- | --- | --- |
| 0.266 | 0.019 | 2.31 | 8.11 | A | A | A |
| 0.24 | 0.03 | -7.69 | 13.08 | A | A | A |
| 0.28 | 0.03 | 7.69 | 11.38 | A | A | A |
| 0.267 | 0.017 | 2.69 | 7.44 | A | A | A |
| 0.258 | 0.010 | -0.77 | 5.46 | A | A | A |
| 0.26 | 0.01 | 0.00 | 5.44 | A | A | A |
| 0.26 | 0.02 | 0.00 | 8.60 | A | A | A |
| 0.20 | 0.02 | -23.08 | 10.71 | A | A | A |
| 0.270 | 0.035 | 3.85 | 13.52 | A | A | A |
| 0.27 | 0.08 | 3.85 | 29.88 | A | A | A |
| 0.227 | 0.023 | -12.69 | 10.84 | A | A | A |
| 0.344 | 0.017 | 32.31 | 6.26 | N | N | N |
| 0.289 | 0.015 | 11.15 | 6.46 | A | A | A |
| 6.108 | 0.472 | 2249.23 | 8.63 | N | N | N |
| 0.2826 | 0.0316 | 8.69 | 11.82 | A | A | A |
| 0.263 | 0.016 | 1.15 | 7.20 | A | A | A |
| 0.19 | 0.01 | -26.92 | 6.52 | A | N | W |
| 0.22 | 0.02 | -15.38 | 9.87 | A | A | A |
| 0.23 | 0.11 | -11.54 | 47.98 | A | N | W |
| 0.27 | 0.03 | 3.85 | 11.76 | A | A | A |
| 0.32 | 0.04 | 23.08 | 13.08 | A | A | A |
| 0.26 | 0.01 | 0.00 | 5.44 | A | A | A |
| 0.258 | 0.04 | -0.77 | 15.97 | A | A | A |
| 0.23 | 0.01 | -11.54 | 5.80 | A | A | A |
| 0.0389 | 0.0003 | -85.04 | 3.92 | N | N | N |
| 0.192 | 0.0228 | -26.15 | 12.48 | A | A | A |
| 0.21 | 0.02 | -19.23 | 10.27 | A | A | A |
| 0.19 | 0.05 | -26.92 | 26.60 | A | A | A |
| 0.242 | 0.027 | -6.92 | 11.80 | A | A | A |
| 0.13 | 0.05 | -50.00 | 38.65 | N | N | N |
| 0.232 | 0.030 | -10.77 | 13.49 | A | A | A |
| 0.15 | 0.03 | -42.31 | 20.37 | N | A | N |
| 0.50 | 0.18 | 92.31 | 36.20 | N | N | N |
| 0.268 | 0.0130 | 3.08 | 6.19 | A | A | A |
| 0.246 | 0.015 | -5.38 | 7.21 | A | A | A |
| 0.233 | 0.026 | -10.38 | 11.80 | A | A | A |
| 0.30 | 0.03 | 15.38 | 10.71 | A | A | A |
| 0.153 | 0.029 | -41.15 | 19.34 | N | A | N |
| 0.19 | 0.02 | -26.92 | 11.21 | A | A | A |
| 0.22 | 0.02 | -15.38 | 9.87 | A | A | A |
| 0.23 | 0.03 | -11.54 | 13.60 | A | A | A |
| 0.3 | 0.06 | 15.38 | 20.37 | A | A | A |
| 0.28 | 0.04 | 7.69 | 14.79 | A | A | A |
| 0.39 | 0.06 | 50.00 | 15.86 | N | N | N |
| 0.250 | 0.004 | -3.85 | 4.17 | A | A | A |
| 5.45 | 1.09 | 1996.15 | 20.37 | N | N | N |
| 0.249 | 0.007 | -4.23 | 4.76 | A | A | A |
| 0.16 | 0.01 | -38.46 | 7.34 | N | N | N |
| 0.21 | 0.04 | -19.23 | 19.43 | A | A | A |
| 0.21 | 0.04 | -19.23 | 19.43 | A | A | A |
| 0.246 | 0.023 | -5.38 | 10.11 | A | A | A |
| 0.18 | 0.04 | -30.77 | 22.55 | N | A | N |
| 0.57 | 0.05 | 119.23 | 9.58 | N | N | N |
| 0.27 | 0.05 | 3.85 | 18.91 | A | A | A |
| 0.28 | 0.01 | 7.69 | 5.25 | A | A | A |
| 0.348 | 0.0382 | 33.85 | 11.63 | N | N | N |
| 0.22 | 0.02 | -15.38 | 9.87 | A | A | A |
| 0.192 | 0.0461 | -26.15 | 24.32 | A | A | A |
| 0.240 | 0.020 | -7.69 | 9.18 | A | A | A |

TABLE 8. Evaluation Results for U-235 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg by alpha

Labcode

1

19

21

27

35

43

44

54

57

63

67

68

69

72

75

79

81

85

89

102

105

106

112

119

123

126

127

129

130

135

140

145

166

167

172

178

180

184

185

186

188

189

198

199

202

203

210

215

223

225

232

235

245

247

248

258

262

263

266

268

TABLE 8. Evaluation Results for U-235 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg by alpha

| Reported value [Bq/kg ] 0.204 | Reported uncertainty (k = 1) [Bq/kg ] 0.028 | Relative bias [%] -21.54 | P-Test [%] 14.25 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 0.239 | 0.025 | -8.08 | 11.14 | A | A | A |
| 0.255 | 0.0538 | -1.92 | 21.45 | A | A | A |
| 0.263 | 0.01 | 1.15 | 5.41 | A | A | A |
| 0.256 | 0.013 | -1.54 | 6.37 | A | A | A |
| 0.23 | 0.04 | -11.54 | 17.81 | A | A | A |
| 0.29 | 0.02 | 11.54 | 7.90 | A | A | A |
| 0.30 | 0.04 | 15.38 | 13.88 | A | A | A |
| 0.25 | 0.02 | -3.85 | 8.88 | A | A | A |

Labcode

270

272

279

288

290

300

315

330

333

Activity concentration  [ Bq/kg ]

5

6

7

Reported Results for U-238 determination in Sample 1

TABLE 9. Target values

| Parameter | Value |
| --- | --- |
| Activity Concentration [ Bq/kg ] | 5.60 |
| Uncertainty [ Bq/kg ] | 0.28 |
| Maximum Acceptable Relative Bias, [%] | 30 |

Warning 2% Not Accepted 3%

Accepted 23%

Not Reported 72%

8

9

FIGURE 9. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

2

3

4

76

210

210

336

166

185

223

140

104

270

25

129

68

89

123

232

266

333

57

102

135

288

300

135

173

172

279

31

215

215

157

130

268

178

112

203

63

27

85

272

290

54

167

180

81

44

334

43

21

67

235

245

120

186

145

1

248

119

72

258

262

105

40

35

196

39

225

19

188

79

22

189

69

16

315

126

184

330

199

263

75

293

106

202

198

247

104

127

247

38

65

Laboratory code

FIGURE 10. Reported results and evaluation parameters for U-238. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

| Reported value [Bq/kg ] 5.64 | Reported uncertainty (k = 1) [Bq/kg ] 0.29 | Relative bias [%] 0.71 | P-Test [%] 7.17 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 6.03 | 0.01 | 7.68 | 5.00 | A | A | A |
| 5.84 | 0.30 | 4.29 | 7.17 | A | A | A |
| 5.61 | 0.42 | 0.18 | 9.00 | A | A | A |
| 5.91 | 0.54 | 5.54 | 10.42 | A | A | A |
| 4.5 | 0.7 | -19.64 | 16.34 | A | A | A |
| 5.49 | 0.13 | -1.96 | 5.53 | A | A | A |
| 5.3 |  | -5.36 |  | A |  |  |
| 5.79 | 0.37 | 3.39 | 8.11 | A | A | A |
| 5.82 | 0.30 | 3.93 | 7.18 | A | A | A |
| 5.61 | 0.22 | 0.18 | 6.35 | A | A | A |
| 5.58 | 0.28 | -0.36 | 7.08 | A | A | A |
| 5.54 | 0.30 | -1.07 | 7.37 | A | A | A |
| 5.00 | 0.27 | -10.71 | 7.36 | A | A | A |
| 5.47 | 0.60 | -2.32 | 12.05 | A | A | A |
| 86.2 | 24.14 | 1439.29 | 28.45 | N | N | N |
| 5.62 | 0.6 | 0.36 | 11.79 | A | A | A |
| 4.61 | 0.27 | -17.68 | 7.70 | A | A | A |
| 6.004 | 0.137 | 7.21 | 5.50 | A | A | A |
| 5.71 | 0.16 | 1.96 | 5.73 | A | A | A |
| 6.266 | 0.460 | 11.89 | 8.88 | A | A | A |
| 0.399 | 0.017 | -92.88 | 6.57 | N | N | N |
| 5.871 | 0.193 | 4.84 | 5.98 | A | A | A |
| 5.58 | 0.12 | -0.36 | 5.44 | A | A | A |
| 5.50 | 0.18 | -1.79 | 5.98 | A | A | A |
| 4.81 | 0.24 | -14.11 | 7.06 | A | A | A |
| 5.10 | 0.79 | -8.93 | 16.28 | A | A | A |
| 4.35 | 1.11 | -22.32 | 26.00 | A | A | A |
| 5.77 | 0.39 | 3.04 | 8.41 | A | A | A |
| 6.44 | 0.72 | 15.00 | 12.25 | A | A | A |
| 5.4 | 0.3 | -3.57 | 7.47 | A | A | A |
| 5.65 | 0.3 | 0.89 | 7.29 | A | A | A |
| 5.62 | 0.96 | 0.36 | 17.80 | A | A | A |
| 4.89 | 0.28 | -12.68 | 7.60 | A | A | A |
| 6.04 | 0.05 | 7.86 | 5.07 | A | A | A |
| 8.16 | 0.928 | 45.71 | 12.42 | N | N | N |
| 4.61 | 0.46 | -17.68 | 11.16 | A | A | A |
| 5.37 | 0.64 | -4.11 | 12.92 | A | A | A |
| 5.2 | 0.4 | -7.14 | 9.17 | A | A | A |
| 4.26 | 0.23 | -23.93 | 7.36 | A | N | W |
| 5.63 | 0.16 | 0.54 | 5.75 | A | A | A |
| 5.36 | 0.54 | -4.29 | 11.25 | A | A | A |
| 3.5 | 0.6 | -37.50 | 17.86 | N | A | N |
| 5.55 | 1.58 | -0.89 | 28.90 | A | A | A |
| 5.21 | 0.0526 | -6.96 | 5.10 | A | A | A |
| 5.2 | 0.2 | -7.14 | 6.31 | A | A | A |
| 5.39 | 0.17 | -3.75 | 5.91 | A | A | A |
| 5.56 | 0.21 | -0.71 | 6.27 | A | A | A |
| 6.09 | 0.61 | 8.75 | 11.20 | A | A | A |
| 3.52 | 0.60 | -37.14 | 17.76 | N | A | N |
| 5.63 | 0.16 | 0.54 | 5.75 | A | A | A |
| 5.86 | 0.21 | 4.64 | 6.15 | A | A | A |
| 6.0 | 0.6 | 7.14 | 11.18 | A | A | A |
| 5.8 | 0.5 | 3.57 | 9.97 | A | A | A |
| 6.8 | 0.4 | 21.43 | 7.72 | A | N | W |
| 6.24 | 0.18 | 11.43 | 5.77 | A | A | A |
| 6.67 | 0.22 | 19.11 | 5.99 | A | N | W |
| 5.44 | 0.09 | -2.86 | 5.27 | A | A | A |
| 0.44 | 0.17 | -92.14 | 38.96 | N | N | N |
| 5.35 | 0.12 | -4.46 | 5.48 | A | A | A |

TABLE 10. Evaluation Results for U-238 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg by alpha

Labcode

1

16

19

21

22

25

27

31

35

39

43

44

54

57

63

65

67

68

69

72

75

76

79

81

85

89

102

104

105

106

112

119

120

123

126

127

129

130

135

140

145

157

166

167

172

173

178

180

184

185

186

188

189

196

198

199

202

203

210

215

TABLE 10. Evaluation Results for U-238 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg by alpha

| Reported value [Bq/kg ] | Reported uncertainty (k = 1) [Bq/kg ] | Relative bias [%] | P-Test [%] | Trueness evaluation | Precision evaluation | Final Score |
| --- | --- | --- | --- | --- | --- | --- |
| 3.67 | 0.13 | -34.46 | 6.13 | N | N | N |
| 5.83 | 0.75 | 4.11 | 13.80 | A | A | A |
| 4.9 | 0.5 | -12.50 | 11.36 | A | A | A |
| 5.62 | 0.20 | 0.36 | 6.14 | A | A | A |
| 5.62 | 0.76 | 0.36 | 14.42 | A | A | A |
| 10.0 | 0.10 | 78.57 | 5.10 | N | N | N |
| 5.65 | 0.86 | 0.89 | 16.02 | A | A | A |
| 5.76 | 0.11 | 2.86 | 5.35 | A | A | A |
| 5.77 | 0.263 | 3.04 | 6.77 | A | A | A |
| 6.26 | 0.24 | 11.79 | 6.30 | A | A | A |
| 4.90 | 0.439 | -12.50 | 10.26 | A | A | A |
| 5.38 | 0.15 | -3.93 | 5.72 | A | A | A |
| 4.49 | 0.48 | -19.82 | 11.80 | A | A | A |
| 5.51 | 0.343 | -1.61 | 7.98 | A | A | A |
| 5.25 | 0.341 | -6.25 | 8.20 | A | A | A |
| 5.18 | 0.07 | -7.50 | 5.18 | A | A | A |
| 5.51 | 0.27 | -1.61 | 7.00 | A | A | A |
| 5.2 | 0.5 | -7.14 | 10.84 | A | A | A |
| 6.03 | 0.31 | 7.68 | 7.17 | A | A | A |
| 6.1 | 0.9 | 8.93 | 15.58 | A | A | A |
| 4.99 | 0.3 | -10.89 | 7.82 | A | A | A |

Labcode

223

225

232

235

245

247

248

258

262

263

266

268

270

272

279

288

290

300

315

330

333

TABLE 11. Evaluation Results for U-238 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg by gamma

| Reported value [Bq/kg ] | Reported uncertainty (k = 1) [Bq/kg ] | Relative bias [%] | P-Test [%] | Trueness evaluation | Precision evaluation | Final Score |
| --- | --- | --- | --- | --- | --- | --- |
| 12.2 | 1.9 | 117.86 | 16.36 | N | N | N |
| 5.77 | 0.453 | 3.04 | 9.31 | A | A | A |
| 7.24 | 2.43 | 29.29 | 33.93 | A | N | W |
| 5.1 | 0.8 | -8.93 | 16.46 | A | A | A |
| 0.44 | 0.17 | -92.14 | 38.96 | N | N | N |
| 5.35 | 0.11 | -4.46 | 5.41 | A | A | A |
| 6.96 | 0.51 | 24.29 | 8.87 | A | N | W |
| 6.28 | 0.44 | 12.14 | 8.61 | A | A | A |
| 5.6 | 0.8 | 0.00 | 15.14 | A | A | A |
| 3.06 | 0.20 | -45.36 | 8.23 | N | N | N |

Labcode

38

40

104

135

210

215

247

293

334

336

Activity concentration  [ Bq/kg ]

25

Reported Results for Am-241 determination in Sample 1

TABLE 12. Target values

| Parameter | Value |
| --- | --- |
| Activity Concentration [ Bq/kg ] | 22.3 |
| Uncertainty [ Bq/kg ] | 1.1 |
| Maximum Acceptable Relative Bias, [%] | 20 |

Accepted 74%

Not Reported 17%

Warning 3%

Not Accepted 6%

30

FIGURE 11. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

15

20

92

282

152

93

305

110

97

192

197

286

46

17

227

135

315

283

59

308

223

104

262

170

137

277

331

112

241

274

244

249

25

330

22

271

260

285

131

44

12

167

318

11

72

178

140

107

132

237

314

148

35

269

123

159

247

119

239

304

95

254

251

342

30

145

63

134

183

27

300

81

232

64

276

71

18

49

125

6

179

316

50

3

200

275

77

120

290

190

166

234

36

257

228

306

327

186

284

233

279

209

26

222

45

180

67

68

199

115

177

21

20

163

66

79

133

70

171

34

2

225

264

1

248

153

15

173

147

256

57

149

129

108

224

294

Laboratory code

FIGURE 12. Reported results and evaluation parameters for Am-241. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

| Reported value [Bq/kg ] 23.9 | Reported uncertainty (k = 1) [Bq/kg ] 1.0 | Relative bias [%] 7.17 | P-Test [%] 6.47 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 23.66 | 1.63 | 6.10 | 8.47 | A | A | A |
| 22.6 | 2.52 | 1.35 | 12.19 | A | A | A |
| 22.15 | 1.48 | -0.67 | 8.31 | A | A | A |
| 22.5 | 1.2 | 0.90 | 7.26 | A | A | A |
| 43.96 | 4.4 | 97.13 | 11.16 | N | N | N |
| 22.51 | 1.89 | 0.94 | 9.74 | A | A | A |
| 22.9 | 0.7 | 2.69 | 5.80 | A | A | A |
| 21.5 | 1.7 | -3.59 | 9.32 | A | A | A |
| 21.4 | 1.0 | -4.04 | 6.79 | A | A | A |
| 23.8 | 1.9 | 6.73 | 9.38 | A | A | A |
| 23.0 | 1.5 | 3.14 | 8.18 | A | A | A |
| 24.0 | 0.9 | 7.62 | 6.20 | A | A | A |
| 23.34 | 2.12 | 4.66 | 10.34 | A | A | A |
| 19.5 | 1.2 | -12.56 | 7.89 | A | A | A |
| 22.41 | 1.46 | 0.49 | 8.17 | A | A | A |
| 20.3 | 1.8 | -8.97 | 10.15 | A | A | A |
| 23.3 | 2.59 | 4.48 | 12.16 | A | A | A |
| 23.3 | 0.9 | 4.48 | 6.27 | A | A | A |
| 21.22 | 0.98 | -4.84 | 6.76 | A | A | A |
| 22.1 | 0.7 | -0.90 | 5.86 | A | A | A |
| 21.19 | 2.2 | -4.98 | 11.49 | A | A | A |
| 23.1 | 1.52 | 3.59 | 8.22 | A | A | A |
| 22.29 | 1.32 | -0.04 | 7.71 | A | A | A |
| 21.2 | 2.3 | -4.93 | 11.92 | A | A | A |
| 21.7 | 1.2 | -2.69 | 7.41 | A | A | A |
| 22.2 | 1.1 | -0.45 | 6.99 | A | A | A |
| 20.3 | 1.3 | -8.97 | 8.08 | A | A | A |
| 21.33 | 0.35 | -4.35 | 5.20 | A | A | A |
| 23.7 | 3.5 | 6.28 | 15.57 | A | A | A |
| 23.6 | 4.25 | 5.83 | 18.67 | A | A | A |
| 21.8 | 1.8 | -2.24 | 9.62 | A | A | A |
| 22.9 | 1.8 | 2.69 | 9.28 | A | A | A |
| 22.7 | 1.3 | 1.79 | 7.56 | A | A | A |
| 22.4 | 2.5 | 0.45 | 12.20 | A | A | A |
| 21.3 | 0.9 | -4.48 | 6.50 | A | A | A |
| 22.2 | 1.45 | -0.45 | 8.18 | A | A | A |
| 21.6 | 3.03 | -3.14 | 14.87 | A | A | A |
| 22.8 | 1.7 | 2.24 | 8.94 | A | A | A |
| 21.4 | 0.5 | -4.04 | 5.46 | A | A | A |
| 23.1 | 1.54 | 3.59 | 8.29 | A | A | A |
| 19.4 | 0.2 | -13.00 | 5.04 | A | A | A |
| 21.4 | 0.7 | -4.04 | 5.92 | A | A | A |
| 22.9 | 1.2 | 2.69 | 7.20 | A | A | A |
| 22.45 | 0.62 | 0.67 | 5.65 | A | A | A |
| 22.55 | 0.61 | 1.12 | 5.63 | A | A | A |
| 23.7 | 1.5 | 6.28 | 8.02 | A | A | A |
| 22.8 | 0.8 | 2.24 | 6.05 | A | A | A |
| 19.3 | 0.6 | -13.45 | 5.83 | A | A | A |
| 23.4 | 2.9 | 4.93 | 13.34 | A | A | A |
| 24.7 | 2.7 | 10.76 | 11.99 | A | A | A |
| 22.50 | 3.59 | 0.90 | 16.70 | A | A | A |
| 20.3 | 1.12 | -8.97 | 7.40 | A | A | A |
| 21.88 | 1.31 | -1.88 | 7.76 | A | A | A |
| 22.1 | 1.6 | -0.90 | 8.76 | A | A | A |
| 22.2 | 0.67 | -0.45 | 5.78 | A | A | A |
| 22.32 | 2.02 | 0.09 | 10.31 | A | A | A |
| 23.4 | 1.6 | 4.93 | 8.43 | A | A | A |
| 23.2 | 0.8 | 4.04 | 6.02 | A | A | A |
| 23.2 | 1.5 | 4.04 | 8.13 | A | A | A |

TABLE 13. Evaluation Results for Am-241 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

1

2

3

5

6

7

8

10

11

12

13

14

15

16

17

18

19

20

21

22

23

25

26

27

28

29

30

31

32

33

34

35

36

37

38

39

40

42

43

44

45

46

47

48

49

50

51

52

54

55

57

58

59

60

62

63

64

66

67

68

| Reported value [Bq/kg ] 23.5 | Reported uncertainty (k = 1) [Bq/kg ] 0.2 | Relative bias [%] 5.38 | P-Test [%] 5.01 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 22.4 | 1 | 0.45 | 6.65 | A | A | A |
| 21.51 | 0.30 | -3.54 | 5.13 | A | A | A |
| 22.3 | 1.7 | 0.00 | 9.08 | A | A | A |
| 21.02 | 0.38 | -5.74 | 5.25 | A | A | A |
| 22.7 | 1.0 | 1.79 | 6.61 | A | A | A |
| 23.4 | 1.2 | 4.93 | 7.12 | A | A | A |
| 15.9 | 0.8 | -28.70 | 7.05 | N | N | N |
| 22.3 | 0.7 | 0.00 | 5.85 | A | A | A |
| 21.22 | 1.14 | -4.84 | 7.29 | A | A | A |
| 23.62 | 1.78 | 5.92 | 9.01 | A | A | A |
| 22.3 | 1.49 | 0.00 | 8.31 | A | A | A |
| 22.86 | 1.39 | 2.51 | 7.83 | A | A | A |
| 23.5 | 1.3 | 5.38 | 7.41 | A | A | A |
| 21.44 | 1.02 | -3.86 | 6.85 | A | A | A |
| 21.13 | 0.95 | -5.25 | 6.67 | A | A | A |
| 1.74 | 0.22 | -92.20 | 13.57 | N | N | N |
| 13.40 | 1.51 | -39.91 | 12.30 | N | N | N |
| 20.85 | 0.95 | -6.50 | 6.72 | A | A | A |
| 22.1 | 1.5 | -0.90 | 8.39 | A | A | A |
| 17.88 | 0.54 | -19.82 | 5.78 | A | N | W |
| 22.44 | 1.58 | 0.63 | 8.60 | A | A | A |
| 20.94 | 0.81 | -6.10 | 6.27 | A | A | A |
| 30.7 |  | 37.67 |  | N |  | N |
| 21.8 | 1.2 | -2.24 | 7.39 | A | A | A |
| 20.54 | 2.15 | -7.89 | 11.57 | A | A | A |
| 23.13 | 1.61 | 3.72 | 8.53 | A | A | A |
| 23.21 | 1.25 | 4.08 | 7.30 | A | A | A |
| 21.6 | 1.4 | -3.14 | 8.15 | A | A | A |
| 31.8 | 2 | 42.60 | 7.99 | N | N | N |
| 17.12 | 2.21 | -23.23 | 13.82 | N | A | N |
| 21 | 1 | -5.83 | 6.86 | A | A | A |
| 22.5 | 1.6 | 0.90 | 8.65 | A | A | A |
| 23.22 | 0.23 | 4.13 | 5.03 | A | A | A |
| 22.18 | 4.00 | -0.54 | 18.70 | A | A | A |
| 21.5 | 2.0 | -3.59 | 10.53 | A | A | A |
| 22 | 0.8 | -1.35 | 6.13 | A | A | A |
| 22.75 | 0.91 | 2.02 | 6.35 | A | A | A |
| 23.0 | 0.8 | 3.14 | 6.04 | A | A | A |
| 21.9 | 1.3 | -1.79 | 7.72 | A | A | A |
| 22.46 | 2.53 | 0.72 | 12.30 | A | A | A |
| 16.7 | 0.5 | -25.11 | 5.77 | N | N | N |
| 27.86 | 4.70 | 24.93 | 17.58 | N | A | N |
| 22.8 | 0.8 | 2.24 | 6.05 | A | A | A |
| 21.34 | 0.94 | -4.30 | 6.61 | A | A | A |
| 21.7 | 2.0 | -2.69 | 10.45 | A | A | A |
| 23.4 | 2.4 | 4.93 | 11.38 | A | A | A |
| 22.2 | 0.8 | -0.45 | 6.11 | A | A | A |
| 19.8 | 2.6 | -11.21 | 14.03 | A | A | A |
| 23 | 2 | 3.14 | 10.00 | A | A | A |
| 20.85 | 0.95 | -6.50 | 6.72 | A | A | A |
| 22.1 | 0.67 | -0.90 | 5.79 | A | A | A |
| 21.6 | 0.9 | -3.14 | 6.46 | A | A | A |
| 20.57 | 2.33 | -7.76 | 12.35 | A | A | A |
| 20.482 | 4.741 | -8.15 | 23.67 | A | N | W |
| 23.3 | 1.2 | 4.48 | 7.13 | A | A | A |
| 22.2 | 1.4 | -0.45 | 8.01 | A | A | A |
| 24.34 | 1.36 | 9.15 | 7.45 | A | A | A |
| 24.37 | 0.73 | 9.28 | 5.77 | A | A | A |
| 21.8 | 1.8 | -2.24 | 9.62 | A | A | A |

TABLE 13. Evaluation Results for Am-241 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

70

71

72

73

75

77

79

80

81

82

83

84

85

86

89

90

92

93

94

95

97

98

99

101

102

104

105

106

107

108

110

112

114

115

116

118

119

120

122

123

125

126

129

130

131

132

133

134

135

136

137

139

140

141

142

144

145

146

147

148

| Reported value [Bq/kg ] 26.7 | Reported uncertainty (k = 1) [Bq/kg ] 1.11 | Relative bias [%] 19.73 | P-Test [%] 6.45 | Trueness evaluation A | Precision evaluation N | Final Score W |
| --- | --- | --- | --- | --- | --- | --- |
| 22.3 | 1.3 | 0.00 | 7.64 | A | A | A |
| 11.6 |  | -47.98 |  | N |  | N |
| 23.952 | 2.018 | 7.41 | 9.76 | A | A | A |
| 17.89 | 0.46 | -19.78 | 5.56 | A | N | W |
| 22.2 | 1.0 | -0.45 | 6.68 | A | A | A |
| 23.0 | 1.2 | 3.14 | 7.18 | A | A | A |
| 23.3 | 0.7 | 4.48 | 5.78 | A | A | A |
| 22.03 | 0.66 | -1.21 | 5.77 | A | A | A |
| 22.0 | 1.3 | -1.35 | 7.70 | A | A | A |
| 22.7 | 2.0 | 1.79 | 10.10 | A | A | A |
| 23.03 | 1.7 | 3.27 | 8.88 | A | A | A |
| 23.92 | 0.66 | 7.26 | 5.65 | A | A | A |
| 23.4 | 2.3 | 4.93 | 11.00 | A | A | A |
| 23.49 | 0.32 | 5.34 | 5.12 | A | A | A |
| 22.8 | 2.4 | 2.24 | 11.62 | A | A | A |
| 21.50 | 1.16 | -3.59 | 7.31 | A | A | A |
| 21.4 | 1.2 | -4.04 | 7.47 | A | A | A |
| 22.0 | 2.6 | -1.35 | 12.81 | A | A | A |
| 20.8 | 2.1 | -6.73 | 11.24 | A | A | A |
| 23.5 | 0.82 | 5.38 | 6.04 | A | A | A |
| 19.7 | 2.36 | -11.66 | 12.96 | A | A | A |
| 24.012 | 2.42 | 7.68 | 11.22 | A | A | A |
| 24 | 2 | 7.62 | 9.68 | A | A | A |
| 17.4 | 2.1 | -21.97 | 13.04 | N | A | N |
| 23.25 | 1.77 | 4.26 | 9.07 | A | A | A |
| 21.57 | 0.47 | -3.27 | 5.39 | A | A | A |
| 22.5 | 1.2 | 0.90 | 7.26 | A | A | A |
| 23.16 | 0.86 | 3.86 | 6.17 | A | A | A |
| 23.9 | 1.2 | 7.17 | 7.04 | A | A | A |
| 22.23 | 2.67 | -0.31 | 12.98 | A | A | A |
| 22.2 | 1.5 | -0.45 | 8.37 | A | A | A |
| 22.9 | 2.5 | 2.69 | 11.98 | A | A | A |
| 22.81 | 2.40 | 2.29 | 11.62 | A | A | A |
| 23.0 | 1.1 | 3.14 | 6.87 | A | A | A |
| 22 | 1 | -1.35 | 6.71 | A | A | A |
| 21 | 2.3 | -5.83 | 12.01 | A | A | A |
| 22.8 | 1.7 | 2.24 | 8.94 | A | A | A |
| 18.2 | 1.3 | -18.39 | 8.68 | A | A | A |
| 24.38 | 0.822 | 9.33 | 5.97 | A | A | A |
| 18.7953 | 1.5084 | -15.72 | 9.42 | A | A | A |
| 22.2 | 0.6 | -0.45 | 5.62 | A | A | A |
| 23.21 | 0.96 | 4.08 | 6.44 | A | A | A |
| 22.67 | 1.24 | 1.66 | 7.37 | A | A | A |
| 22.65 | 1.61 | 1.57 | 8.65 | A | A | A |
| 22.3 | 1.7 | 0.00 | 9.08 | A | A | A |
| 9.74 | 0.22 | -56.32 | 5.43 | N | N | N |
| 21.564 | 1.669 | -3.30 | 9.18 | A | A | A |
| 18.83 | 2.04 | -15.56 | 11.90 | A | A | A |
| 23.10 | 2.14 | 3.59 | 10.50 | A | A | A |
| 23.40 | 2.30 | 4.93 | 11.00 | A | A | A |
| 23.1 | 1.2 | 3.59 | 7.16 | A | A | A |
| 21.5 | 1.3 | -3.59 | 7.80 | A | A | A |
| 21.1 | 2.2 | -5.38 | 11.53 | A | A | A |
| 21.8 | 1.15 | -2.24 | 7.22 | A | A | A |
| 18.30 | 0.30 | -17.94 | 5.20 | A | N | W |
| 23.10 | 2.34 | 3.59 | 11.27 | A | A | A |
| 20.50 | 2.11 | -8.07 | 11.41 | A | A | A |
| 35.2 | 0.11 | 57.85 | 4.94 | N | N | N |
| 23.7 | 1.0 | 6.28 | 6.49 | A | A | A |

TABLE 13. Evaluation Results for Am-241 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

149

151

152

153

154

155

156

157

158

159

160

161

162

163

165

166

167

168

169

170

171

172

173

174

176

177

178

179

180

181

182

183

184

185

186

188

189

190

192

196

197

198

199

200

201

204

205

206

207

209

210

213

214

217

218

221

222

223

224

225

| Reported value [Bq/kg ] 26.56 | Reported uncertainty (k = 1) [Bq/kg ] 0.2 | Relative bias [%] 19.10 | P-Test [%] 4.99 | Trueness evaluation A | Precision evaluation N | Final Score W |
| --- | --- | --- | --- | --- | --- | --- |
| 19.6 | 0.2 | -12.11 | 5.04 | A | A | A |
| 22.9 | 0.6 | 2.69 | 5.59 | A | A | A |
| 22.3 | 1.2 | 0.00 | 7.30 | A | A | A |
| 23.02 | 1.5 | 3.23 | 8.17 | A | A | A |
| 22.84 | 0.72 | 2.42 | 5.85 | A | A | A |
| 23.1 | 1.2 | 3.59 | 7.16 | A | A | A |
| 24.5 | 1.28 | 9.87 | 7.19 | A | A | A |
| 21.73 | 0.35 | -2.56 | 5.19 | A | A | A |
| 22.07 | 1.08 | -1.03 | 6.95 | A | A | A |
| 21.6 | 1.2 | -3.14 | 7.43 | A | A | A |
| 21.0 | 1.7 | -5.83 | 9.48 | A | A | A |
| 24 | 2.4 | 7.62 | 11.15 | A | A | A |
| 21.1 | 1.72 | -5.38 | 9.53 | A | A | A |
| 22.2 | 1.4 | -0.45 | 8.01 | A | A | A |
| 23.55 | 0.78 | 5.61 | 5.94 | A | A | A |
| 22.0 | 0.46 | -1.35 | 5.36 | A | A | A |
| 23.9 | 1.5 | 7.17 | 7.98 | A | A | A |
| 21.1 | 2.2 | -5.38 | 11.53 | A | A | A |
| 22.16 | 1.84 | -0.63 | 9.66 | A | A | A |
| 22.2 | 2.0 | -0.45 | 10.27 | A | A | A |
| 22.1 | 0.44 | -0.90 | 5.32 | A | A | A |
| 20.529 | 1.3297 | -7.94 | 8.14 | A | A | A |
| 24.5 | 1.34 | 9.87 | 7.37 | A | A | A |
| 22.9 | 1.0 | 2.69 | 6.59 | A | A | A |
| 19.40 | 1.00 | -13.00 | 7.13 | A | A | A |
| 21.9 | 2 | -1.79 | 10.38 | A | A | A |
| 21.30 | 1.60 | -4.48 | 8.99 | A | A | A |
| 21.7 | 1.1 | -2.69 | 7.07 | A | A | A |
| 20.6 | 1.23 | -7.62 | 7.74 | A | A | A |
| 21.8 | 0.3 | -2.24 | 5.12 | A | A | A |
| 23.71 | 2.160 | 6.32 | 10.36 | A | A | A |
| 27.0 | 2.3 | 21.08 | 9.84 | N | A | N |
| 21.83 | 0.63 | -2.11 | 5.71 | A | A | A |
| 23.4 | 1.3 | 4.93 | 7.43 | A | A | A |
| 21.3 | 1.1 | -4.48 | 7.14 | A | A | A |
| 23.2 | 1.0 | 4.04 | 6.55 | A | A | A |
| 21.2 | 1.3 | -4.93 | 7.87 | A | A | A |
| 21 | 1.51 | -5.83 | 8.72 | A | A | A |
| 22.7 | 2.6 | 1.79 | 12.47 | A | A | A |
| 22.4 | 0.8 | 0.45 | 6.09 | A | A | A |
| 20.90 | 0.56 | -6.28 | 5.61 | A | A | A |
| 23.04 | 2.48 | 3.32 | 11.84 | A | A | A |
| 21.51 | 2.37 | -3.54 | 12.07 | A | A | A |
| 9.0 | 0.45 | -59.64 | 7.02 | N | N | N |
| 20.27 | 3.241 | -9.10 | 16.73 | A | A | A |
| 23.0 | 1.5 | 3.14 | 8.18 | A | A | A |
| 21.305 | 0.83549 | -4.46 | 6.30 | A | A | A |
| 19.1 | 1.3 | -14.35 | 8.41 | A | A | A |
| 21.0 | 0.27 | -5.83 | 5.10 | A | A | A |
| 22.8 | 0.30 | 2.24 | 5.11 | A | A | A |
| 22.4 | 1.59 | 0.45 | 8.64 | A | A | A |
| 19.87 | 1.17 | -10.90 | 7.68 | A | A | A |
| 23.2 | 1.1 | 4.04 | 6.84 | A | A | A |
| 44.7 | 2.7 | 100.45 | 7.80 | N | N | N |
| 22.4 | 3.1 | 0.45 | 14.69 | A | A | A |
| 22.45 | 1.35 | 0.67 | 7.78 | A | A | A |
| 20.7 | 1.0 | -7.17 | 6.90 | A | A | A |
| 21.3 | 1.3 | -4.48 | 7.85 | A | A | A |
| 22.3 | 2.6 | 0.00 | 12.66 | A | A | A |

TABLE 13. Evaluation Results for Am-241 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

226

227

228

232

233

234

235

236

237

239

240

241

243

244

245

246

247

248

249

251

252

254

255

256

257

258

259

260

261

262

263

264

268

269

270

271

272

273

274

275

276

277

279

281

282

283

284

285

286

288

290

291

292

293

294

295

296

298

299

300

TABLE 13. Evaluation Results for Am-241 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg by gamma

| Reported value [Bq/kg ] 20.14 | Reported uncertainty (k = 1) [Bq/kg ] 0.71 | Relative bias [%] -9.69 | P-Test [%] 6.06 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 23.1 | 0.5 | 3.59 | 5.39 | A | A | A |
| 23.18 | 5.44 | 3.95 | 23.98 | A | N | W |
| 22.1 | 1.4 | -0.90 | 8.03 | A | A | A |
| 16.6 | 0.7 | -25.56 | 6.49 | N | N | N |
| 22.97 | 3.50 | 3.00 | 16.02 | A | A | A |
| 22.58 | 0.27 | 1.26 | 5.08 | A | A | A |
| 20.38 | 0.25 | -8.61 | 5.08 | A | A | A |
| 19.52 | 1.51 | -12.47 | 9.17 | A | A | A |
| 34.7 | 4.0 | 55.61 | 12.54 | N | N | N |
| 21.8 | 3.5 | -2.24 | 16.80 | A | A | A |
| 20.1 | 1.5 | -9.87 | 8.95 | A | A | A |
| 22.5 | 0.7 | 0.90 | 5.83 | A | A | A |
| 20.84 | 0.34 | -6.55 | 5.20 | A | A | A |
| 21.5 | 0.7 | -3.59 | 5.91 | A | A | A |
| 21.79 | 8.405 | -2.29 | 38.89 | A | N | W |
| 12.78 | 1.11 | -42.69 | 9.99 | N | N | N |
| 23.0 | 0.6 | 3.14 | 5.58 | A | A | A |
| 23.1 | 1.0 | 3.59 | 6.56 | A | A | A |
| 21.2 | 2.1 | -4.93 | 11.07 | A | A | A |
| 20.96 | 0.48 | -6.01 | 5.44 | A | A | A |
| 21.0 | 1.6 | -5.83 | 9.08 | A | A | A |
| 22.5 | 1.6 | 0.90 | 8.65 | A | A | A |
| 5.96 | 0.45 | -73.27 | 9.02 | N | N | N |
| 23.25 | 0.17 | 4.26 | 4.99 | A | A | A |
| 22.73 | 1.47 | 1.93 | 8.13 | A | A | A |
| 22.18 | 0.35 | -0.54 | 5.18 | A | A | A |

Labcode

301

302

303

304

305

306

307

308

309

311

314

315

316

317

318

319

322

327

328

330

331

333

334

336

338

341

342

Activity concentration  [ Bq/kg ]

15

Reported Results for Sr-90 determination in Sample 2

TABLE 14. Target values

| Parameter | Value |
| --- | --- |
| Activity Concentration [ Bq/kg ] | 14.2 |
| Uncertainty [ Bq/kg ] | 0.7 |
| Maximum Acceptable Relative Bias, [%] | 30 |

Not Accepted 5%

Warning 2%

Accepted 32%

Not Reported 61%

20

FIGURE 13. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

5

10

302

63

104

97

157

73

21

56

116

34

152

192

215

215

237

135

119

54

166

185

25

263

296

32

72

67

270

35

193

133

112

195

169

44

47

14

206

23

151

137

167

43

94

279

328

95

196

201

277

114

284

11

262

248

218

268

40

100

28

203

3

291

22

199

180

17

125

149

102

1

106

327

245

232

255

145

184

79

300

150

58

120

82

175

89

96

141

264

46

298

105

5

57

188

26

287

85

99

19

186

126

293

129

334

290

75

189

198

45

330

317

87

247

202

42

4

295

342

341

83

205

27

289

304

Laboratory code

FIGURE 14. Reported results and evaluation parameters for Sr-90. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

| Reported value [Bq/kg ] 14.07 | Reported uncertainty (k = 1) [Bq/kg ] 2.34 | Relative bias [%] -0.92 | P-Test [%] 17.35 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 13.9 | 0.61 | -2.11 | 6.60 | A | A | A |
| 18.65 | 1.10 | 31.34 | 7.69 | N | N | N |
| 14.79 | 0.49 | 4.15 | 5.94 | A | A | A |
| 13.76 | 1.03 | -3.10 | 8.96 | A | A | A |
| 13.44 | 0.54 | -5.35 | 6.36 | A | A | A |
| 14.0 | 0.5 | -1.41 | 6.09 | A | A | A |
| 15.3 | 0.8 | 7.75 | 7.19 | A | A | A |
| 9.4 | 1.09 | -33.80 | 12.60 | N | N | N |
| 13.91 | 0.74 | -2.04 | 7.25 | A | A | A |
| 13.45 | 0.54 | -5.28 | 6.36 | A | A | A |
| 12.66 | 2.0 | -10.85 | 16.55 | A | A | A |
| 14.9 | 0.95 | 4.93 | 8.06 | A | A | A |
| 31.0 | 2.6 | 118.31 | 9.73 | N | N | N |
| 13.9 | 0.98 | -2.11 | 8.60 | A | A | A |
| 12.82 | 0.42 | -9.72 | 5.92 | A | A | A |
| 11.8 | 1.9 | -16.90 | 16.84 | A | A | A |
| 13.1 | 1.0 | -7.75 | 9.09 | A | A | A |
| 13.84 | 0.78 | -2.54 | 7.49 | A | A | A |
| 18.5 | 3.61 | 30.28 | 20.13 | N | A | N |
| 13.5 | 0.7 | -4.93 | 7.15 | A | A | A |
| 13.4 | 0.6 | -5.63 | 6.66 | A | A | A |
| 16.7 | 1.74 | 17.61 | 11.53 | A | A | A |
| 14.7 | 0.7 | 3.52 | 6.85 | A | A | A |
| 13.4 | 0.6 | -5.63 | 6.66 | A | A | A |
| 12.5 | 0.2 | -11.97 | 5.18 | A | A | A |
| 10.2 | 0.8 | -28.17 | 9.26 | A | N | W |
| 14.8 | 1.5 | 4.23 | 11.27 | A | A | A |
| 14.30 | 2.56 | 0.70 | 18.57 | A | A | A |
| 3.29 | 0.16 | -76.83 | 6.92 | N | N | N |
| 12.87 | 1.01 | -9.37 | 9.27 | A | A | A |
| 12.85 | 0.24 | -9.51 | 5.27 | A | A | A |
| 9.2 | 1.6 | -35.21 | 18.08 | N | A | N |
| 16.20 | 1.53 | 14.08 | 10.65 | A | A | A |
| 14.23 | 0.71 | 0.21 | 7.01 | A | A | A |
| 14.32 | 0.79 | 0.85 | 7.40 | A | A | A |
| 22.16 | 1.56 | 56.06 | 8.59 | N | N | N |
| 14.94 | 0.43 | 5.21 | 5.71 | A | A | A |
| 17.26 | 0.79 | 21.55 | 6.73 | A | N | W |
| 14.51 | 0.8 | 2.18 | 7.40 | A | A | A |
| 13.50 | 0.72 | -4.93 | 7.26 | A | A | A |
| 13.6 | 0.7 | -4.23 | 7.13 | A | A | A |
| 14.52 | 0.83 | 2.25 | 7.55 | A | A | A |
| 15.18 | 0.48 | 6.90 | 5.86 | A | A | A |
| 13.85 | 0.81 | -2.46 | 7.65 | A | A | A |
| 14.03 | 4.94 | -1.20 | 35.55 | A | N | W |
| 7.11 | 2.88 | -49.93 | 40.81 | N | N | N |
| 14.72 | 1.50 | 3.66 | 11.32 | A | A | A |
| 14.1 | 1.0 | -0.70 | 8.64 | A | A | A |
| 13.2 | 0.8 | -7.04 | 7.81 | A | A | A |
| 13.7 | 0.85 | -3.52 | 7.92 | A | A | A |
| 11.67 | 0.36 | -17.82 | 5.82 | A | N | W |
| 12.3 | 1.2 | -13.38 | 10.93 | A | A | A |
| 14.3 | 1.33 | 0.70 | 10.53 | A | A | A |
| 14.01 | 2.82 | -1.34 | 20.72 | A | A | A |
| 15.9 | 0.09 | 11.97 | 4.96 | A | A | A |
| 15.94 | 0.37 | 12.25 | 5.45 | A | A | A |
| 13.2 | 0.48 | -7.04 | 6.13 | A | A | A |
| 12.1 | 1.8 | -14.79 | 15.67 | A | A | A |
| 13.50 | 0.72 | -4.93 | 7.26 | A | A | A |

TABLE 15. Evaluation Results for Sr-90 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg by beta

Labcode

1

3

4

5

11

14

17

19

21

22

23

25

26

27

28

32

34

35

40

42

43

44

45

46

47

54

56

57

58

63

67

72

73

75

79

82

83

85

87

89

94

95

96

99

100

102

104

105

106

112

114

116

119

120

125

126

129

133

135

137

| Reported value [Bq/kg ] 14.69 | Reported uncertainty (k = 1) [Bq/kg ] 0.46 | Relative bias [%] 3.45 | P-Test [%] 5.84 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 14.2 | 0.8 | 0.00 | 7.49 | A | A | A |
| 14.02 | 0.22 | -1.27 | 5.17 | A | A | A |
| 14.3 | 1.2 | 0.70 | 9.73 | A | A | A |
| 13.46 | 0.53 | -5.21 | 6.31 | A | A | A |
| 11.8 | 3.1 | -16.90 | 26.73 | A | A | A |
| 7.66 | 1.07 | -46.06 | 14.81 | N | N | N |
| 12.6 | 2.3 | -11.27 | 18.91 | A | A | A |
| 13.50 | 0.52 | -4.93 | 6.26 | A | A | A |
| 13.4 | 1.3 | -5.63 | 10.88 | A | A | A |
| 14.4 | 0.2 | 1.41 | 5.12 | A | A | A |
| 13.95 | 3.57 | -1.76 | 26.06 | A | A | A |
| 14.2 | 1.4 | 0.00 | 11.02 | A | A | A |
| 12.61 | 2.27 | -11.20 | 18.66 | A | A | A |
| 15.4 | 0.5 | 8.45 | 5.90 | A | A | A |
| 14.8 | 0.8 | 4.23 | 7.32 | A | A | A |
| 16.3 | 1.3 | 14.79 | 9.38 | A | A | A |
| 12.0 | 0.53 | -15.49 | 6.62 | A | A | A |
| 13.16 | 0.66 | -7.32 | 7.03 | A | A | A |
| 13.3 | 1.1 | -6.34 | 9.63 | A | A | A |
| 13.62 | 0.99 | -4.08 | 8.78 | A | A | A |
| 16.4 | 1.1 | 15.49 | 8.32 | A | A | A |
| 13.94 | 3.16 | -1.83 | 23.20 | A | A | A |
| 13.7 | 0.7 | -3.52 | 7.10 | A | A | A |
| 17.62 | 1.11 | 24.08 | 8.00 | A | N | W |
| 13.9 | 1.0 | -2.11 | 8.72 | A | A | A |
| 28.90 | 0.61 | 103.52 | 5.36 | N | N | N |
| 13.44 | 0.08 | -5.35 | 4.97 | A | A | A |
| 12 | 3 | -15.49 | 25.48 | A | A | A |
| 13.8 | 0.37 | -2.82 | 5.61 | A | A | A |
| 14.2 | 1.8 | 0.00 | 13.60 | A | A | A |
| 12.04 | 0.46 | -15.21 | 6.24 | A | A | A |
| 14.15 | 0.62 | -0.35 | 6.60 | A | A | A |
| 17.45 | 0.82 | 22.89 | 6.81 | A | N | W |
| 13.8 | 2.3 | -2.82 | 17.38 | A | A | A |
| 14.2 | 0.94 | 0.00 | 8.25 | A | A | A |
| 13.8 | 0.325 | -2.82 | 5.46 | A | A | A |
| 12.72 | 0.51 | -10.42 | 6.35 | A | A | A |
| 14.7 | 1.42 | 3.52 | 10.84 | A | A | A |
| 13.8 | 0.85 | -2.82 | 7.89 | A | A | A |
| 13.08 | 0.71 | -7.89 | 7.33 | A | A | A |
| 13.70 | 0.71 | -3.52 | 7.15 | A | A | A |
| 13.6 | 0.68 | -4.23 | 7.02 | A | A | A |
| 13.7 | 1.0 | -3.52 | 8.81 | A | A | A |
| 14.9 | 1.12 | 4.93 | 8.99 | A | A | A |
| 104.89 | 8.85 | 638.66 | 9.77 | N | N | N |
| 16.0 | 1.28 | 12.68 | 9.40 | A | A | A |
| 13.9 | 0.7 | -2.11 | 7.05 | A | A | A |
| 15.9 | 1.1 | 11.97 | 8.49 | A | A | A |
| 19.8 | 1.0 | 39.44 | 7.06 | N | N | N |
| 12.79 | 0.8 | -9.93 | 7.96 | A | A | A |
| 14.7 | 0.7 | 3.52 | 6.85 | A | A | A |
| 14.3 | 1.1 | 0.70 | 9.14 | A | A | A |
| 2.1 | 0.3 | -85.21 | 15.11 | N | N | N |
| 193 | 12 | 1259.15 | 7.93 | N | N | N |
| 16.89 | 1.03 | 18.94 | 7.84 | A | A | A |
| 14.1 | 0.2 | -0.70 | 5.13 | A | A | A |
| 13.6 | 0.8 | -4.23 | 7.67 | A | A | A |
| 16.8 | 2.4 | 18.31 | 15.11 | A | A | A |
| 16 | 3 | 12.68 | 19.39 | A | A | A |

TABLE 15. Evaluation Results for Sr-90 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg by beta

Labcode

141

145

149

150

151

152

157

166

167

169

175

180

184

185

186

188

189

192

193

195

196

198

199

201

202

203

205

206

215

218

232

237

245

247

248

255

262

263

264

268

270

277

279

284

287

289

290

291

293

295

296

298

300

302

304

317

327

328

330

334

TABLE 15. Evaluation Results for Sr-90 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg by beta

Labcode

Reported value

[Bq/kg ]

Reported uncertainty

(k = 1) [Bq/kg ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

341	20.37	6.80	43.45	33.74	N	N	N

342	20.35	1.04	43.31	7.10	N	N	N

Reported Results for Sb-125 determination in Sample 2

TABLE 16. Target values

34

Activity concentration  [ Bq/kg ]

FIGURE 16. Reported results and evaluation parameters for Sb-125. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

50	60	70	80	90	100

333

214

96

261

227

144

230

294

239

82

165

42

132

126

158

283

145

154

213

14

72

174

317

296

37

FIGURE 15. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

120

133

97

80

142

205

89

302

285

87

337

186

75

277

170

93

286

304

298

137

94

85

307

299

52

232

12

60

81

64

207

23

328

106

244

40

69

254

169

184

38

196

46

32

Parameter

Activity Concentration [ Bq/kg ] Uncertainty [ Bq/kg ]

Maximum Acceptable Relative Bias, [%]

269

153

47

315

217

19

274

44

22

71

223

20

122

183

301

5

Accepted 69%

26

67

112

130

10

Not Accepted 2%

258

66

21

107

105

233

263

235

222

182

62

163

295

Laboratory code

156

95

35

168

31

148

119

240

245

255

200

Warning 2%

201

248

114

234

300

79

275

24

271

303

125

281

155

151

260

90

36

329

228

171

309

30

173

256

264

29

57

241

134

63

292

115

Not Reported 27%

273

86

139

11

84

160

179

141

290

43

98

176

Value 73.0

3.6

20

210

180

25

279

190

306

334

251

48

28

59

6

185

55

108

77

54

146

140

268

34

118

99

152

270

342

51

188

136

246

284

209

224

123

110

257

162

161

252

272

104

314

218

27

293

202

1

287

327

247

318

2

147

192

178

15

197

129

198

311

341

68

226

316

308

135

249

331

199

330

159

58

243

225

| Reported value [Bq/kg ] 76.9 | Reported uncertainty (k = 1) [Bq/kg ] 2.6 | Relative bias [%] 5.34 | P-Test [%] 5.98 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 77.41 | 3.39 | 6.04 | 6.60 | A | A | A |
| 70.89 | 4.31 | -2.89 | 7.83 | A | A | A |
| 74.2 | 2.4 | 1.64 | 5.90 | A | A | A |
| 71.1 | 2.6 | -2.60 | 6.14 | A | A | A |
| 73.3 | 4.5 | 0.41 | 7.87 | A | A | A |
| 69 | 4 | -5.48 | 7.61 | A | A | A |
| 65.4 | 4.3 | -10.41 | 8.22 | A | A | A |
| 78.2 | 1.9 | 7.12 | 5.50 | A | A | A |
| 70.4 | 4.7 | -3.56 | 8.30 | A | A | A |
| 70.7 | 5.27 | -3.15 | 8.94 | A | A | A |
| 71.3 | 1.9 | -2.33 | 5.61 | A | A | A |
| 70.51 | 2.27 | -3.41 | 5.89 | A | A | A |
| 69.3 | 0.9 | -5.07 | 5.10 | A | A | A |
| 72.4 | 2.755 | -0.82 | 6.23 | A | A | A |
| 73.82 | 7.4 | 1.12 | 11.17 | A | A | A |
| 70.9 | 2.82 | -2.88 | 6.34 | A | A | A |
| 76.46 | 2.10 | 4.74 | 5.64 | A | A | A |
| 74.1 | 7.1 | 1.51 | 10.78 | A | A | A |
| 73 | 4 | 0.00 | 7.37 | A | A | A |
| 72.8 | 3.3 | -0.27 | 6.70 | A | A | A |
| 71.9 | 6.6 | -1.51 | 10.42 | A | A | A |
| 70.00 | 0.98 | -4.11 | 5.13 | A | A | A |
| 74.6 | 6.6 | 2.19 | 10.13 | A | A | A |
| 71.8 | 6.1 | -1.64 | 9.82 | A | A | A |
| 72.7 | 5.8 | -0.41 | 9.38 | A | A | A |
| 66.1 | 3.1 | -9.45 | 6.81 | A | A | A |
| 69.9 | 4.3 | -4.25 | 7.88 | A | A | A |
| 69.7 | 2.96 | -4.52 | 6.51 | A | A | A |
| 63.2 | 9.16 | -13.42 | 15.31 | A | A | A |
| 73.5 | 6.9 | 0.68 | 10.60 | A | A | A |
| 70.5 | 4.4 | -3.42 | 7.95 | A | A | A |
| 70 | 5 | -4.11 | 8.68 | A | A | A |
| 70.2 | 3.3 | -3.84 | 6.81 | A | A | A |
| 74.1 | 2.4 | 1.51 | 5.90 | A | A | A |
| 75.0 | 4.9 | 2.74 | 8.19 | A | A | A |
| 69.0 | 0.3 | -5.48 | 4.95 | A | A | A |
| 74.4 | 1.5 | 1.92 | 5.33 | A | A | A |
| 74.3 | 3.8 | 1.78 | 7.10 | A | A | A |
| 73.0 | 3.1 | 0.00 | 6.51 | A | A | A |
| 83.50 | 13.36 | 14.38 | 16.74 | A | A | A |
| 74.2 | 3.27 | 1.64 | 6.61 | A | A | A |
| 69.02 | 2.96 | -5.45 | 6.54 | A | A | A |
| 71.7 | 3.2 | -1.78 | 6.65 | A | A | A |
| 73.1 | 1.46 | 0.14 | 5.32 | A | A | A |
| 69.2 | 4.1 | -5.21 | 7.71 | A | A | A |
| 71.28 | 0.89 | -2.36 | 5.09 | A | A | A |
| 70.93 | 0.95 | -2.84 | 5.11 | A | A | A |
| 79.5 | 3.3 | 8.90 | 6.45 | A | A | A |
| 69.77 | 3.55 | -4.42 | 7.09 | A | A | A |
| 70.6 | 5 | -3.29 | 8.63 | A | A | A |
| 65.41 | 0.70 | -10.40 | 5.05 | A | A | A |
| 67.86 | 0.64 | -7.04 | 5.02 | A | A | A |
| 74.3 | 2.6 | 1.78 | 6.05 | A | A | A |
| 72.3 | 1.5 | -0.96 | 5.35 | A | A | A |
| 67.1 | 1.0 | -8.08 | 5.15 | A | A | A |
| 69.1 | 2.1 | -5.34 | 5.79 | A | A | A |
| 62.77 | 3.14 | -14.01 | 7.02 | A | A | A |
| 73.4 | 3.2 | 0.55 | 6.58 | A | A | A |
| 68.7 | 6.9 | -5.89 | 11.19 | A | A | A |

TABLE 17. Evaluation Results for Sb-125 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

1

2

5

6

10

11

12

14

15

19

20

21

22

23

24

25

26

27

28

29

30

31

32

34

35

36

37

38

40

42

43

44

46

47

48

51

52

54

55

57

58

59

60

62

63

64

66

67

68

69

71

72

75

77

79

80

81

82

84

85

| Reported value [Bq/kg ] 73.3 | Reported uncertainty (k = 1) [Bq/kg ] 2.5 | Relative bias [%] 0.41 | P-Test [%] 6.00 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 67.53 | 1.82 | -7.49 | 5.62 | A | A | A |
| 67.3 | 3.40 | -7.81 | 7.06 | A | A | A |
| 72.61 | 2.24 | -0.53 | 5.82 | A | A | A |
| 68.40 | 1.71 | -6.30 | 5.53 | A | A | A |
| 68.56 | 4.25 | -6.08 | 7.92 | A | A | A |
| 71.7 | 5.2 | -1.78 | 8.77 | A | A | A |
| 53.10 | 2.47 | -27.26 | 6.78 | N | N | N |
| 66.97 | 0.85 | -8.26 | 5.09 | A | A | A |
| 73.57 | 3.78 | 0.78 | 7.12 | A | A | A |
| 74.7 | 2.3 | 2.33 | 5.81 | A | A | A |
| 76.02 | 5.44 | 4.14 | 8.69 | A | A | A |
| 71.41 | 3.81 | -2.18 | 7.27 | A | A | A |
| 69.7 | 3.6 | -4.52 | 7.14 | A | A | A |
| 71.4 | 1.6 | -2.19 | 5.42 | A | A | A |
| 74.3 | 1.6 | 1.78 | 5.38 | A | A | A |
| 75.40 | 3.34 | 3.29 | 6.63 | A | A | A |
| 71 | 3 | -2.74 | 6.49 | A | A | A |
| 72.2 | 6.0 | -1.10 | 9.66 | A | A | A |
| 73.2 | 3.0 | 0.27 | 6.41 | A | A | A |
| 74.6 | 3.4 | 2.19 | 6.72 | A | A | A |
| 72 | 2.3 | -1.37 | 5.88 | A | A | A |
| 66.70 | 2.27 | -8.63 | 5.99 | A | A | A |
| 70.8 | 4.0 | -3.01 | 7.50 | A | A | A |
| 75.4 | 1.9 | 3.29 | 5.54 | A | A | A |
| 72.47 | 4.61 | -0.73 | 8.05 | A | A | A |
| 64 | 1 | -12.33 | 5.17 | A | A | A |
| 78.64 | 3.18 | 7.73 | 6.38 | A | A | A |
| 71.0 | 4.0 | -2.74 | 7.49 | A | A | A |
| 63.5 | 5.7 | -13.01 | 10.24 | A | A | A |
| 66.9 | 4.6 | -8.36 | 8.46 | A | A | A |
| 73 | 4 | 0.00 | 7.37 | A | A | A |
| 81 | 7 | 10.96 | 9.95 | A | A | A |
| 75 | 5 | 2.74 | 8.29 | A | A | A |
| 68.56 | 4.25 | -6.08 | 7.92 | A | A | A |
| 73.3 | 2.6 | 0.41 | 6.07 | A | A | A |
| 74.5 | 4.4 | 2.05 | 7.69 | A | A | A |
| 73.49 | 2.65 | 0.67 | 6.11 | A | A | A |
| 67.139 | 3.074 | -8.03 | 6.73 | A | A | A |
| 59.3 | 3.1 | -18.77 | 7.19 | A | N | W |
| 65.2 | 3.2 | -10.68 | 6.96 | A | A | A |
| 74.45 | 4.67 | 1.99 | 7.98 | A | A | A |
| 77.49 | 4.11 | 6.15 | 7.24 | A | A | A |
| 72.0 | 4.1 | -1.37 | 7.53 | A | A | A |
| 72.5 | 3.9 | -0.68 | 7.30 | A | A | A |
| 74.8 | 2.32 | 2.47 | 5.83 | A | A | A |
| 70.188 | 2.334 | -3.85 | 5.95 | A | A | A |
| 65.4 | 1.9 | -10.41 | 5.72 | A | A | A |
| 72.5 | 2.7 | -0.68 | 6.18 | A | A | A |
| 71.7 | 4.0 | -1.78 | 7.45 | A | A | A |
| 64.83 | 1.35 | -11.19 | 5.35 | A | A | A |
| 83.5 | 1.9 | 14.38 | 5.43 | A | N | W |
| 73.4 | 1.8 | 0.55 | 5.51 | A | A | A |
| 75.7 | 6.1 | 3.70 | 9.45 | A | A | A |
| 75.7 | 2.1 | 3.70 | 5.66 | A | A | A |
| 71.7 | 1.5 | -1.78 | 5.36 | A | A | A |
| 62.78 | 1.18 | -14.00 | 5.28 | A | N | W |
| 71.9 | 4.6 | -1.51 | 8.08 | A | A | A |
| 69.8 | 4.3 | -4.38 | 7.89 | A | A | A |
| 68.3 | 4.6 | -6.44 | 8.35 | A | A | A |

TABLE 17. Evaluation Results for Sb-125 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

86

87

89

90

93

94

95

96

97

98

99

104

105

106

107

108

110

112

114

115

118

119

120

122

123

125

126

129

130

132

133

134

135

136

137

139

140

141

142

144

145

146

147

148

151

152

153

154

155

156

158

159

160

161

162

163

165

168

169

170

| Reported value [Bq/kg ] 72.7 | Reported uncertainty (k = 1) [Bq/kg ] 1.66 | Relative bias [%] -0.41 | P-Test [%] 5.43 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 72.81 | 1.7 | -0.26 | 5.46 | A | A | A |
| 66 | 3 | -9.59 | 6.71 | A | A | A |
| 73.7 | 2.9 | 0.96 | 6.31 | A | A | A |
| 77.62 | 3.18 | 6.33 | 6.41 | A | A | A |
| 73.4 | 1.5 | 0.55 | 5.34 | A | A | A |
| 73.80 | 0.83 | 1.10 | 5.06 | A | A | A |
| 71.7 | 6.46 | -1.78 | 10.27 | A | A | A |
| 70.8 | 4.6 | -3.01 | 8.16 | A | A | A |
| 69.9 | 9.1 | -4.25 | 13.92 | A | A | A |
| 74.21 | 10.39 | 1.66 | 14.84 | A | A | A |
| 67.8 | 3.3 | -7.12 | 6.93 | A | A | A |
| 75 | 9 | 2.74 | 12.97 | A | A | A |
| 73.9 | 1.3 | 1.23 | 5.24 | A | A | A |
| 77.5 | 4.7 | 6.16 | 7.82 | A | A | A |
| 70.00 | 2.394 | -4.11 | 6.00 | A | A | A |
| 78.3555 | 3.5843 | 7.34 | 6.73 | A | A | A |
| 78.8 | 5.1 | 7.95 | 8.14 | A | A | A |
| 82.96 | 2.98 | 13.64 | 6.10 | A | A | A |
| 72.2 | 4.8 | -1.10 | 8.28 | A | A | A |
| 72.2 | 4.7 | -1.10 | 8.17 | A | A | A |
| 76.69 | 7.48 | 5.05 | 10.93 | A | A | A |
| 67.20 | 1.68 | -7.95 | 5.53 | A | A | A |
| 69.29 | 2.88 | -5.08 | 6.45 | A | A | A |
| 75.37 | 6.12 | 3.25 | 9.50 | A | A | A |
| 73.72 | 6.16 | 0.99 | 9.70 | A | A | A |
| 65.4 | 3.3 | -10.41 | 7.06 | A | A | A |
| 44.2 | 4.7 | -39.45 | 11.72 | N | N | N |
| 70.2 | 4.4 | -3.84 | 7.98 | A | A | A |
| 76.4 | 0.32 | 4.66 | 4.95 | A | A | A |
| 71.60 | 4.87 | -1.92 | 8.40 | A | A | A |
| 70.70 | 1.79 | -3.15 | 5.54 | A | A | A |
| 75.4 | 4.5 | 3.29 | 7.74 | A | A | A |
| 93.1 | 4.9 | 27.53 | 7.21 | N | N | N |
| 80.34 | 0.2 | 10.05 | 4.94 | A | A | A |
| 57.2 | 0.1 | -21.64 | 4.93 | N | N | N |
| 72.7 | 1.9 | -0.41 | 5.58 | A | A | A |
| 59.81 | 2.78 | -18.07 | 6.78 | A | N | W |
| 69 | 5 | -5.48 | 8.77 | A | A | A |
| 71.43 | 3.6 | -2.15 | 7.05 | A | A | A |
| 72.28 | 2.59 | -0.99 | 6.10 | A | A | A |
| 71.6 | 3.9 | -1.92 | 7.35 | A | A | A |
| 62.67 | 1.50 | -14.15 | 5.48 | A | N | W |
| 72.1 | 3.6 | -1.23 | 7.02 | A | A | A |
| 73.0 | 4.9 | 0.00 | 8.33 | A | A | A |
| 84.5 | 8.4 | 15.75 | 11.10 | A | A | A |
| 69.7 | 5.03 | -4.52 | 8.74 | A | A | A |
| 72.1 | 4.3 | -1.23 | 7.74 | A | A | A |
| 75.1 | 2.3 | 2.88 | 5.81 | A | A | A |
| 77.15 | 2.22 | 5.68 | 5.71 | A | A | A |
| 72.2 | 3.5 | -1.10 | 6.92 | A | A | A |
| 81.5 | 2.4 | 11.64 | 5.74 | A | A | A |
| 74.04 | 1.77 | 1.42 | 5.48 | A | A | A |
| 75.8 | 6.7 | 3.84 | 10.12 | A | A | A |
| 69.77 | 2.6 | -4.42 | 6.18 | A | A | A |
| 72.114 | 4.3880 | -1.21 | 7.83 | A | A | A |
| 72.9 | 3.11 | -0.14 | 6.52 | A | A | A |
| 75.6 | 2.6 | 3.56 | 6.01 | A | A | A |
| 71.11 | 2.63 | -2.59 | 6.16 | A | A | A |
| 72.58 | 5.45 | -0.58 | 8.98 | A | A | A |

TABLE 17. Evaluation Results for Sb-125 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

171

173

174

176

178

179

180

182

183

184

185

186

188

190

192

196

197

198

199

200

201

202

205

207

209

210

213

214

217

218

222

223

224

225

226

227

228

230

232

233

234

235

239

240

241

243

244

245

246

247

248

249

251

252

254

255

256

257

258

260

TABLE 17. Evaluation Results for Sb-125 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg by gamma

| Reported value [Bq/kg ] 56.5 | Reported uncertainty (k = 1) [Bq/kg ] 8.0 | Relative bias [%] -22.60 | P-Test [%] 14.99 | Trueness evaluation N | Precision evaluation A | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 71.5 | 1.8 | -2.05 | 5.54 | A | A | A |
| 72.96 | 6.817 | -0.05 | 10.57 | A | A | A |
| 74.5 | 4.1 | 2.05 | 7.39 | A | A | A |
| 70.1 | 2.2 | -3.97 | 5.85 | A | A | A |
| 74.8 | 4.1 | 2.47 | 7.37 | A | A | A |
| 72.4 | 1.5 | -0.82 | 5.35 | A | A | A |
| 76.0 | 4.9 | 4.11 | 8.12 | A | A | A |
| 73.2 | 2.2 | 0.27 | 5.78 | A | A | A |
| 70.4 | 3.78 | -3.56 | 7.29 | A | A | A |
| 72.34 | 4.5 | -0.90 | 7.94 | A | A | A |
| 68.10 | 0.70 | -6.71 | 5.04 | A | A | A |
| 73.84 | 2.31 | 1.15 | 5.84 | A | A | A |
| 72.48 | 7.248 | -0.71 | 11.15 | A | A | A |
| 65.14 | 10.86 | -10.77 | 17.39 | A | A | A |
| 75.2 | 4.6 | 3.01 | 7.86 | A | A | A |
| 67.491 | 1.4636 | -7.55 | 5.39 | A | A | A |
| 68.4 | 3 | -6.30 | 6.60 | A | A | A |
| 76.9 | 6.21 | 5.34 | 9.46 | A | A | A |
| 73.5 | 0.68 | 0.68 | 5.02 | A | A | A |
| 73.15 | 4.36 | 0.21 | 7.74 | A | A | A |
| 76.5 | 3.3 | 4.79 | 6.55 | A | A | A |
| 60.3 | 3.0 | -17.40 | 7.01 | A | A | A |
| 71.7 | 2.2 | -1.78 | 5.81 | A | A | A |
| 66.04 | 2.18 | -9.53 | 5.93 | A | A | A |
| 68.5 | 3.7 | -6.16 | 7.31 | A | A | A |
| 68.9 | 1.2 | -5.62 | 5.23 | A | A | A |
| 72.3 | 4.3 | -0.96 | 7.73 | A | A | A |
| 70.89 | 3.05 | -2.89 | 6.54 | A | A | A |
| 67.4 | 2.1 | -7.67 | 5.83 | A | A | A |
| 72.44 | 5.91 | -0.77 | 9.53 | A | A | A |
| 68.5 | 2.3 | -6.16 | 5.97 | A | A | A |
| 73.97 | 8.21 | 1.33 | 12.15 | A | A | A |
| 68.73 | 0.71 | -5.85 | 5.04 | A | A | A |
| 80.92 | 1.00 | 10.85 | 5.08 | A | A | A |
| 72.79 | 5.78 | -0.29 | 9.35 | A | A | A |
| 79.1 | 2.8 | 8.36 | 6.07 | A | A | A |
| 76.1 | 8.9 | 4.25 | 12.69 | A | A | A |
| 70.2 | 3.9 | -3.84 | 7.43 | A | A | A |
| 80.4 | 2.7 | 10.14 | 5.97 | A | A | A |
| 66.02 | 2.12 | -9.56 | 5.88 | A | A | A |
| 77.3 | 2.5 | 5.89 | 5.90 | A | A | A |
| 77.1 | 1.5 | 5.62 | 5.30 | A | A | A |
| 69.3 | 3.8 | -5.07 | 7.37 | A | A | A |
| 72.7 | 5.4 | -0.41 | 8.92 | A | A | A |
| 83.0 | 8.3 | 13.70 | 11.15 | A | A | A |
| 82.92 | 2.58 | 13.59 | 5.83 | A | A | A |
| 30.1 | 0.7 | -58.77 | 5.45 | N | N | N |
| 74 | 3 | 1.37 | 6.38 | A | A | A |
| 67.8 | 2.4 | -7.12 | 6.07 | A | A | A |
| 79.38 | 2.17 | 8.74 | 5.64 | A | A | A |
| 74.89 | 2.99 | 2.59 | 6.35 | A | A | A |

Labcode

261

263

264

268

269

270

271

272

273

274

275

277

279

281

283

284

285

286

287

290

292

293

294

295

296

298

299

300

301

302

303

304

306

307

308

309

311

314

315

316

317

318

327

328

329

330

331

333

334

337

341

342

Activity concentration  [ Bq/kg ]

40

45

Reported Results for Cs-134 determination in Sample 2

TABLE 18. Target values

| Parameter | Value |
| --- | --- |
| Activity Concentration [ Bq/kg ] | 40.0 |
| Uncertainty [ Bq/kg ] | 2.0 |
| Maximum Acceptable Relative Bias, [%] | 20 |

Accepted 79%

Not Reported 10%

Warning 3%

Not Accepted 8%

50

55

FIGURE 17. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

25

30

35

101

282

154

214

144

343

13

308

116

243

330

285

277

221

80

47

93

305

197

206

294

295

115

236

106

20

304

259

286

77

120

344

342

189

319

167

108

279

74

309

162

165

306

283

89

171

119

246

147

130

199

145

64

274

23

131

91

102

137

159

148

90

51

223

174

196

6

264

247

271

169

40

224

202

250

151

19

293

107

3

155

244

248

180

276

57

232

26

183

207

272

156

328

329

114

205

290

228

291

36

10

122

32

44

35

67

43

173

188

71

234

209

2

30

118

18

262

251

48

201

298

141

311

334

72

270

245

192

58

275

318

166

233

296

28

249

190

16

226

198

225

338

287

322

Laboratory code

FIGURE 18. Reported results and evaluation parameters for Cs-134. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

| Reported value [Bq/kg ] 43.3 | Reported uncertainty (k = 1) [Bq/kg ] 1.4 | Relative bias [%] 8.25 | P-Test [%] 5.95 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 40.18 | 1.55 | 0.45 | 6.32 | A | A | A |
| 38.8 | 0.95 | -3.00 | 5.57 | A | A | A |
| 33.42 | 2.28 | -16.45 | 8.46 | A | A | A |
| 39.33 | 2.33 | -1.68 | 7.75 | A | A | A |
| 38.1 | 1.4 | -4.75 | 6.21 | A | A | A |
| 55.16 | 2.46 | 37.90 | 6.70 | N | N | N |
| 37.74 | 1.55 | -5.65 | 6.47 | A | A | A |
| 39.6 | 2.1 | -1.00 | 7.29 | A | A | A |
| 40.2 | 2.0 | 0.50 | 7.05 | A | A | A |
| 36 | 3 | -10.00 | 9.72 | A | A | A |
| 29.2 | 0.4 | -27.00 | 5.18 | N | N | N |
| 38.8 | 2.5 | -3.00 | 8.16 | A | A | A |
| 43.6 | 0.8 | 9.00 | 5.33 | A | A | A |
| 43.12 | 1.63 | 7.80 | 6.27 | A | A | A |
| 38.0 | 2.1 | -5.00 | 7.45 | A | A | A |
| 40.34 | 2.76 | 0.85 | 8.47 | A | A | A |
| 38.7 | 2.6 | -3.25 | 8.37 | A | A | A |
| 35.1 | 2.13 | -12.25 | 7.86 | A | A | A |
| 35.1 | 0.94 | -12.25 | 5.67 | A | A | A |
| 40.58 | 1.35 | 1.45 | 6.01 | A | A | A |
| 37.4 | 0.6 | -6.50 | 5.25 | A | A | A |
| 40.4 | 1.401 | 1.00 | 6.08 | A | A | A |
| 38.93 | 3.9 | -2.68 | 11.20 | A | A | A |
| 39.1 | 1.50 | -2.25 | 6.30 | A | A | A |
| 39.77 | 0.87 | -0.57 | 5.46 | A | A | A |
| 41.9 | 3.9 | 4.75 | 10.57 | A | A | A |
| 39 | 2 | -2.50 | 7.16 | A | A | A |
| 40.2 | 1.8 | 0.50 | 6.71 | A | A | A |
| 37.7 | 2.1 | -5.75 | 7.49 | A | A | A |
| 39.66 | 0.50 | -0.85 | 5.16 | A | A | A |
| 41.2 | 2.8 | 3.00 | 8.44 | A | A | A |
| 40.7 | 4 | 1.75 | 11.03 | A | A | A |
| 39.7 | 3.4 | -0.75 | 9.92 | A | A | A |
| 39.6 | 3.3 | -1.00 | 9.72 | A | A | A |
| 38.5 | 2.1 | -3.75 | 7.40 | A | A | A |
| 39.3 | 2.3 | -1.75 | 7.70 | A | A | A |
| 36.1 | 1.3 | -9.75 | 6.16 | A | A | A |
| 38.4 | 1.16 | -4.00 | 5.84 | A | A | A |
| 36.2 | 3.44 | -9.50 | 10.74 | A | A | A |
| 39.8 | 3.4 | -0.50 | 9.90 | A | A | A |
| 39.7 | 1.5 | -0.75 | 6.27 | A | A | A |
| 34.0 | 2.8 | -15.00 | 9.63 | A | A | A |
| 33 | 2 | -17.50 | 7.86 | A | A | A |
| 33.7 | 1.4 | -15.75 | 6.50 | A | A | A |
| 40.5 | 1.1 | 1.25 | 5.69 | A | A | A |
| 38.69 | 0.38 | -3.28 | 5.10 | A | A | A |
| 38.71 | 0.39 | -3.22 | 5.10 | A | A | A |
| 37.8 | 2.2 | -5.50 | 7.67 | A | A | A |
| 38.1 | 1.6 | -4.75 | 6.53 | A | A | A |
| 39.67 | 0.44 | -0.82 | 5.12 | A | A | A |
| 41.4 | 2.1 | 3.50 | 7.12 | A | A | A |
| 46.1 | 1.6 | 15.25 | 6.09 | A | A | A |
| 39.0 | 1.7 | -2.50 | 6.63 | A | A | A |
| 41.30 | 6.61 | 3.25 | 16.77 | A | A | A |
| 38.7 | 1.74 | -3.25 | 6.72 | A | A | A |
| 36.76 | 1.52 | -8.10 | 6.49 | A | A | A |
| 36.4 | 1.6 | -9.00 | 6.66 | A | A | A |
| 37.1 | 0.56 | -7.25 | 5.22 | A | A | A |
| 37.2 | 1.81 | -7.00 | 6.98 | A | A | A |

TABLE 19. Evaluation Results for Cs-134 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

1

2

3

4

5

6

7

8

10

11

12

13

14

15

16

17

18

19

20

21

22

23

24

25

26

27

28

29

30

31

32

33

34

35

36

37

38

39

40

42

43

44

45

46

47

48

49

50

51

52

54

55

56

57

58

59

60

62

63

64

| Reported value [Bq/kg ] 41.75 | Reported uncertainty (k = 1) [Bq/kg ] 0.46 | Relative bias [%] 4.38 | P-Test [%] 5.12 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 39.76 | 0.34 | -0.60 | 5.07 | A | A | A |
| 42.3 | 1.1 | 5.75 | 5.64 | A | A | A |
| 34.94 | 0.75 | -12.65 | 5.44 | A | A | A |
| 7.43 | 0.39 | -81.42 | 7.25 | N | N | N |
| 40 | 3 | 0.00 | 9.01 | A | A | A |
| 41.03 | 0.88 | 2.58 | 5.44 | A | A | A |
| 39.3 | 1.2 | -1.75 | 5.86 | A | A | A |
| 36.29 | 0.74 | -9.28 | 5.40 | A | A | A |
| 36.89 | 0.27 | -7.77 | 5.05 | A | A | A |
| 35.6 | 1.2 | -11.00 | 6.03 | A | A | A |
| 39.8 | 1 | -0.50 | 5.60 | A | A | A |
| 33.6 | 0.5 | -16.00 | 5.22 | A | N | W |
| 38.9 | 1.2 | -2.75 | 5.88 | A | A | A |
| 36.2 | 1.69 | -9.50 | 6.84 | A | A | A |
| 34.72 | 0.6 | -13.20 | 5.29 | A | A | A |
| 35.9 | 1.54 | -10.25 | 6.59 | A | A | A |
| 40.2 | 2.3 | 0.50 | 7.60 | A | A | A |
| 41.2 | 1.5 | 3.00 | 6.19 | A | A | A |
| 35.76 | 1.02 | -10.60 | 5.76 | A | A | A |
| 36.70 | 1.85 | -8.25 | 7.10 | A | A | A |
| 37.78 | 1.04 | -5.55 | 5.71 | A | A | A |
| 37.48 | 0.45 | -6.30 | 5.14 | A | A | A |
| 33.80 | 0.74 | -15.50 | 5.46 | A | N | W |
| 37.66 | 1.92 | -5.85 | 7.14 | A | A | A |
| 39.6 | 2.4 | -1.00 | 7.86 | A | A | A |
| 26.55 | 1.14 | -33.62 | 6.59 | N | N | N |
| 34.66 | 0.47 | -13.35 | 5.18 | A | A | A |
| 39.20 | 2.00 | -2.00 | 7.14 | A | A | A |
| 35.91 | 0.78 | -10.23 | 5.45 | A | A | A |
| 5.8 |  | -85.50 |  | N |  | N |
| 37.5 | 1.2 | -6.25 | 5.94 | A | A | A |
| 35.13 | 1.75 | -12.17 | 7.06 | A | A | A |
| 39.36 | 2.11 | -1.60 | 7.33 | A | A | A |
| 35.0 | 1.8 | -12.50 | 7.17 | A | A | A |
| 38.7 | 0.8 | -3.25 | 5.41 | A | A | A |
| 36.2 | 3.9 | -9.50 | 11.88 | A | A | A |
| 39.99 | 1.32 | -0.02 | 5.99 | A | A | A |
| 42 | 2 | 5.00 | 6.90 | A | A | A |
| 39.3 | 3.5 | -1.75 | 10.21 | A | A | A |
| 34.7 | 1.2 | -13.25 | 6.08 | A | A | A |
| 30.54 | 1.08 | -23.65 | 6.12 | N | N | N |
| 40.3 | 1.6 | 0.75 | 6.38 | A | A | A |
| 36.86 | 0.8 | -7.85 | 5.45 | A | A | A |
| 35.77 | 1.21 | -10.57 | 6.04 | A | A | A |
| 39.6 | 1.6 | -1.00 | 6.43 | A | A | A |
| 40.8 | 1.1 | 2.00 | 5.68 | A | A | A |
| 36.42 | 0.53 | -8.95 | 5.21 | A | A | A |
| 39.07 | 2.63 | -2.32 | 8.39 | A | A | A |
| 32.2 | 0.4 | -19.50 | 5.15 | A | N | W |
| 37.59 | 0.09 | -6.02 | 5.01 | A | A | A |
| 37.0 | 3.0 | -7.50 | 9.53 | A | A | A |
| 37.44 | 2.54 | -6.40 | 8.43 | A | A | A |
| 29.0 | 2.6 | -27.50 | 10.27 | N | N | N |
| 39.1 | 3.7 | -2.25 | 10.70 | A | A | A |
| 40 | 2 | 0.00 | 7.07 | A | A | A |
| 40 | 3 | 0.00 | 9.01 | A | A | A |
| 37 | 3 | -7.50 | 9.53 | A | A | A |
| 37.66 | 2.24 | -5.85 | 7.77 | A | A | A |
| 40.8 | 1.5 | 2.00 | 6.21 | A | A | A |

TABLE 19. Evaluation Results for Cs-134 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

66

67

68

69

70

71

72

73

74

75

77

79

80

81

82

83

84

85

86

87

89

90

91

93

94

95

96

97

98

99

101

102

104

105

106

107

108

110

112

114

115

116

118

119

120

122

123

124

125

126

129

130

131

132

133

134

135

136

137

139

| Reported value [Bq/kg ] 38.2 | Reported uncertainty (k = 1) [Bq/kg ] 2.2 | Relative bias [%] -4.50 | P-Test [%] 7.63 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 40.78 | 1.92 | 1.95 | 6.87 | A | A | A |
| 30.428 | 2.068 | -23.93 | 8.44 | N | N | N |
| 27.0 | 1.8 | -32.50 | 8.33 | N | N | N |
| 37.1 | 2.8 | -7.25 | 9.05 | A | A | A |
| 40.38 | 2.23 | 0.95 | 7.45 | A | A | A |
| 36.91 | 1.82 | -7.73 | 7.02 | A | A | A |
| 37.7 | 2.0 | -5.75 | 7.29 | A | A | A |
| 31.9 | 1.46 | -20.25 | 6.78 | N | N | N |
| 41.0 | 1.3 | 2.50 | 5.92 | A | A | A |
| 38.6 | 2.1 | -3.50 | 7.39 | A | A | A |
| 35.6 | 1.01 | -11.00 | 5.75 | A | A | A |
| 39.159 | 1.757 | -2.10 | 6.72 | A | A | A |
| 19.13 | 0.57 | -52.18 | 5.82 | N | N | N |
| 38.9 | 1.4 | -2.75 | 6.16 | A | A | A |
| 39.2 | 3 | -2.00 | 9.14 | A | A | A |
| 39.6 | 0.5 | -1.00 | 5.16 | A | A | A |
| 41.43 | 0.87 | 3.57 | 5.42 | A | A | A |
| 37.67 | 0.68 | -5.82 | 5.32 | A | A | A |
| 39.6 | 2.5 | -1.00 | 8.05 | A | A | A |
| 41.3 | 3.3 | 3.25 | 9.43 | A | A | A |
| 36.4 | 1.3 | -9.00 | 6.14 | A | A | A |
| 39.2 | 1.7 | -2.00 | 6.62 | A | A | A |
| 36.41 | 0.33 | -8.98 | 5.08 | A | A | A |
| 41.5 | 4.3 | 3.75 | 11.50 | A | A | A |
| 36.20 | 1.38 | -9.50 | 6.29 | A | A | A |
| 37.8 | 2.2 | -5.50 | 7.67 | A | A | A |
| 38.3 | 2.3 | -4.25 | 7.81 | A | A | A |
| 37.3 | 2.3 | -6.75 | 7.94 | A | A | A |
| 36.7 | 0.9 | -8.25 | 5.57 | A | A | A |
| 35.4 | 2.75 | -11.50 | 9.24 | A | A | A |
| 39.87 | 1.23 | -0.33 | 5.88 | A | A | A |
| 38 | 2 | -5.00 | 7.26 | A | A | A |
| 38.5 | 1.7 | -3.75 | 6.67 | A | A | A |
| 49.3 | 0.14 | 23.25 | 5.01 | N | N | N |
| 38.16 | 0.41 | -4.60 | 5.11 | A | A | A |
| 39.0 | 0.71 | -2.50 | 5.32 | A | A | A |
| 38.95 | 0.41 | -2.62 | 5.11 | A | A | A |
| 39.7 | 2.0 | -0.75 | 7.10 | A | A | A |
| 39.56 | 3.17 | -1.10 | 9.45 | A | A | A |
| 39.1 | 2.3 | -2.25 | 7.72 | A | A | A |
| 47.3 | 5.2 | 18.25 | 12.08 | A | A | A |
| 41.50 | 4.36 | 3.75 | 11.64 | A | A | A |
| 36.4 | 1.8 | -9.00 | 7.03 | A | A | A |
| 40 | 5 | 0.00 | 13.46 | A | A | A |
| 36 | 3 | -10.00 | 9.72 | A | A | A |
| 42.3 | 0.7 | 5.75 | 5.27 | A | A | A |
| 41.2 | 2.5 | 3.00 | 7.86 | A | A | A |
| 37.02 | 0.81 | -7.45 | 5.46 | A | A | A |
| 40.3 | 2.7 | 0.75 | 8.36 | A | A | A |
| 38.03 | 1.270 | -4.92 | 6.01 | A | A | A |
| 34.0711 | 1.8350 | -14.82 | 7.35 | A | A | A |
| 45.5 | 1.6 | 13.75 | 6.11 | A | A | A |
| 37.08 | 1.84 | -7.30 | 7.04 | A | A | A |
| 39.72 | 2.35 | -0.70 | 7.75 | A | A | A |
| 40.6 | 2.3 | 1.50 | 7.56 | A | A | A |
| 38.55 | 2.32 | -3.63 | 7.82 | A | A | A |
| 38.2 | 2.7 | -4.50 | 8.66 | A | A | A |
| 39.36 | 0.87 | -1.60 | 5.47 | A | A | A |
| 34.266 | 0.866 | -14.34 | 5.60 | A | A | A |

TABLE 19. Evaluation Results for Cs-134 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

140

141

142

144

145

146

147

148

149

150

151

152

153

154

155

156

157

158

159

160

161

162

163

165

166

167

168

169

170

171

172

173

174

176

177

178

179

180

181

182

183

184

185

186

188

189

190

192

193

195

196

197

198

199

200

201

202

204

205

206

| Reported value [Bq/kg ] 39.18 | Reported uncertainty (k = 1) [Bq/kg ] 1.46 | Relative bias [%] -2.05 | P-Test [%] 6.24 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 40.07 | 3.26 | 0.18 | 9.55 | A | A | A |
| 33.37 | 2.79 | -16.58 | 9.74 | A | A | A |
| 37.443 | 0.565 | -6.39 | 5.22 | A | A | A |
| 37.1 | 2.1 | -7.25 | 7.55 | A | A | A |
| 26.1 | 2.5 | -34.75 | 10.81 | N | N | N |
| 36.7 | 2.3 | -8.25 | 8.02 | A | A | A |
| 38.3 | 0.21 | -4.25 | 5.03 | A | A | A |
| 33.40 | 1.00 | -16.50 | 5.83 | A | N | W |
| 39.20 | 1.67 | -2.00 | 6.57 | A | A | A |
| 37.90 | 0.93 | -5.25 | 5.57 | A | A | A |
| 38.5 | 0.2 | -3.75 | 5.03 | A | A | A |
| 46.1 | 1.8 | 15.25 | 6.34 | A | A | A |
| 43.41 | 0.2 | 8.52 | 5.02 | A | A | A |
| 29.3 | 0.2 | -26.75 | 5.05 | N | N | N |
| 39.4 | 1.0 | -1.50 | 5.61 | A | A | A |
| 30.55 | 1.65 | -23.62 | 7.36 | N | N | N |
| 39 | 4 | -2.50 | 11.41 | A | A | A |
| 41.5 | 1.92 | 3.75 | 6.81 | A | A | A |
| 40.04 | 1.09 | 0.10 | 5.69 | A | A | A |
| 40.5 | 2.2 | 1.25 | 7.38 | A | A | A |
| 34.86 | 0.66 | -12.85 | 5.35 | A | A | A |
| 41.50 | 0.60 | 3.75 | 5.20 | A | A | A |
| 17.12 | 0.41 | -57.20 | 5.54 | N | N | N |
| 39.4 | 2.0 | -1.50 | 7.13 | A | A | A |
| 36.0 | 2.2 | -10.00 | 7.90 | A | A | A |
| 31 | 3.1 | -22.50 | 11.18 | N | A | N |
| 38.9 | 2.78 | -2.75 | 8.72 | A | A | A |
| 41.2 | 3.7 | 3.00 | 10.28 | A | A | A |
| 36.9 | 1.4 | -7.75 | 6.28 | A | A | A |
| 38.2 | 1.2 | -4.50 | 5.90 | A | A | A |
| 38.9 | 2.0 | -2.75 | 7.17 | A | A | A |
| 42 | 1.1 | 5.00 | 5.64 | A | A | A |
| 38.6 | 7.8 | -3.50 | 20.82 | A | N | W |
| 40.40 | 0.56 | 1.00 | 5.19 | A | A | A |
| 36.3 | 2.9 | -9.25 | 9.42 | A | A | A |
| 33.6 | 0.8 | -16.00 | 5.54 | A | N | W |
| 38.099 | 2.3382 | -4.75 | 7.92 | A | A | A |
| 38.7 | 1.78 | -3.25 | 6.79 | A | A | A |
| 38.6 | 1.3 | -3.50 | 6.03 | A | A | A |
| 38.96 | 1.46 | -2.60 | 6.25 | A | A | A |
| 35.4 | 3 | -11.50 | 9.84 | A | A | A |
| 37.85 | 2.84 | -5.37 | 9.02 | A | A | A |
| 42.6 | 1.8 | 6.50 | 6.55 | A | A | A |
| 40.4 | 2.25 | 1.00 | 7.48 | A | A | A |
| 38.0 | 0.8 | -5.00 | 5.43 | A | A | A |
| 38.16 | 3.479 | -4.60 | 10.40 | A | A | A |
| 40.1 | 0.84 | 0.25 | 5.42 | A | A | A |
| 37.4 | 2.1 | -6.50 | 7.52 | A | A | A |
| 40.05 | 0.44 | 0.12 | 5.12 | A | A | A |
| 41.1 | 2.2 | 2.75 | 7.32 | A | A | A |
| 38.2 | 1.1 | -4.50 | 5.77 | A | A | A |
| 39.2 | 2.6 | -2.00 | 8.31 | A | A | A |
| 35.4 | 1.7 | -11.50 | 6.93 | A | A | A |
| 37.3 | 1.99 | -6.75 | 7.31 | A | A | A |
| 41.4 | 2.5 | 3.50 | 7.84 | A | A | A |
| 39.0 | 2.7 | -2.50 | 8.54 | A | A | A |
| 33.20 | 0.29 | -17.00 | 5.08 | A | N | W |
| 36.24 | 1.41 | -9.40 | 6.34 | A | A | A |
| 38.58 | 3.558 | -3.55 | 10.49 | A | A | A |

TABLE 19. Evaluation Results for Cs-134 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

207

209

210

211

213

214

217

218

221

222

223

224

225

226

227

228

230

232

233

234

235

236

237

239

240

241

243

244

245

246

247

248

249

250

251

252

254

255

256

257

258

259

260

261

262

263

264

265

268

269

270

271

272

273

274

275

276

277

279

281

TABLE 19. Evaluation Results for Cs-134 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg by gamma

| Reported value [Bq/kg ] 17 | Reported uncertainty (k = 1) [Bq/kg ] 0.31 | Relative bias [%] -57.50 | P-Test [%] 5.32 | Trueness evaluation N | Precision evaluation N | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 36.67 | 5.984 | -8.32 | 17.07 | A | A | A |
| 34.5 | 2.0 | -13.75 | 7.66 | A | A | A |
| 32.463 | 0.69741 | -18.84 | 5.44 | A | N | W |
| 35.4 | 1.74 | -11.50 | 7.01 | A | A | A |
| 49.7 | 2.55 | 24.25 | 7.16 | N | N | N |
| 39.4 | 0.28 | -1.50 | 5.05 | A | A | A |
| 39.5 | 2.79 | -1.25 | 8.65 | A | A | A |
| 33.91 | 1.99 | -15.23 | 7.71 | A | A | A |
| 38.7 | 1.9 | -3.25 | 7.01 | A | A | A |
| 34.5 | 1.8 | -13.75 | 7.23 | A | A | A |
| 34.6 | 0.9 | -13.50 | 5.64 | A | A | A |
| 41.65 | 1.48 | 4.12 | 6.13 | A | A | A |
| 40.7 | 2.4 | 1.75 | 7.73 | A | A | A |
| 27.6 | 0.4 | -31.00 | 5.21 | N | N | N |
| 37.5 | 2.2 | -6.25 | 7.71 | A | A | A |
| 38.90 | 1.32 | -2.75 | 6.04 | A | A | A |
| 34.3 | 1.5 | -14.25 | 6.64 | A | A | A |
| 39.46 | 3.39 | -1.35 | 9.94 | A | A | A |
| 35.3 | 2.3 | -11.75 | 8.21 | A | A | A |
| 34 | 1 | -15.00 | 5.80 | A | N | W |
| 36.59 | 2.61 | -8.52 | 8.71 | A | A | A |
| 36.91 | 0.33 | -7.73 | 5.08 | A | A | A |
| 29.73 | 0.39 | -25.67 | 5.17 | N | N | N |
| 36.36 | 1.67 | -9.10 | 6.79 | A | A | A |
| 40.8 | 4.2 | 2.00 | 11.44 | A | A | A |
| 36.6 | 4.9 | -8.50 | 14.29 | A | A | A |
| 34.2 | 2.6 | -14.50 | 9.10 | A | A | A |
| 37.3 | 2.0 | -6.75 | 7.33 | A | A | A |
| 36.27 | 0.35 | -9.32 | 5.09 | A | A | A |
| 41.4 | 2.5 | 3.50 | 7.84 | A | A | A |
| 36.047 | 9.011 | -9.88 | 25.49 | A | N | W |
| 87.7 | 6.75 | 119.25 | 9.18 | N | N | N |
| 41.1 | 0.8 | 2.75 | 5.37 | A | A | A |
| 39.2 | 2.1 | -2.00 | 7.33 | A | A | A |
| 39.3 | 2.6 | -1.75 | 8.29 | A | A | A |
| 32.0 | 3.2 | -20.00 | 11.18 | A | A | A |
| 41.10 | 0.72 | 2.75 | 5.30 | A | A | A |
| 33.8 | 1.6 | -15.50 | 6.89 | A | A | A |
| 40.9 | 1.8 | 2.25 | 6.66 | A | A | A |
| 25.19 | 2.19 | -37.02 | 10.03 | N | N | N |
| 36.7 | 1.3 | -8.25 | 6.13 | A | A | A |
| 49.30 | 0.14 | 23.25 | 5.01 | N | N | N |
| 40.73 | 1.10 | 1.82 | 5.68 | A | A | A |
| 35.95 | 1.26 | -10.12 | 6.11 | A | A | A |
| 27.8397 |  | -30.40 |  | N |  | N |
| 35.9 | 2.3 | -10.25 | 8.13 | A | A | A |

Labcode

282

283

284

285

286

287

290

291

292

293

294

295

296

298

299

300

301

302

303

304

305

306

307

308

309

311

314

315

316

317

318

319

322

327

328

329

330

331

333

334

335

337

338

341

342

343

344

Activity concentration  [ Bq/kg ]

45

50

Reported Results for Cs-137 determination in Sample 2

TABLE 20. Target values

| Parameter | Value |
| --- | --- |
| Activity Concentration [ Bq/kg ] | 44.1 |
| Uncertainty [ Bq/kg ] | 2.2 |
| Maximum Acceptable Relative Bias, [%] | 20 |

Accepted 83%

Not Reported 10%

Warning 2%

Not Accepted 5%

55

60

FIGURE 19. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

30

35

40

319

336

221

239

154

294

46

120

172

133

47

170

305

137

93

39

290

72

38

75

96

169

19

6

183

230

5

235

50

3

163

12

87

160

217

102

34

213

236

52

206

304

179

168

110

18

106

228

91

300

182

146

14

159

259

180

211

158

204

202

260

188

295

130

35

155

36

244

48

255

276

222

114

237

141

245

173

125

99

44

119

20

79

37

59

268

95

45

329

248

51

201

195

171

261

29

140

71

134

251

157

318

200

279

152

74

2

233

256

285

303

136

55

54

24

327

28

7

298

264

246

284

184

142

68

162

306

98

309

13

307

314

189

116

16

293

147

214

199

341

56

340

226

302

Laboratory code

FIGURE 20. Reported results and evaluation parameters for Cs-137. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

| Reported value [Bq/kg ] 47.2 | Reported uncertainty (k = 1) [Bq/kg ] 1.6 | Relative bias [%] 7.03 | P-Test [%] 6.03 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 45.61 | 1.76 | 3.42 | 6.31 | A | A | A |
| 42.9 | 2.60 | -2.72 | 7.85 | A | A | A |
| 48.98 | 3.14 | 11.07 | 8.12 | A | A | A |
| 42.76 | 2.57 | -3.04 | 7.81 | A | A | A |
| 42.5 | 2.5 | -3.63 | 7.71 | A | A | A |
| 46.33 | 1.78 | 5.06 | 6.30 | A | A | A |
| 47.29 | 2.88 | 7.23 | 7.87 | A | A | A |
| 44.3 | 2.3 | 0.45 | 7.20 | A | A | A |
| 44.8 | 2.5 | 1.59 | 7.49 | A | A | A |
| 43 | 2 | -2.49 | 6.82 | A | A | A |
| 47.9 | 1.4 | 8.62 | 5.78 | A | A | A |
| 43.7 | 2.8 | -0.91 | 8.12 | A | A | A |
| 47.7 | 0.9 | 8.16 | 5.33 | A | A | A |
| 48.86 | 2.29 | 10.79 | 6.84 | A | A | A |
| 42.1 | 2.3 | -4.54 | 7.40 | A | A | A |
| 43.47 | 2.84 | -1.43 | 8.22 | A | A | A |
| 42.5 | 2.8 | -3.63 | 8.26 | A | A | A |
| 44.4 | 3.98 | 0.68 | 10.26 | A | A | A |
| 44.5 | 1.4 | 0.91 | 5.90 | A | A | A |
| 44.40 | 1.64 | 0.68 | 6.21 | A | A | A |
| 43.4 | 1.1 | -1.59 | 5.60 | A | A | A |
| 46.3 | 1.532 | 4.99 | 5.99 | A | A | A |
| 42.38 | 4.3 | -3.90 | 11.31 | A | A | A |
| 42.0 | 1.61 | -4.76 | 6.29 | A | A | A |
| 44.11 | 0.82 | 0.02 | 5.32 | A | A | A |
| 46.3 | 4.5 | 4.99 | 10.92 | A | A | A |
| 45 | 2 | 2.04 | 6.68 | A | A | A |
| 44.1 | 2.0 | 0.00 | 6.74 | A | A | A |
| 43.3 | 2.6 | -1.81 | 7.81 | A | A | A |
| 45.07 | 0.54 | 2.20 | 5.13 | A | A | A |
| 46.4 | 3.1 | 5.22 | 8.34 | A | A | A |
| 43.2 | 3.7 | -2.04 | 9.91 | A | A | A |
| 44.0 | 3.5 | -0.23 | 9.39 | A | A | A |
| 44.1 | 3.5 | 0.00 | 9.37 | A | A | A |
| 44.6 | 3.3 | 1.13 | 8.92 | A | A | A |
| 42.1 | 3.3 | -4.54 | 9.29 | A | A | A |
| 41.7 | 1.3 | -5.44 | 5.88 | A | A | A |
| 42.3 | 1.10 | -4.08 | 5.63 | A | A | A |
| 39.4 | 5.51 | -10.66 | 14.85 | A | A | A |
| 43.9 | 3.4 | -0.45 | 9.21 | A | A | A |
| 44.4 | 1.0 | 0.68 | 5.47 | A | A | A |
| 44.6 | 2.3 | 1.13 | 7.18 | A | A | A |
| 39 | 2 | -11.56 | 7.15 | A | A | A |
| 40.5 | 1.3 | -8.16 | 5.93 | A | A | A |
| 44.1 | 1.1 | 0.00 | 5.58 | A | A | A |
| 42.79 | 0.61 | -2.97 | 5.19 | A | A | A |
| 42.87 | 0.61 | -2.79 | 5.19 | A | A | A |
| 44.8 | 2.9 | 1.59 | 8.17 | A | A | A |
| 43.2 | 1.2 | -2.04 | 5.71 | A | A | A |
| 46.13 | 0.76 | 4.60 | 5.25 | A | A | A |
| 46.1 | 2.3 | 4.54 | 7.06 | A | A | A |
| 59.2 | 1.9 | 34.24 | 5.93 | N | N | N |
| 45.0 | 3.0 | 2.04 | 8.33 | A | A | A |
| 45.60 | 7.29 | 3.40 | 16.75 | A | A | A |
| 44.6 | 1.56 | 1.13 | 6.09 | A | A | A |
| 43.50 | 1.96 | -1.36 | 6.72 | A | A | A |
| 70.44 | 2.97 | 59.73 | 6.53 | N | N | N |
| 44.2 | 1.9 | 0.23 | 6.59 | A | A | A |
| 46.2 | 2.5 | 4.76 | 7.36 | A | A | A |

TABLE 21. Evaluation Results for Cs-137 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

1

2

3

4

5

6

7

8

10

11

12

13

14

15

16

17

18

19

20

21

22

23

24

25

26

27

28

29

30

31

32

33

34

35

36

37

38

39

40

42

43

44

45

46

47

48

49

50

51

52

54

55

56

57

58

59

60

61

62

63

| Reported value [Bq/kg ] 43.47 | Reported uncertainty (k = 1) [Bq/kg ] 2.9 | Relative bias [%] -1.43 | P-Test [%] 8.33 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 45.0 | 0.7 | 2.04 | 5.23 | A | A | A |
| 44.36 | 0.97 | 0.59 | 5.45 | A | A | A |
| 47.1 | 1.0 | 6.80 | 5.42 | A | A | A |
| 43.63 | 0.60 | -1.07 | 5.17 | A | A | A |
| 22.3 | 1.6 | -49.43 | 8.74 | N | N | N |
| 45 | 2 | 2.04 | 6.68 | A | A | A |
| 42.01 | 0.55 | -4.74 | 5.16 | A | A | A |
| 43.5 | 1.8 | -1.36 | 6.48 | A | A | A |
| 45.49 | 1.28 | 3.15 | 5.73 | A | A | A |
| 42.19 | 0.46 | -4.33 | 5.11 | A | A | A |
| 44.8 | 1.3 | 1.59 | 5.77 | A | A | A |
| 44.5 | 1.9 | 0.91 | 6.57 | A | A | A |
| 41.4 | 0.6 | -6.12 | 5.19 | A | A | A |
| 42.4 | 1.3 | -3.85 | 5.86 | A | A | A |
| 43.29 | 2.33 | -1.84 | 7.34 | A | A | A |
| 41.02 | 0.8 | -6.98 | 5.36 | A | A | A |
| 45.2 | 1.98 | 2.49 | 6.64 | A | A | A |
| 44.8 | 2.5 | 1.59 | 7.49 | A | A | A |
| 44.1 | 1.5 | 0.00 | 6.04 | A | A | A |
| 43.07 | 1.42 | -2.34 | 5.98 | A | A | A |
| 44.40 | 2.20 | 0.68 | 7.03 | A | A | A |
| 44.25 | 1.33 | 0.34 | 5.82 | A | A | A |
| 43.55 | 0.38 | -1.25 | 5.06 | A | A | A |
| 41.50 | 1.73 | -5.90 | 6.50 | A | A | A |
| 39.62 | 1.49 | -10.16 | 6.25 | A | A | A |
| 44.6 | 2.8 | 1.13 | 8.02 | A | A | A |
| 42.31 | 1.20 | -4.06 | 5.74 | A | A | A |
| 43.38 | 0.57 | -1.63 | 5.16 | A | A | A |
| 47.73 | 2.51 | 8.23 | 7.25 | A | A | A |
| 44.4 | 1.0 | 0.68 | 5.47 | A | A | A |
| 63.4 |  | 43.76 |  | N |  | N |
| 43.16 | 1.58 | -2.13 | 6.19 | A | A | A |
| 53.20 | 2.54 | 20.63 | 6.91 | N | N | N |
| 44.06 | 2.47 | -0.09 | 7.50 | A | A | A |
| 43.5 | 2.3 | -1.36 | 7.27 | A | A | A |
| 43.9 | 1.0 | -0.45 | 5.48 | A | A | A |
| 44 | 0.5 | -0.23 | 5.12 | A | A | A |
| 43.43 | 2.24 | -1.52 | 7.18 | A | A | A |
| 44.2 | 3.9 | 0.23 | 10.14 | A | A | A |
| 44.4 | 1.4 | 0.68 | 5.90 | A | A | A |
| 48.57 | 1.69 | 10.14 | 6.08 | A | A | A |
| 45.0 | 2.7 | 2.04 | 7.80 | A | A | A |
| 44.4 | 0.9 | 0.68 | 5.38 | A | A | A |
| 39.46 | 1.33 | -10.52 | 6.02 | A | A | A |
| 44.4 | 1.8 | 0.68 | 6.43 | A | A | A |
| 46.6 | 1.9 | 5.67 | 6.44 | A | A | A |
| 40.52 | 0.60 | -8.12 | 5.20 | A | A | A |
| 44.38 | 4.62 | 0.63 | 11.54 | A | A | A |
| 40.9 | 0.9 | -7.26 | 5.45 | A | A | A |
| 45.67 | 2.52 | 3.56 | 7.44 | A | A | A |
| 44.0 | 2.0 | -0.23 | 6.75 | A | A | A |
| 43.71 | 1.55 | -0.88 | 6.12 | A | A | A |
| 45.3 | 4.1 | 2.72 | 10.33 | A | A | A |
| 40.4 | 2.5 | -8.39 | 7.95 | A | A | A |
| 45 | 2 | 2.04 | 6.68 | A | A | A |
| 49 | 4 | 11.11 | 9.57 | A | A | A |
| 46 | 3 | 4.31 | 8.21 | A | A | A |
| 41.21 | 1.74 | -6.55 | 6.54 | A | A | A |
| 44.2 | 1.4 | 0.23 | 5.91 | A | A | A |

TABLE 21. Evaluation Results for Cs-137 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

64

66

67

68

69

70

71

72

73

74

75

77

79

80

81

82

83

84

85

86

87

89

90

91

93

94

95

96

97

98

99

101

102

104

105

106

107

108

110

114

115

116

118

119

120

122

123

124

125

126

129

130

131

132

133

134

135

136

137

139

| Reported value [Bq/kg ] 45.0 | Reported uncertainty (k = 1) [Bq/kg ] 1.3 | Relative bias [%] 2.04 | P-Test [%] 5.76 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 44.26 | 2.72 | 0.36 | 7.92 | A | A | A |
| 46.903 | 2.377 | 6.36 | 7.11 | A | A | A |
| 46.8 | 2.4 | 6.12 | 7.15 | A | A | A |
| 43.2 | 2.1 | -2.04 | 6.97 | A | A | A |
| 43.67 | 2.36 | -0.98 | 7.35 | A | A | A |
| 49.23 | 2.41 | 11.63 | 6.99 | A | A | A |
| 43.7 | 2.4 | -0.91 | 7.42 | A | A | A |
| 41.5 | 1.89 | -5.90 | 6.75 | A | A | A |
| 44.3 | 1.1 | 0.45 | 5.57 | A | A | A |
| 43.7 | 2.3 | -0.91 | 7.25 | A | A | A |
| 45.3 | 1.26 | 2.72 | 5.71 | A | A | A |
| 43.863 | 2.671 | -0.54 | 7.87 | A | A | A |
| 38.1 | 1.07 | -13.61 | 5.72 | A | A | A |
| 44.1 | 1.6 | 0.00 | 6.17 | A | A | A |
| 44.6 | 2.2 | 1.13 | 7.02 | A | A | A |
| 45.2 | 0.7 | 2.49 | 5.22 | A | A | A |
| 43.89 | 0.92 | -0.48 | 5.41 | A | A | A |
| 43.7 | 1.2 | -0.91 | 5.69 | A | A | A |
| 43.1 | 1.4 | -2.27 | 5.95 | A | A | A |
| 46.9 | 2.8 | 6.35 | 7.78 | A | A | A |
| 47.2 | 2.3 | 7.03 | 6.97 | A | A | A |
| 43.0 | 1.8 | -2.49 | 6.51 | A | A | A |
| 44.91 | 0.48 | 1.84 | 5.10 | A | A | A |
| 44 | 4.6 | -0.23 | 11.58 | A | A | A |
| 43.20 | 1.78 | -2.04 | 6.47 | A | A | A |
| 43.4 | 2.2 | -1.59 | 7.11 | A | A | A |
| 42.4 | 3.3 | -3.85 | 9.24 | A | A | A |
| 40.8 | 2.5 | -7.48 | 7.90 | A | A | A |
| 44.9 | 1.07 | 1.81 | 5.53 | A | A | A |
| 40.0 | 2.55 | -9.30 | 8.09 | A | A | A |
| 44.31 | 2.25 | 0.48 | 7.12 | A | A | A |
| 46 | 3 | 4.31 | 8.21 | A | A | A |
| 46.3 | 2.4 | 4.99 | 7.19 | A | A | A |
| 42.99 | 0.44 | -2.52 | 5.09 | A | A | A |
| 43.3 | 0.75 | -1.81 | 5.28 | A | A | A |
| 43.78 | 0.60 | -0.73 | 5.17 | A | A | A |
| 45.2 | 2.3 | 2.49 | 7.13 | A | A | A |
| 43.62 | 3.49 | -1.09 | 9.43 | A | A | A |
| 42.6 | 2.6 | -3.40 | 7.88 | A | A | A |
| 46.8 | 4.2 | 6.12 | 10.27 | A | A | A |
| 43.98 | 4.62 | -0.27 | 11.63 | A | A | A |
| 48.0 | 2.3 | 8.84 | 6.92 | A | A | A |
| 44 | 5 | -0.23 | 12.41 | A | A | A |
| 48.4 | 4 | 9.75 | 9.65 | A | A | A |
| 44.6 | 1.1 | 1.13 | 5.57 | A | A | A |
| 46.3 | 2.7 | 4.99 | 7.67 | A | A | A |
| 42.18 | 1.31 | -4.35 | 5.88 | A | A | A |
| 44.8 | 3.9 | 1.59 | 10.03 | A | A | A |
| 42.68 | 1.396 | -3.22 | 5.97 | A | A | A |
| 47.9004 | 0.9527 | 8.62 | 5.37 | A | A | A |
| 48.5 | 1.9 | 9.98 | 6.34 | A | A | A |
| 50.74 | 0.86 | 15.06 | 5.27 | A | N | W |
| 45.21 | 2.35 | 2.52 | 7.20 | A | A | A |
| 44.8 | 2.7 | 1.59 | 7.82 | A | A | A |
| 43.96 | 3.38 | -0.32 | 9.17 | A | A | A |
| 43.9 | 3.1 | -0.45 | 8.65 | A | A | A |
| 43.69 | 0.92 | -0.93 | 5.41 | A | A | A |
| 43.289 | 1.289 | -1.84 | 5.81 | A | A | A |
| 43.81 | 2.78 | -0.66 | 8.07 | A | A | A |

TABLE 21. Evaluation Results for Cs-137 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

140

141

142

144

145

146

147

148

149

150

151

152

153

154

155

156

157

158

159

160

161

162

163

165

166

167

168

169

170

171

172

173

174

176

178

179

180

181

182

183

184

185

186

188

189

190

192

193

195

196

197

198

199

200

201

202

204

205

206

207

| Reported value [Bq/kg ] 45.26 | Reported uncertainty (k = 1) [Bq/kg ] 3.80 | Relative bias [%] 2.63 | P-Test [%] 9.77 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 43.45 | 3.42 | -1.47 | 9.32 | A | A | A |
| 43.851 | 0.831 | -0.56 | 5.34 | A | A | A |
| 43.2 | 2.1 | -2.04 | 6.97 | A | A | A |
| 49.7 | 1.0 | 12.70 | 5.38 | A | A | A |
| 43.1 | 3.1 | -2.27 | 8.75 | A | A | A |
| 46.0 | 0.31 | 4.31 | 5.03 | A | A | A |
| 35.10 | 0.50 | -20.41 | 5.19 | N | N | N |
| 44.20 | 2.66 | 0.23 | 7.82 | A | A | A |
| 42.80 | 2.58 | -2.95 | 7.82 | A | A | A |
| 38.24 | 0.18 | -13.29 | 5.01 | A | N | W |
| 60.4 | 1.9 | 36.96 | 5.90 | N | N | N |
| 66.76 | 0.2 | 51.38 | 5.00 | N | N | N |
| 37.5 | 0.2 | -14.97 | 5.02 | A | N | W |
| 43.5 | 1.3 | -1.36 | 5.82 | A | A | A |
| 42.64 | 2.21 | -3.31 | 7.19 | A | A | A |
| 45 | 3 | 2.04 | 8.33 | A | A | A |
| 45.63 | 1.98 | 3.47 | 6.61 | A | A | A |
| 45.41 | 1.24 | 2.97 | 5.69 | A | A | A |
| 42.8 | 2.3 | -2.95 | 7.33 | A | A | A |
| 43.2 | 1.28 | -2.04 | 5.80 | A | A | A |
| 44.21 | 0.70 | 0.25 | 5.23 | A | A | A |
| 36.19 | 0.87 | -17.94 | 5.54 | A | N | W |
| 43.6 | 2.2 | -1.13 | 7.10 | A | A | A |
| 44.0 | 3.2 | -0.23 | 8.82 | A | A | A |
| 50 | 1 | 13.38 | 5.37 | A | A | A |
| 44.1 | 3.18 | 0.00 | 8.77 | A | A | A |
| 44.3 | 2.8 | 0.45 | 8.05 | A | A | A |
| 46.5 | 2.2 | 5.44 | 6.88 | A | A | A |
| 43.19 | 1.02 | -2.06 | 5.52 | A | A | A |
| 44.7 | 2.2 | 1.36 | 7.01 | A | A | A |
| 49.6 | 3 | 12.47 | 7.84 | A | A | A |
| 44 | 8.8 | -0.23 | 20.61 | A | N | W |
| 45.08 | 0.48 | 2.22 | 5.10 | A | A | A |
| 46.3 | 3.2 | 4.99 | 8.52 | A | A | A |
| 40.4 | 0.8 | -8.39 | 5.37 | A | A | A |
| 44.192 | 2.5808 | 0.21 | 7.68 | A | A | A |
| 45.7 | 2.04 | 3.63 | 6.69 | A | A | A |
| 44.2 | 1.6 | 0.23 | 6.16 | A | A | A |
| 43.26 | 1.60 | -1.90 | 6.21 | A | A | A |
| 43.7 | 4 | -0.91 | 10.42 | A | A | A |
| 43.98 | 3.30 | -0.27 | 9.01 | A | A | A |
| 45.0 | 1.1 | 2.04 | 5.56 | A | A | A |
| 47.9 | 2.68 | 8.62 | 7.50 | A | A | A |
| 44.6 | 0.4 | 1.13 | 5.07 | A | A | A |
| 46.47 | 4.246 | 5.37 | 10.41 | A | A | A |
| 44.8 | 0.83 | 1.59 | 5.32 | A | A | A |
| 44.6 | 2.4 | 1.13 | 7.34 | A | A | A |
| 43.04 | 0.64 | -2.40 | 5.21 | A | A | A |
| 44.5 | 2.3 | 0.91 | 7.18 | A | A | A |
| 44.7 | 1.5 | 1.36 | 6.01 | A | A | A |
| 46.5 | 3.0 | 5.44 | 8.16 | A | A | A |
| 44.6 | 2.4 | 1.13 | 7.34 | A | A | A |
| 44.6 | 2.37 | 1.13 | 7.29 | A | A | A |
| 45.72 | 2.8 | 3.67 | 7.90 | A | A | A |
| 44.2 | 1.5 | 0.23 | 6.03 | A | A | A |
| 42.50 | 0.27 | -3.63 | 5.03 | A | A | A |
| 45.28 | 2.43 | 2.68 | 7.33 | A | A | A |
| 43.14 | 4.314 | -2.18 | 11.18 | A | A | A |
| 21 | 0.38 | -52.38 | 5.31 | N | N | N |

TABLE 21. Evaluation Results for Cs-137 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

209

210

211

213

214

217

218

221

222

223

224

225

226

227

228

230

232

233

234

235

236

237

239

240

241

243

244

245

246

247

248

249

250

251

252

254

255

256

257

258

259

260

261

262

263

264

265

268

269

270

271

272

273

274

275

276

277

279

281

282

TABLE 21. Evaluation Results for Cs-137 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg by gamma

| Reported value [Bq/kg ] 45.63 | Reported uncertainty (k = 1) [Bq/kg ] 8.239 | Relative bias [%] 3.47 | P-Test [%] 18.73 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 46.8 | 2.3 | 6.12 | 7.00 | A | A | A |
| 45.757 | 0.82507 | 3.76 | 5.30 | A | A | A |
| 42.6 | 2.64 | -3.40 | 7.96 | A | A | A |
| 47.1 | 1.71 | 6.80 | 6.17 | A | A | A |
| 41.9 | 1.00 | -4.99 | 5.53 | A | A | A |
| 44.1 | 2.96 | 0.00 | 8.36 | A | A | A |
| 40.15 | 2.34 | -8.96 | 7.67 | A | A | A |
| 49.0 | 2.5 | 11.11 | 7.14 | A | A | A |
| 38.8 | 2.3 | -12.02 | 7.75 | A | A | A |
| 44 | 1.4 | -0.23 | 5.92 | A | A | A |
| 43.56 | 1.63 | -1.22 | 6.24 | A | A | A |
| 46.4 | 1.9 | 5.22 | 6.45 | A | A | A |
| 43.2 | 1.1 | -2.04 | 5.60 | A | A | A |
| 43.6 | 3.5 | -1.13 | 9.45 | A | A | A |
| 43.07 | 1.82 | -2.34 | 6.54 | A | A | A |
| 423 | 0.6 | 859.18 | 4.99 | N | N | N |
| 45.91 | 5.85 | 4.10 | 13.68 | A | A | A |
| 43.3 | 1.8 | -1.81 | 6.49 | A | A | A |
| 41 | 2 | -7.03 | 6.98 | A | A | A |
| 47.65 | 3.51 | 8.05 | 8.90 | A | A | A |
| 47.91 | 0.45 | 8.64 | 5.08 | A | A | A |
| 41.81 | 0.38 | -5.19 | 5.07 | A | A | A |
| 47.81 | 2.64 | 8.41 | 7.44 | A | A | A |
| 43.1 | 1.2 | -2.27 | 5.71 | A | A | A |
| 48.0 | 4.0 | 8.84 | 9.71 | A | A | A |
| 43.0 | 3.4 | -2.49 | 9.35 | A | A | A |
| 44.2 | 1.0 | 0.23 | 5.48 | A | A | A |
| 42.87 | 0.92 | -2.79 | 5.43 | A | A | A |
| 45.2 | 3 | 2.49 | 8.30 | A | A | A |
| 15.438 | 4.015 | -64.99 | 26.48 | N | N | N |
| 35.8 | 3.19 | -18.82 | 10.21 | A | A | A |
| 46.3 | 1.1 | 4.99 | 5.53 | A | A | A |
| 47.8 | 2.3 | 8.39 | 6.93 | A | A | A |
| 44.7 | 3.0 | 1.36 | 8.36 | A | A | A |
| 58.9 | 5.9 | 33.56 | 11.19 | N | N | N |
| 48.74 | 1.37 | 10.52 | 5.73 | A | A | A |
| 46.4 | 2.1 | 5.22 | 6.74 | A | A | A |
| 45.2 | 2.0 | 2.49 | 6.67 | A | A | A |
| 38.96 | 3.18 | -11.66 | 9.57 | A | A | A |
| 21.7 | 0.64 | -50.79 | 5.80 | N | N | N |
| 42.5 | 1.4 | -3.63 | 5.98 | A | A | A |
| 62 | 11 | 40.59 | 18.43 | N | A | N |
| 53.72 | 1.95 | 21.81 | 6.17 | N | N | N |
| 45.77 | 3.82 | 3.79 | 9.72 | A | A | A |
| 47.9754 |  | 8.79 |  | A |  |  |
| 46.1 | 3.3 | 4.54 | 8.73 | A | A | A |

Labcode

283

284

285

286

287

290

291

292

293

294

295

296

298

299

300

301

302

303

304

305

306

307

308

309

311

314

315

316

317

318

319

322

327

328

329

330

331

333

334

335

336

337

340

341

342

343

344

Activity concentration  [ Bq/kg ]

5

6

Reported Results for Po-210 determination in Sample 2

TABLE 22. Target values

| Parameter | Value |
| --- | --- |
| Activity Concentration [ Bq/kg ] | 5.21 |
| Uncertainty [ Bq/kg ] | 0.26 |
| Maximum Acceptable Relative Bias, [%] | 30 |

Warning 0%

Not Accepted 4%

Accepted 8%

Not Reported 88%

7

8

FIGURE 21. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

2

3

4

223

235

89

232

166

185

107

127

344

270

35

44

268

188

272

112

115

133

114

43

72

11

85

248

67

186

199

22

178

202

184

225

330

140

247

307

65

104

57

Laboratory code

FIGURE 22. Reported results and evaluation parameters for Po-210. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

TABLE 23. Evaluation Results for Po-210 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg by alpha

| Reported value [Bq/kg ] 5.30 | Reported uncertainty (k = 1) [Bq/kg ] 0.24 | Relative bias [%] 1.73 | P-Test [%] 6.74 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 5.88 | 0.82 | 12.86 | 14.81 | A | A | A |
| 4.54 | 0.33 | -12.86 | 8.82 | A | A | A |
| 5.11 | 0.32 | -1.92 | 8.01 | A | A | A |
| 4.70 | 0.31 | -9.79 | 8.27 | A | A | A |
| 20.8 | 1.4 | 299.23 | 8.38 | N | N | N |
| 14.37 | 7.28 | 175.82 | 50.91 | N | N | N |
| 5.42 | 0.09 | 4.03 | 5.26 | A | A | A |
| 5.23 | 0.33 | 0.38 | 8.04 | A | A | A |
| 5.33 | 0.23 | 2.30 | 6.60 | A | A | A |
| 2.81 | 0.25 | -46.07 | 10.20 | N | N | N |
| 16.88 | 0.07 | 223.99 | 5.01 | N | N | N |
| 4.21 | 0.42 | -19.19 | 11.15 | A | A | A |
| 4.9 | 0.6 | -5.95 | 13.22 | A | A | A |
| 5.09 | 0.51 | -2.30 | 11.19 | A | A | A |
| 4.98 | 0.23 | -4.41 | 6.80 | A | A | A |
| 4.24 | 0.59 | -18.62 | 14.78 | A | A | A |
| 5.08 | 0.20 | -2.50 | 6.36 | A | A | A |
| 10.70 | 0.59 | 105.37 | 7.44 | N | N | N |
| 3.8 | 0.7 | -27.06 | 19.09 | A | A | A |
| 5.89 | 0.20 | 13.05 | 6.04 | A | A | A |
| 6.70 | 0.70 | 28.60 | 11.58 | A | A | A |
| 3.80 | 0.63 | -27.06 | 17.31 | A | A | A |
| 5.6 | 0.3 | 7.49 | 7.32 | A | A | A |
| 4.9 | 0.3 | -5.95 | 7.90 | A | A | A |
| 5.81 | 0.54 | 11.52 | 10.55 | A | A | A |
| 5.95 | 0.19 | 14.20 | 5.92 | A | A | A |
| 0.52 | 0.05 | -90.02 | 10.83 | N | N | N |
| 7.7 | 1.4 | 47.79 | 18.85 | N | A | N |
| 2.9 | 0.4 | -44.34 | 14.67 | N | N | N |
| 2.74 | 0.21 | -47.41 | 9.15 | N | N | N |
| 11.26 | 0.20 | 116.12 | 5.30 | N | N | N |
| 5.4 | 0.3 | 3.65 | 7.47 | A | A | A |
| 4.82 | 0.44 | -7.49 | 10.40 | A | A | A |
| 4.54 | 0.31 | -12.86 | 8.46 | A | A | A |
| 4.9 | 0.31 | -5.95 | 8.06 | A | A | A |
| 13.25 | 0.48 | 154.32 | 6.17 | N | N | N |
| 9.2 | 1.3 | 76.58 | 14.99 | N | N | N |
| 4.5 | 0.3 | -13.63 | 8.33 | A | A | A |

Labcode

11

22

35

43

44

57

65

67

72

85

89

104

107

112

114

115

127

133

140

166

178

184

185

186

188

199

202

223

225

232

235

247

248

268

270

272

307

330

344

Activity concentration  [ Bq/kg ]

5

6

Reported Results for Pb-210 determination in Sample 2

TABLE 24. Target values

| Parameter | Value |
| --- | --- |
| Activity Concentration [ Bq/kg ] | 5.23 |
| Uncertainty [ Bq/kg ] | 0.26 |
| Maximum Acceptable Relative Bias, [%] | 30 |

Warning 2%

Not Accepted 5%

Accepted 12%

Not Reported 81%

7

8

FIGURE 23. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

2

3

4

243

89

89

158

261

57

166

127

185

66

145

233

189

202

247

272

85

85

334

119

268

247

72

270

173

119

29

114

222

248

11

300

330

43

125

173

170

131

133

79

301

67

203

262

135

40

200

184

135

184

104

180

307

206

186

98

299

65

149

Laboratory code

FIGURE 24. Reported results and evaluation parameters for Pb-210. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

TABLE 25. Evaluation Results for Pb-210 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg by beta

| Reported value [Bq/kg ] | Reported uncertainty (k = 1) [Bq/kg ] | Relative bias [%] | P-Test [%] | Trueness evaluation | Precision evaluation | Final Score |
| --- | --- | --- | --- | --- | --- | --- |
| 5.21 | 0.45 | -0.38 | 9.97 | A | A | A |
| 5.32 | 0.34 | 1.72 | 8.10 | A | A | A |
| 3.6 | 0.4 | -31.17 | 12.17 | N | A | N |
| 23.31 | 10.06 | 345.70 | 43.44 | N | N | N |
| 4.99 | 0.21 | -4.59 | 6.51 | A | A | A |
| 4.66 | 0.11 | -10.90 | 5.50 | A | A | A |
| 2.81 | 0.25 | -46.27 | 10.19 | N | N | N |
| 7.36 | 1.34 | 40.73 | 18.87 | N | A | N |
| 5.12 | 0.7 | -2.10 | 14.55 | A | A | A |
| 4.71 | 0.5 | -9.94 | 11.72 | A | A | A |
| 3.88 | 0.376 | -25.81 | 10.89 | A | A | A |
| 5.57 | 0.36 | 6.50 | 8.15 | A | A | A |
| 6.1 | 1.3 | 16.63 | 21.88 | A | A | A |
| 4.18 | 0.60 | -20.08 | 15.19 | A | A | A |
| 3.8 | 0.8 | -27.34 | 21.63 | A | A | A |
| 5.38 | 1 | 2.87 | 19.24 | A | A | A |
| 6.70 | 0.80 | 28.11 | 12.93 | A | A | A |
| 3.90 | 0.72 | -25.43 | 19.12 | A | A | A |
| 10.0 | 0.3 | 91.20 | 5.81 | N | N | N |
| 4.4 | 0.9 | -15.87 | 21.05 | A | A | A |
| 4.52 | 0.28 | -13.58 | 7.94 | A | A | A |
| 5.73 | 0.44 | 9.56 | 9.15 | A | A | A |
| 4.88 | 0.51 | -6.69 | 11.57 | A | A | A |
| 5.2 | 0.3 | -0.57 | 7.62 | A | A | A |
| 5.78 | 1.92 | 10.52 | 33.59 | A | N | W |
| 4.81 | 0.18 | -8.03 | 6.22 | A | A | A |
| 4.99 | 0.40 | -4.59 | 9.43 | A | A | A |
| 4.63 | 0.75 | -11.47 | 16.94 | A | A | A |
| 5.3 | 0.6 | 1.34 | 12.36 | A | A | A |
| 8.54 | 0.74 | 63.29 | 9.99 | N | N | N |
| 5.3 | 1.0 | 1.34 | 19.51 | A | A | A |

Labcode

11

43

57

65

72

85

89

104

114

119

127

133

135

145

166

173

184

185

186

189

202

203

247

248

262

268

270

272

300

307

330

TABLE 26. Evaluation Results for Pb-210 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg by gamma

| Reported value [Bq/kg ] | Reported uncertainty (k = 1) [Bq/kg ] | Relative bias [%] | P-Test [%] | Trueness evaluation | Precision evaluation | Final Score |
| --- | --- | --- | --- | --- | --- | --- |
| 5.1 | 1.7 | -2.49 | 33.70 | A | N | W |
| 6.29 | 0.595 | 20.27 | 10.69 | A | A | A |
| 3.9 | 0.7 | -25.43 | 18.62 | A | A | A |
| 5.63 | 1.21 | 7.65 | 22.06 | A | A | A |
| 5.57 | 0.99 | 6.50 | 18.46 | A | A | A |
| 4.68 | 0.40 | -10.52 | 9.89 | A | A | A |
| 2.81 | 0.25 | -46.27 | 10.19 | N | N | N |
| 13.01 | 1.46 | 148.76 | 12.27 | N | N | N |
| 5.08 | 2.4 | -2.87 | 47.50 | A | N | W |
| 5.32 | 1.75 | 1.72 | 33.27 | A | N | W |
| 5.43 | 0.71 | 3.82 | 13.99 | A | A | A |
| 6.7 | 2.4 | 28.11 | 36.16 | A | N | W |
| 26.2 | 6.98 | 400.96 | 27.10 | N | N | N |
| 3.53 | 0.13 | -32.50 | 6.19 | N | N | N |
| 5.4 | 1.6 | 3.25 | 30.04 | A | N | W |
| 5.07 | 1.32 | -3.06 | 26.51 | A | A | A |
| 7.54 | 1.46 | 44.17 | 19.99 | N | A | N |
| 6.79 | 0.85 | 29.83 | 13.47 | A | A | A |
| 6.35 | 3.25 | 21.41 | 51.42 | A | N | W |
| 9.509 | 2.086 | 81.82 | 22.49 | N | N | N |
| 5.14 | 1.00 | -1.72 | 20.08 | A | A | A |
| 4.21 | 0.78 | -19.50 | 19.18 | A | A | A |
| 1.8 | 0.2 | -65.58 | 12.17 | N | N | N |

Labcode

29

40

66

67

79

85

89

98

119

125

131

135

149

158

170

173

180

184

200

206

222

233

243

TABLE 26. Evaluation Results for Pb-210 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg by gamma

| Reported value [Bq/kg ] 4.60 | Reported uncertainty (k = 1) [Bq/kg ] 0.4 | Relative bias [%] -12.05 | P-Test [%] 10.02 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 3.55 | 0.40 | -32.12 | 12.32 | N | N | N |
| 15.8 | 1.5 | 202.10 | 10.72 | N | N | N |
| 5.622 | 1.294 | 7.50 | 23.55 | A | A | A |
| 4.7 | 1.0 | -10.13 | 21.85 | A | A | A |

Labcode

247

261

299

301

334

Reported Results for Ra-226 determination in Sample 2

TABLE 27. Target values

| Parameter | Value |
| --- | --- |
| Activity Concentration [ Bq/kg ] | 6.32 |
| Uncertainty [ Bq/kg ] | 0.32 |
| Maximum Acceptable Relative Bias, [%] | 30 |

Warning 3%

Not Accepted 11%

Activity concentration  [ Bq/kg ]

6

8

Accepted 19%

Not Reported 67%

10

FIGURE 25. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

4

172

140

264

335

167

89

89

225

202

333

72

206

196

247

202

126

222

247

131

201

334

1

159

38

178

29

6

67

149

155

330

330

145

184

184

294

157

43

290

144

127

135

235

55

245

200

262

112

35

293

272

258

44

215

215

17

316

133

296

248

100

268

188

183

102

37

40

11

37

135

40

102

307

224

85

149

171

173

27

104

186

85

104

270

57

8

39

147

69

136

321

110

189

54

108

31

185

166

33

323

160

65

129

129

226

Laboratory code

FIGURE 26. Reported results and evaluation parameters for Ra-226. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

| Reported value [Bq/kg ] 5.10 | Reported uncertainty (k = 1) [Bq/kg ] 0.10 | Relative bias [%] -19.30 | P-Test [%] 5.43 | Trueness evaluation A | Precision evaluation N | Final Score W |
| --- | --- | --- | --- | --- | --- | --- |
| 6.81 | 0.60 | 7.75 | 10.16 | A | A | A |
| 5.5 | 0.2 | -12.97 | 6.23 | A | A | A |
| 11 | 5 | 74.05 | 45.74 | N | N | N |
| 6.41 | 0.40 | 1.42 | 8.04 | A | A | A |
| 6.65 | 0.62 | 5.22 | 10.61 | A | A | A |
| 6.92 | 0.28 | 9.49 | 6.48 | A | A | A |
| 6.1 | 0.4 | -3.48 | 8.28 | A | A | A |
| 6.49 | 0.39 | 2.69 | 7.86 | A | A | A |
| 8.2 | 0.4 | 29.75 | 7.03 | A | N | W |
| 21.32 | 2.67 | 237.34 | 13.51 | N | N | N |
| 5.71 | 0.16 | -9.65 | 5.79 | A | A | A |
| 9.219 | 0.824 | 45.87 | 10.27 | N | N | N |
| 3.67 | 0.02 | -41.93 | 5.09 | N | N | N |
| 7.03 | 0.17 | 11.23 | 5.61 | A | A | A |
| 2.81 | 0.43 | -55.54 | 16.12 | N | N | N |
| 6.55 | 0.57 | 3.64 | 10.07 | A | A | A |
| 6.65 | 0.29 | 5.22 | 6.68 | A | A | A |
| 7.66 | 0.65 | 21.20 | 9.88 | A | A | A |
| 6.4 | 0.2 | 1.27 | 5.95 | A | A | A |
| 6.15 | 0.712 | -2.69 | 12.64 | A | A | A |
| 25.25 | 19.46 | 299.53 | 77.24 | N | N | N |
| 6.51 | 0.40 | 3.01 | 7.96 | A | A | A |
| 6.9 | 1.2 | 9.18 | 18.11 | A | A | A |
| 9.7 | 0.3 | 53.48 | 5.93 | N | N | N |
| 0.14 | 0.03 | -97.78 | 22.02 | N | N | N |
| 5.91 | 0.39 | -6.49 | 8.32 | A | A | A |
| 5.85 | 0.57 | -7.44 | 10.98 | A | A | A |
| 5.9 | 1.0 | -6.65 | 17.69 | A | A | A |
| 6.03 | 0.50 | -4.59 | 9.72 | A | A | A |
| 14.0 | 0.9 | 121.52 | 8.18 | N | N | N |
| 2.81 | 0.62 | -55.54 | 22.64 | N | A | N |
| 0.078 | 0.023 | -98.77 | 29.92 | N | N | N |
| 7.34 | 0.5 | 16.14 | 8.49 | A | A | A |
| 6.0 | 0.9 | -5.06 | 15.83 | A | A | A |
| 11.10 | 2.13 | 75.63 | 19.85 | N | N | N |
| 7.8 | 0.3 | 23.42 | 6.36 | A | N | W |
| 6.6 | 0.7 | 4.43 | 11.75 | A | A | A |
| 10.6 | 0.6 | 67.72 | 7.59 | N | N | N |
| 4.11 | 0.28 | -34.97 | 8.49 | N | N | N |
| 4.35 | 0.15 | -31.17 | 6.13 | N | N | N |
| 6.5 | 0.6 | 2.85 | 10.53 | A | A | A |
| 3.5 | 0.3 | -44.62 | 9.96 | N | N | N |
| 6.3 | 0.5 | -0.32 | 9.41 | A | A | A |
| 6.31 | 0.48 | -0.16 | 9.14 | A | A | A |
| 4.12 | 0.15 | -34.81 | 6.24 | N | N | N |
| 6.52 | 0.42 | 3.16 | 8.19 | A | A | A |
| 6.48 | 0.32 | 2.53 | 7.07 | A | A | A |
| 6.38 | 0.263 | 0.95 | 6.53 | A | A | A |
| 6.59 | 0.59 | 4.27 | 10.29 | A | A | A |
| 8.15 | 0.70 | 28.96 | 9.97 | A | N | W |
| 6.45 | 0.74 | 2.06 | 12.54 | A | A | A |
| 6.1 | 0.39 | -3.48 | 8.16 | A | A | A |
| 6.51 | 0.6 | 3.01 | 10.52 | A | A | A |
| 6.93 | 0.10 | 9.65 | 5.26 | A | A | A |
| 9.72 | 1.35 | 53.80 | 14.78 | N | N | N |
| 5.9 | 1.2 | -6.65 | 20.96 | A | A | A |
| 3.62 | 0.34 | -42.72 | 10.67 | N | N | N |

TABLE 28. Evaluation Results for Ra-226 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg by alpha

Labcode

1

11

29

31

35

37

40

43

44

57

65

67

69

72

85

89

100

102

104

112

127

129

133

135

136

140

145

149

155

157

160

167

172

173

184

185

186

188

189

196

202

215

225

235

245

247

248

258

262

268

270

272

290

296

307

321

330

333

TABLE 29. Evaluation Results for Ra-226 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg by gamma

| Reported value [Bq/kg ] 5.57 | Reported uncertainty (k = 1) [Bq/kg ] 0.38 | Relative bias [%] -11.87 | P-Test [%] 8.50 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 8.41 | 0.33 | 33.07 | 6.41 | N | N | N |
| 6.5 | 0.3 | 2.85 | 6.85 | A | A | A |
| 7.44 | 0.80 | 17.72 | 11.89 | A | A | A |
| 12.2 | 1.5 | 93.04 | 13.30 | N | N | N |
| 6.82 | 0.88 | 7.91 | 13.86 | A | A | A |
| 5.42 | 0.82 | -14.24 | 15.95 | A | A | A |
| 8.5 | 0.5 | 34.49 | 7.76 | N | N | N |
| 6.71 | 0.711 | 6.17 | 11.74 | A | A | A |
| 10.8 | 5.4 | 70.89 | 50.26 | N | N | N |
| 6.3 | 0.4 | -0.32 | 8.12 | A | A | A |
| 7.80 | 0.84 | 23.42 | 11.90 | A | A | A |
| 2.81 | 0.43 | -55.54 | 16.12 | N | N | N |
| 6.92 | 1.61 | 9.49 | 23.81 | A | A | A |
| 7.87 | 1.18 | 24.53 | 15.83 | A | A | A |
| 10.9 | 1.5 | 72.47 | 14.66 | N | N | N |
| 9.93 | 0.52 | 57.12 | 7.28 | N | N | N |
| 4.4 | 0.2 | -30.38 | 6.80 | N | N | N |
| 25.25 | 19.50 | 299.53 | 77.39 | N | N | N |
| 4.77 | 0.21 | -24.53 | 6.71 | A | N | W |
| 6.2 | 1.6 | -1.90 | 26.30 | A | A | A |
| 6.12 | 0.46 | -3.16 | 9.06 | A | A | A |
| 8.54 | 1.75 | 35.13 | 21.11 | N | A | N |
| 7.10 | 1.96 | 12.34 | 28.07 | A | A | A |
| 5.33 | 0.50 | -15.66 | 10.66 | A | A | A |
| 7.18 | 1.85 | 13.61 | 26.26 | A | A | A |
| 5.48 | 0.33 | -13.29 | 7.87 | A | A | A |
| 6.6 | 1.6 | 4.43 | 24.77 | A | A | A |
| 5.96 | 0.72 | -5.70 | 13.10 | A | A | A |
| 6.35 | 2.02 | 0.47 | 32.21 | A | N | W |
| 4.8 | 0.9 | -24.05 | 19.42 | A | A | A |
| 3.50 | 0.28 | -44.62 | 9.47 | N | N | N |
| 4.017 | 3.612 | -36.44 | 90.06 | N | N | N |
| 6.5 | 0.6 | 2.85 | 10.53 | A | A | A |
| 4.55 | 0.43 | -28.01 | 10.72 | A | N | W |
| 7 | 0.6 | 10.76 | 9.96 | A | A | A |
| 30.11 | 0.2 | 376.42 | 5.11 | N | N | N |
| 4.56 | 0.50 | -27.85 | 12.08 | A | A | A |
| 2.349 | 0.2797 | -62.83 | 12.94 | N | N | N |
| 6.43 | 0.39 | 1.74 | 7.90 | A | A | A |
| 6.0 | 2.6 | -5.06 | 43.63 | A | N | W |
| 6.5 | 1 | 2.85 | 16.20 | A | A | A |
| 12.25 | 0.32 | 93.83 | 5.70 | N | N | N |
| 5.9 | 0.5 | -6.65 | 9.87 | A | A | A |
| 4.9 | 0.3 | -22.47 | 7.94 | A | N | W |
| 2.40 | 0.19 | -62.03 | 9.40 | N | N | N |

Labcode

6

8

17

27

33

37

38

39

40

54

55

85

89

102

104

108

110

126

129

131

135

144

147

149

159

171

178

183

184

200

201

202

206

215

222

224

226

247

264

293

294

316

323

330

334

335

Reported Results for K-40 determination in Sample 4

TABLE 30. Target values

59

Activity concentration [ Bq/kg ]

300	400	500	600	700	800	900

FIGURE 28. Reported results and evaluation parameters for K-40. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

172

76

336

243

30

249

33

92

335

311

239

38

306

159

294

170

223

85

178

314

209

125

142

FIGURE 27. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

342

183

317

167

39

90

217

60

146

222

260

224

333

79

343

18

233

133

169

43

157

21

66

270

316

75

8

173

302

25

50

120

231

6

104

309

163

130

46

226

51

Not Accepted 8%

12

Parameter

Activity Concentration [ Bq/kg ] Uncertainty [ Bq/kg ]

Maximum Acceptable Relative Bias, [%]

112

174

232

201

307

107

308

158

300

179

218

128

10

245

71

Warning 9%

73

204

299

329

155

54

180

11

275

141

135

228

263

273

221

154

268

Laboratory code

182

Accepted 51%

22

156

69

271

137

168

259

185

246

298

134

166

2

7

5

17

296

32

247

23

122

27

145

151

19

248

197

62

303

328

162

230

140

105

264

Not Reported 32%

34

258

337

293

47

1

295

42

52

31

119

58

199

188

53

202

Value 584

20

30

132

131

279

86

144

200

277

234

205

48

153

251

126

227

81

244

74

55

28

190

57

147

272

67

240

108

63

123

77

72

213

29

184

235

102

210

206

118

186

36

344

281

176

165

14

171

83

84

70

99

97

301

334

149

87

315

89

252

13

323

68

330

96

110

214

318

241

93

236

| Reported value [Bq/kg ] 630.6 | Reported uncertainty (k = 1) [Bq/kg ] 35.8 | Relative bias [%] 7.98 | P-Test [%] 6.63 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 602.09 | 24.76 | 3.10 | 5.35 | A | A | A |
| 604.03 | 34.07 | 3.43 | 6.60 | A | A | A |
| 572 | 24.0 | -2.05 | 5.42 | A | A | A |
| 602.9 | 28.2 | 3.24 | 5.80 | A | A | A |
| 568.00 | 24.80 | -2.74 | 5.55 | A | A | A |
| 585 | 18 | 0.17 | 4.60 | A | A | A |
| 589 | 46 | 0.86 | 8.53 | A | A | A |
| 579 | 25 | -0.86 | 5.51 | A | A | A |
| 837 | 21.2 | 43.32 | 4.26 | N | N | N |
| 708 | 80 | 21.23 | 11.81 | A | A | A |
| 604.4 | 41.6 | 3.49 | 7.69 | A | A | A |
| 555.45 | 39.41 | -4.89 | 7.88 | A | A | A |
| 615 | 42 | 5.31 | 7.64 | A | A | A |
| 561 | 33.3 | -3.94 | 6.85 | A | A | A |
| 595 | 32.2 | 1.88 | 6.40 | A | A | A |
| 607 | 33 | 3.94 | 6.43 | A | A | A |
| 570.3 | 57.1 | -2.35 | 10.58 | A | A | A |
| 613.0 | 38.8 | 4.97 | 7.20 | A | A | A |
| 658 | 49 | 12.67 | 8.20 | A | A | A |
| 680 | 80 | 16.44 | 12.25 | A | A | A |
| 298 | 25 | -48.97 | 9.06 | N | N | N |
| 635 | 44 | 8.73 | 7.73 | A | A | A |
| 606.41 | 18.47 | 3.84 | 4.58 | A | A | A |
| 353.8 | 24.5 | -39.42 | 7.73 | N | N | N |
| 624 | 31.5 | 6.85 | 6.10 | A | A | A |
| 697 | 63 | 19.35 | 9.67 | A | A | A |
| 457 | 38 | -21.75 | 8.99 | A | A | A |
| 535 | 26 | -8.39 | 5.95 | A | A | A |
| 632.2 | 60.00 | 8.25 | 10.09 | A | A | A |
| 560 | 54 | -4.11 | 10.23 | A | A | A |
| 576 | 18 | -1.37 | 4.64 | A | A | A |
| 630 | 20 | 7.88 | 4.67 | A | A | A |
| 648 | 35 | 10.96 | 6.40 | A | A | A |
| 570.38 | 16.68 | -2.33 | 4.50 | A | A | A |
| 577.7 | 31.9 | -1.08 | 6.50 | A | A | A |
| 634 | 33 | 8.56 | 6.23 | A | A | A |
| 640.2 | 36.1 | 9.62 | 6.60 | A | A | A |
| 588 | 10 | 0.68 | 3.82 | A | A | A |
| 656.9 | 35.4 | 12.48 | 6.39 | A | A | A |
| 659 | 71 | 12.84 | 11.31 | A | A | A |
| 635.80 | 95.37 | 8.87 | 15.39 | A | A | A |
| 542.69 | 27.43 | -7.07 | 6.11 | A | A | A |
| 616 | 27.0 | 5.48 | 5.56 | A | A | A |
| 665 | 20 | 13.87 | 4.56 | A | N | W |
| 562.91 | 12.68 | -3.61 | 4.10 | A | A | A |
| 662.12 | 22.85 | 13.38 | 4.86 | A | N | W |
| 876 | 37 | 50.00 | 5.44 | N | N | N |
| 599.00 | 26.45 | 2.57 | 5.59 | A | A | A |
| 738.7 | 39.1 | 26.49 | 6.30 | A | N | W |
| 586 | 11 | 0.34 | 3.91 | A | A | A |
| 676.7 | 16.7 | 15.87 | 4.22 | A | N | W |
| 587 | 24 | 0.51 | 5.33 | A | A | A |
| 656.6 | 33.81 | 12.43 | 6.18 | A | A | A |
| 563.88 | 24.43 | -3.45 | 5.52 | A | A | A |
| 51.5 | 7.18 | -91.18 | 14.36 | N | N | N |
| 673 | 23 | 15.24 | 4.84 | A | N | W |
| 551 | 30.5 | -5.65 | 6.51 | A | A | A |
| 652 | 19 | 11.64 | 4.50 | A | N | W |
| 716.82 | 30.29 | 22.74 | 5.44 | A | N | W |

TABLE 31. Evaluation Results for K-40 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

1

2

5

6

7

8

10

11

12

13

14

17

18

19

21

22

23

25

27

28

29

30

31

32

33

34

36

38

39

42

43

46

47

48

50

51

52

53

54

55

57

58

60

62

63

66

67

68

69

70

71

72

73

74

75

76

77

79

81

83

| Reported value [Bq/kg ] 734 | Reported uncertainty (k = 1) [Bq/kg ] 70.3 | Relative bias [%] 25.68 | P-Test [%] 10.17 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 513 | 103 | -12.16 | 20.37 | A | A | A |
| 644 | 35 | 10.27 | 6.42 | A | A | A |
| 793.08 | 6.43 | 35.80 | 3.52 | N | N | N |
| 800.1 | 40.0 | 37.00 | 6.06 | N | N | N |
| 535.04 | 39.95 | -8.38 | 8.21 | A | A | A |
| 358.8 | 11.97 | -38.56 | 4.78 | N | N | N |
| 1280 | 71 | 119.18 | 6.52 | N | N | N |
| 892.69 | 15.38 | 52.86 | 3.83 | N | N | N |
| 753.68 | 13.12 | 29.05 | 3.84 | A | N | W |
| 750 | 30 | 28.42 | 5.27 | A | N | W |
| 688.1 | 18.5 | 17.83 | 4.35 | A | N | W |
| 572.67 | 22.48 | -1.94 | 5.21 | A | A | A |
| 622.2 | 37.7 | 6.54 | 6.96 | A | A | A |
| 583 | 20 | -0.17 | 4.85 | A | A | A |
| 664 | 13 | 13.70 | 3.94 | A | N | W |
| 910.19 | 18.04 | 55.85 | 3.96 | N | N | N |
| 580 | 20 | -0.68 | 4.86 | A | A | A |
| 693.2 | 29.3 | 18.70 | 5.44 | A | N | W |
| 635 | 30 | 8.73 | 5.84 | A | A | A |
| 570.9 | 18.73 | -2.24 | 4.74 | A | A | A |
| 607 | 16 | 3.94 | 4.32 | A | A | A |
| 666 | 28 | 14.04 | 5.42 | A | N | W |
| 524.78 | 51.35 | -10.14 | 10.37 | A | A | A |
| 649 | 15 | 11.13 | 4.13 | A | N | W |
| 584.52 | 17.26 | 0.09 | 4.52 | A | A | A |
| 575 | 20 | -1.54 | 4.88 | A | A | A |
| 642.72 | 29.33 | 10.05 | 5.71 | A | A | A |
| 642.5 | 57.8 | 10.02 | 9.63 | A | A | A |
| 557 | 38.0 | -4.62 | 7.63 | A | A | A |
| 602 | 16 | 3.08 | 4.33 | A | A | A |
| 590 | 40 | 1.03 | 7.60 | A | A | A |
| 599.79 | 35.78 | 2.70 | 6.88 | A | A | A |
| 620 | 19 | 6.16 | 4.60 | A | A | A |
| 589.4 | 26.3 | 0.92 | 5.62 | A | A | A |
| 525.26 | 28.522 | -10.06 | 6.42 | A | A | A |
| 644.1 | 9.4 | 10.29 | 3.72 | A | N | W |
| 613 | 48 | 4.97 | 8.55 | A | A | A |
| 543.6 | 22.5 | -6.92 | 5.37 | A | A | A |
| 659.78 | 35.54 | 12.98 | 6.38 | A | A | A |
| 781.37 | 43.185 | 33.80 | 6.50 | N | N | N |
| 613.8 | 33.3 | 5.10 | 6.42 | A | A | A |
| 648.451 | 28.823 | 11.04 | 5.61 | A | A | A |
| 593 | 26 | 1.54 | 5.56 | A | A | A |
| 588 | 41 | 0.68 | 7.77 | A | A | A |
| 595 | 24 | 1.88 | 5.29 | A | A | A |
| 561 | 23 | -3.94 | 5.34 | A | A | A |
| 583.58 | 20.92 | -0.07 | 4.96 | A | A | A |
| 497 | 41 | -14.90 | 8.93 | A | A | A |
| 617 | 23 | 5.65 | 5.06 | A | A | A |
| 574.6 | 33.6 | -1.61 | 6.78 | A | A | A |
| 707.95 | 14.30 | 21.22 | 3.98 | A | N | W |
| 602 | 78 | 3.08 | 13.40 | A | A | A |
| 533.00 | 38.45 | -8.73 | 7.99 | A | A | A |
| 600 | 30 | 2.74 | 6.06 | A | A | A |
| 559 | 55 | -4.28 | 10.42 | A | A | A |
| 508 | 22 | -13.01 | 5.52 | A | A | A |
| 710 | 29 | 21.58 | 5.33 | A | N | W |
| 50.2 | 2.74 | -91.40 | 6.44 | N | N | N |
| 569 | 27.2 | -2.57 | 5.88 | A | A | A |

TABLE 31. Evaluation Results for K-40 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

84

85

86

87

89

90

92

93

96

97

99

102

104

105

107

108

110

112

118

119

120

122

123

125

126

128

130

131

132

133

134

135

137

140

141

142

144

145

146

147

149

151

153

154

155

156

157

158

159

162

163

165

166

167

168

169

170

171

172

173

| Reported value [Bq/kg ] 580 | Reported uncertainty (k = 1) [Bq/kg ] 25 | Relative bias [%] -0.68 | P-Test [%] 5.51 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 705 | 23 | 20.72 | 4.73 | A | N | W |
| 518 | 29 | -11.30 | 6.56 | A | A | A |
| 584 | 22 | 0.00 | 5.09 | A | A | A |
| 588.63 | 31.54 | 0.79 | 6.36 | A | A | A |
| 594 | 54 | 1.71 | 9.71 | A | A | A |
| 528 | 40 | -9.59 | 8.31 | A | A | A |
| 683 | 82 | 16.95 | 12.48 | A | A | A |
| 601.60 | 78.21 | 3.01 | 13.44 | A | A | A |
| 694 | 19 | 18.84 | 4.38 | A | N | W |
| 640 | 50 | 9.59 | 8.53 | A | A | A |
| 658.5 | 27.1 | 12.76 | 5.35 | A | A | A |
| 615.9755 | 13.5813 | 5.48 | 4.07 | A | A | A |
| 637.67 | 13.5 | 9.19 | 4.03 | A | A | A |
| 644.6 | 57.8 | 10.38 | 9.60 | A | A | A |
| 581.6 | 24.6 | -0.41 | 5.44 | A | A | A |
| 640.2 | 49.1 | 9.62 | 8.40 | A | A | A |
| 587 | 52 | 0.51 | 9.50 | A | A | A |
| 646.37 | 14.85 | 10.68 | 4.12 | A | N | W |
| 692.66 | 29.19 | 18.61 | 5.43 | A | N | W |
| 519.50 | 42.83 | -11.04 | 8.93 | A | A | A |
| 688.43 | 29.27 | 17.88 | 5.46 | A | N | W |
| 680 | 32 | 16.44 | 5.82 | A | N | W |
| 1020 | 48 | 74.66 | 5.82 | N | N | N |
| 540.1 | 44.9 | -7.52 | 8.99 | A | A | A |
| 584 | 60 | 0.00 | 10.83 | A | A | A |
| 591.40 | 7.90 | 1.27 | 3.68 | A | A | A |
| 545.00 | 22.20 | -6.68 | 5.32 | A | A | A |
| 509.00 | 22.60 | -12.84 | 5.61 | A | A | A |
| 548 | 23 | -6.16 | 5.42 | A | A | A |
| 576.10 | 2.38 | -1.35 | 3.45 | A | A | A |
| 651.2 | 3 | 11.51 | 3.46 | A | N | W |
| 590 | 40 | 1.03 | 7.60 | A | A | A |
| 619.47 | 21.58 | 6.07 | 4.89 | A | A | A |
| 570.95 | 10.11 | -2.23 | 3.86 | A | A | A |
| 580 | 25 | -0.68 | 5.51 | A | A | A |
| 555.48 | 25.14 | -4.88 | 5.68 | A | A | A |
| 646.21 | 30.73 | 10.65 | 5.86 | A | A | A |
| 684.5 | 40.1 | 17.21 | 6.79 | A | A | A |
| 1881.4 | 129.07 | 222.16 | 7.67 | N | N | N |
| 426.69 | 10.24 | -26.94 | 4.18 | A | N | W |
| 664 | 34 | 13.70 | 6.16 | A | A | A |
| 1268 | 94.8 | 117.12 | 8.22 | N | N | N |
| 276 | 27 | -52.74 | 10.36 | N | N | N |
| 654 | 47 | 11.99 | 7.96 | A | A | A |
| 585.7 | 36.6 | 0.29 | 7.13 | A | A | A |
| 602 | 31 | 3.08 | 6.18 | A | A | A |
| 606.8 | 54.3 | 3.90 | 9.58 | A | A | A |
| 615.4 | 25.6 | 5.38 | 5.39 | A | A | A |
| 315 | 16 | -46.06 | 6.13 | N | N | N |
| 648.71 | 27.59 | 11.08 | 5.46 | A | A | A |
| 831 | 79 | 42.29 | 10.10 | N | N | N |
| 625.60 | 20.54 | 7.12 | 4.74 | A | A | A |
| 601.2 | 10 | 2.95 | 3.81 | A | A | A |
| 547.52 | 27.4 | -6.25 | 6.06 | A | A | A |
| 590 | 85 | 1.03 | 14.81 | A | A | A |
| 622.4 | 54.88 | 6.58 | 9.46 | A | A | A |
| 593 | 31 | 1.54 | 6.25 | A | A | A |
| 563 | 31 | -3.60 | 6.48 | A | A | A |
| 599.2 | 31.5 | 2.60 | 6.27 | A | A | A |

TABLE 31. Evaluation Results for K-40 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

174

176

178

179

180

182

183

184

185

186

188

190

197

199

200

201

202

204

205

206

209

210

213

214

217

218

221

222

223

224

226

227

228

230

231

232

233

234

235

236

239

240

241

243

244

245

246

247

248

249

251

252

258

259

260

263

264

268

270

271

TABLE 31. Evaluation Results for K-40 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

| Reported value [Bq/kg ] 660 | Reported uncertainty (k = 1) [Bq/kg ] 30 | Relative bias [%] 13.01 | P-Test [%] 5.69 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 591 | 34 | 1.20 | 6.70 | A | A | A |
| 589.35 | 28.5 | 0.92 | 5.93 | A | A | A |
| 645.00 | 9.40 | 10.45 | 3.72 | A | N | W |
| 644 | 36.48 | 10.27 | 6.62 | A | A | A |
| 702.8 | 70.28 | 20.34 | 10.57 | A | A | A |
| 628 | 30 | 7.53 | 5.88 | A | A | A |
| 504 | 42 | -13.70 | 9.01 | A | A | A |
| 632 | 20 | 8.22 | 4.66 | A | A | A |
| 606.33 | 35.80 | 3.82 | 6.83 | A | A | A |
| 602 | 20 | 3.08 | 4.77 | A | A | A |
| 587.0 | 24.8 | 0.51 | 5.44 | A | A | A |
| 584 | 39 | 0.00 | 7.51 | A | A | A |
| 754.0 | 50.2 | 29.11 | 7.49 | A | N | W |
| 569 | 22 | -2.57 | 5.17 | A | A | A |
| 616.53 | 69.44 | 5.57 | 11.77 | A | A | A |
| 489 | 44 | -16.27 | 9.63 | A | A | A |
| 582 | 16 | -0.34 | 4.39 | A | A | A |
| 583.41 | 11.06 | -0.10 | 3.91 | A | A | A |
| 573.94 | 43.92 | -1.72 | 8.38 | A | A | A |
| 407.4 | 13.4 | -30.24 | 4.75 | N | N | N |
| 518.3 | 40.0 | -11.25 | 8.44 | A | A | A |
| 798 | 60 | 36.64 | 8.26 | N | N | N |
| 563.0 | 28.0 | -3.60 | 6.04 | A | A | A |
| 531 | 2.1 | -9.08 | 3.45 | A | N | W |
| 1158 | 45 | 98.29 | 5.18 | N | N | N |
| 855.28 | 0.15 | 46.45 | 3.42 | N | N | N |
| 616.9 | 25.5 | 5.63 | 5.37 | A | A | A |
| 587.3 | 39.6 | 0.57 | 7.56 | A | A | A |
| 885 | 88 | 51.54 | 10.52 | N | N | N |
| 548.1 | 11.4 | -6.15 | 4.01 | A | A | A |
| 760 | 40 | 30.14 | 6.28 | N | N | N |
| 380.93 | 18.47 | -34.77 | 5.94 | N | N | N |
| 70.9 | 4.02 | -87.86 | 6.62 | N | N | N |
| 626.9 | 23.2 | 7.35 | 5.04 | A | A | A |
| 526.22 | 48.98 | -9.89 | 9.92 | A | A | A |
| 551.545 |  | -5.56 |  | A |  |  |
| 698 | 12 | 19.52 | 3.83 | A | N | W |

Labcode

272

273

275

277

279

281

293

294

295

296

298

299

300

301

302

303

306

307

308

309

311

314

315

316

317

318

323

328

329

330

333

334

335

336

337

342

343

344

Reported Results for Cs-134 determination in Sample 4

TABLE 32. Target values

64

Activity concentration [ Bq/kg ]

FIGURE 30. Reported results and evaluation parameters for Cs-134. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

4	6	8	10

172

243

249

295

154

251

343

174

344

226

132

333

149

165

264

124

85

214

277

189

126

42

63

62

FIGURE 29. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

257

259

306

170

140

161

217

17

314

199

104

272

224

97

173

146

265

46

58

223

305

43

292

183

152

300

8

275

227

11

201

341

328

71

144

105

286

70

237

108

38

342

Parameter

Activity Concentration [ Bq/kg ] Uncertainty [ Bq/kg ]

Maximum Acceptable Relative Bias, [%]

137

23

197

153

50

Not Accepted 10%

49

6

180

54

260

2

247

283

64

171

125

158

303

329

282

178

81

120

102

Warning 3%

52

48

268

Accepted 57%

250

112

47

258

128

231

Laboratory code

244

77

18

327

232

107

302

33

221

185

90

147

210

193

60

155

75

298

5

166

10

156

27

200

179

228

163

151

86

130

299

222

167

225

25

67

271

29

31

Not Reported 30%

34

36

188

148

204

12

218

234

205

337

235

246

22

14

Value 6.93

0.50

30

72

182

270

168

51

134

133

206

296

209

28

281

53

30

240

248

19

99

279

89

157

190

141

162

55

84

186

122

330

317

181

245

96

118

145

334

13

39

69

83

123

135

66

1

233

202

79

307

87

301

73

110

309

26

230

294

213

68

239

57

335

74

93

184

131

80

241

4

273

318

236

| Reported value [Bq/kg ] 8.7 | Reported uncertainty (k = 1) [Bq/kg ] 0.6 | Relative bias [%] 25.54 | P-Test [%] 9.98 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 6.69 | 0.40 | -3.46 | 9.37 | A | A | A |
| 16.03 | 11.24 | 131.31 | 70.49 | N | N | N |
| 7.19 | 0.52 | 3.75 | 10.22 | A | A | A |
| 6.64 | 0.35 | -4.18 | 8.94 | A | A | A |
| 6.27 | 0.31 | -9.52 | 8.75 | A | A | A |
| 7.2 | 0.5 | 3.90 | 10.01 | A | A | A |
| 6.30 | 0.45 | -9.09 | 10.15 | A | A | A |
| 7.5 | 0.7 | 8.23 | 11.80 | A | A | A |
| 8.2 | 0.2 | 18.33 | 7.62 | A | A | A |
| 7.65 | 0.47 | 10.39 | 9.48 | A | A | A |
| 5.9 | 0.4 | -14.86 | 9.90 | A | A | A |
| 7.08 | 0.51 | 2.16 | 10.20 | A | A | A |
| 7.81 | 0.78 | 12.70 | 12.32 | A | A | A |
| 7.62 | 0.37 | 9.96 | 8.70 | A | A | A |
| 6.57 | 0.4 | -5.19 | 9.44 | A | A | A |
| 7.34 | 0.8 | 5.92 | 13.07 | A | A | A |
| 9.62 | 1.40 | 38.82 | 16.24 | N | A | N |
| 7.22 | 0.50 | 4.18 | 10.00 | A | A | A |
| 7.79 | 0.81 | 12.41 | 12.66 | A | A | A |
| 7.4 | 0.6 | 6.78 | 10.85 | A | A | A |
| 7.8 | 0.5 | 12.55 | 9.65 | A | A | A |
| 7.4 | 0.6 | 6.78 | 10.85 | A | A | A |
| 7.1 | 0.5 | 2.45 | 10.08 | A | A | A |
| 7.5 | 0.6 | 8.23 | 10.77 | A | A | A |
| 7.5 | 1.1 | 8.23 | 16.35 | A | A | A |
| 6.52 | 0.60 | -5.92 | 11.69 | A | A | A |
| 8.2 | 0.3 | 18.33 | 8.09 | A | A | A |
| 5.6 | 0.58 | -19.19 | 12.62 | A | A | A |
| 6.16 | 0.75 | -11.11 | 14.15 | A | A | A |
| 6.1 | 0.6 | -11.98 | 12.20 | A | A | A |
| 7.0 | 0.3 | 1.01 | 8.39 | A | A | A |
| 6.9 | 0.5 | -0.43 | 10.23 | A | A | A |
| 6.63 | 0.43 | -4.33 | 9.70 | A | A | A |
| 6.62 | 0.38 | -4.47 | 9.22 | A | A | A |
| 7.7 | 0.6 | 11.11 | 10.62 | A | A | A |
| 6.9 | 0.7 | -0.43 | 12.45 | A | A | A |
| 7.8 | 0.7 | 12.55 | 11.52 | A | A | A |
| 6.66 | 0.32 | -3.90 | 8.67 | A | A | A |
| 8.0 | 0.5 | 15.44 | 9.55 | A | A | A |
| 10.8 | 1.4 | 55.84 | 14.84 | N | N | N |
| 6.10 | 0.97 | -11.98 | 17.46 | A | A | A |
| 7.15 | 0.35 | 3.17 | 8.72 | A | A | A |
| 5.7 | 0.5 | -17.75 | 11.36 | A | A | A |
| 5.64 | 0.17 | -18.61 | 7.82 | A | A | A |
| 6.75 | 0.6 | -2.60 | 11.45 | A | A | A |
| 8.53 | 0.16 | 23.09 | 7.45 | A | N | W |
| 7.35 | 0.33 | 6.06 | 8.50 | A | A | A |
| 10.58 | 0.56 | 52.67 | 8.95 | N | N | N |
| 8.23 | 0.22 | 18.76 | 7.69 | A | A | A |
| 6.5 | 0.6 | -6.20 | 11.72 | A | A | A |
| 6.4 | 0.1 | -7.65 | 7.38 | A | A | A |
| 7.66 | 0.33 | 10.53 | 8.40 | A | A | A |
| 9.3 | 0.5 | 34.20 | 9.00 | N | N | N |
| 12.3 | 0.57 | 77.49 | 8.58 | N | N | N |
| 7.16 | 0.38 | 3.32 | 8.96 | A | A | A |
| 7.05 | 0.41 | 1.73 | 9.27 | A | A | A |
| 8.91 | 0.75 | 28.57 | 11.09 | A | A | A |
| 14.3 | 0.6 | 106.35 | 8.35 | N | N | N |
| 6.82 | 0.21 | -1.59 | 7.84 | A | A | A |

TABLE 33. Evaluation Results for Cs-134 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

1

2

4

5

6

8

10

11

12

13

14

17

18

19

22

23

25

26

27

28

29

30

31

33

34

36

38

39

42

43

46

47

48

49

50

51

52

53

54

55

57

58

60

62

63

64

66

67

68

69

70

71

72

73

74

75

77

79

80

81

| Reported value [Bq/kg ] 8.24 | Reported uncertainty (k = 1) [Bq/kg ] 0.41 | Relative bias [%] 18.90 | P-Test [%] 8.76 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 8 | 1.09 | 15.44 | 15.42 | A | A | A |
| 5.3 | 1.1 | -23.52 | 21.97 | A | A | A |
| 7.3 | 0.6 | 5.34 | 10.94 | A | A | A |
| 9.15 | 0.21 | 32.03 | 7.57 | N | N | N |
| 7.89 | 0.40 | 13.85 | 8.82 | A | A | A |
| 7.12 | 0.62 | 2.74 | 11.31 | A | A | A |
| 13 | 1 | 87.59 | 10.55 | N | N | N |
| 8.12 | 0.78 | 17.17 | 12.01 | A | A | A |
| 6.01 | 0.26 | -13.28 | 8.41 | A | A | A |
| 7.84 | 0.70 | 13.13 | 11.48 | A | A | A |
| 6.86 | 0.25 | -1.01 | 8.08 | A | A | A |
| 5.92 | 0.76 | -14.57 | 14.73 | A | A | A |
| 6.49 | 0.51 | -6.35 | 10.67 | A | A | A |
| 7.1 | 0.4 | 2.45 | 9.15 | A | A | A |
| 6.5 | 0.52 | -6.20 | 10.77 | A | A | A |
| 9.40 | 0.61 | 35.64 | 9.70 | N | N | N |
| 7.0 | 0.5 | 1.01 | 10.15 | A | A | A |
| 8.2 | 0.4 | 18.33 | 8.71 | A | A | A |
| 6.84 | 0.29 | -1.30 | 8.37 | A | A | A |
| 8.0 | 0.8 | 15.44 | 12.33 | A | A | A |
| 8.34 | 0.35 | 20.35 | 8.35 | A | A | A |
| 5.21 | 0.20 | -24.82 | 8.17 | A | N | W |
| 6.77 | 0.61 | -2.31 | 11.54 | A | A | A |
| 5.6 | 0.2 | -19.19 | 8.05 | A | A | A |
| 7.02 | 0.37 | 1.30 | 8.94 | A | A | A |
| 7.3 | 0.8 | 5.34 | 13.12 | A | A | A |
| 14.24 | 1.44 | 105.48 | 12.42 | N | N | N |
| 4.9 | 0.4 | -29.29 | 10.89 | A | N | W |
| 7.76 | 0.300 | 11.98 | 8.19 | A | A | A |
| 7.7 | 0.9 | 11.11 | 13.74 | A | A | A |
| 8.4 | 0.6 | 21.21 | 10.15 | A | A | A |
| 6.56 | 0.26 | -5.34 | 8.23 | A | A | A |
| 5.8 | 3.8 | -16.31 | 65.91 | A | N | W |
| 7.90 | 0.54 | 14.00 | 9.94 | A | A | A |
| 6.47 | 0.46 | -6.64 | 10.13 | A | A | A |
| 8.20 | 0.54 | 18.33 | 9.77 | A | A | A |
| 6.062 | 0.315 | -12.53 | 8.89 | A | A | A |
| 7.13 | 0.41 | 2.89 | 9.23 | A | A | A |
| 7.5 | 0.4 | 8.23 | 8.97 | A | A | A |
| 5.0443 | 0.44215 | -27.21 | 11.35 | A | A | A |
| 7.3 | 0.5 | 5.34 | 9.95 | A | A | A |
| 6.2 | 0.41 | -10.53 | 9.79 | A | A | A |
| 6.609 | 0.374 | -4.63 | 9.17 | A | A | A |
| 4.27 | 0.75 | -38.38 | 18.99 | N | A | N |
| 7.16 | 0.49 | 3.32 | 9.94 | A | A | A |
| 7.2 | 0.8 | 3.90 | 13.25 | A | A | A |
| 7.90 | 0.54 | 14.00 | 9.94 | A | A | A |
| 6.79 | 0.30 | -2.02 | 8.46 | A | A | A |
| 5.8 | 0.6 | -16.31 | 12.61 | A | A | A |
| 7.92 | 0.68 | 14.29 | 11.21 | A | A | A |
| 7.3 | 1.4 | 5.34 | 20.49 | A | A | A |
| 5.05 | 0.14 | -27.13 | 7.73 | A | N | W |
| 7.2 | 0.9 | 3.90 | 14.43 | A | A | A |
| 7.31 | 1.13 | 5.48 | 17.06 | A | A | A |
| 7.7 | 0.5 | 11.11 | 9.71 | A | A | A |
| 5.80 | 0.42 | -16.31 | 10.22 | A | A | A |
| 6.76 | 0.42 | -2.45 | 9.52 | A | A | A |
| 0.474 | 0.06 | -93.16 | 14.57 | N | N | N |
| 6.02 | 0.29 | -13.13 | 8.68 | A | A | A |

TABLE 33. Evaluation Results for Cs-134 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

83

84

85

86

87

89

90

93

96

97

99

102

104

105

107

108

110

112

118

120

122

123

124

125

126

128

130

131

132

133

134

135

137

140

141

144

145

146

147

148

149

151

152

153

154

155

156

157

158

161

162

163

165

166

167

168

170

171

172

173

| Reported value [Bq/kg ] 4.8 | Reported uncertainty (k = 1) [Bq/kg ] 0.5 | Relative bias [%] -30.74 | P-Test [%] 12.67 | Trueness evaluation N | Precision evaluation A | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 6.82 | 1.68 | -1.59 | 25.67 | A | A | A |
| 7.25 | 0.60 | 4.62 | 10.98 | A | A | A |
| 6.66 | 0.49 | -3.90 | 10.30 | A | A | A |
| 8.1 | 0.5 | 16.88 | 9.50 | A | A | A |
| 7.67 | 0.7 | 10.68 | 11.63 | A | A | A |
| 6.17 | 0.34 | -10.97 | 9.08 | A | A | A |
| 13.3 | 1.7 | 91.92 | 14.68 | N | N | N |
| 7.12 | 0.93 | 2.74 | 14.92 | A | A | A |
| 8.0 | 0.5 | 15.44 | 9.55 | A | A | A |
| 7.5 | 0.5 | 8.23 | 9.82 | A | A | A |
| 5.53 | 0.48 | -20.20 | 11.29 | A | A | A |
| 7.9 | 0.4 | 14.00 | 8.81 | A | A | A |
| 7.14 | 0.61 | 3.03 | 11.18 | A | A | A |
| 6.6003 | 0.4496 | -4.76 | 9.92 | A | A | A |
| 5.92 | 0.36 | -14.57 | 9.44 | A | A | A |
| 7.24 | 2.52 | 4.47 | 35.55 | A | N | W |
| 6.36 | 0.32 | -8.23 | 8.80 | A | A | A |
| 8.83 | 1.09 | 27.42 | 14.30 | A | A | A |
| 7.5 | 1.3 | 8.23 | 18.78 | A | A | A |
| 7.57 | 0.54 | 9.24 | 10.15 | A | A | A |
| 7.77 | 0.50 | 12.12 | 9.67 | A | A | A |
| 7.79 | 0.65 | 12.41 | 11.03 | A | A | A |
| 7.13 | 0.70 | 2.89 | 12.18 | A | A | A |
| 10.2 | 1.6 | 47.19 | 17.27 | N | N | N |
| 5.3 | 0.5 | -23.52 | 11.88 | A | A | A |
| 5.88 | 0.60 | -15.15 | 12.50 | A | A | A |
| 7.53 | 0.62 | 8.66 | 10.95 | A | A | A |
| 7.11 | 0.40 | 2.60 | 9.15 | A | A | A |
| 7.30 | 0.34 | 5.34 | 8.59 | A | A | A |
| 6.10 | 0.24 | -11.98 | 8.22 | A | A | A |
| 6 | 0.7 | -13.42 | 13.72 | A | A | A |
| 7.32 | 0.84 | 5.63 | 13.56 | A | A | A |
| 4.82 | 0.02 | -30.45 | 7.23 | N | N | N |
| 6.3 | 0.01 | -9.09 | 7.22 | A | A | A |
| 7.25 | 0.44 | 4.62 | 9.43 | A | A | A |
| 9.78 | 0.88 | 41.13 | 11.53 | N | N | N |
| 7.02 | 0.57 | 1.30 | 10.86 | A | A | A |
| 7.1 | 1.0 | 2.45 | 15.82 | A | A | A |
| 8.70 | 0.88 | 25.54 | 12.42 | A | A | A |
| 7.55 | 0.38 | 8.95 | 8.80 | A | A | A |
| 7.6 | 0.6 | 9.67 | 10.70 | A | A | A |
| 21.35 | 2.61 | 208.08 | 14.20 | N | N | N |
| 6.50 | 0.37 | -6.20 | 9.19 | A | A | A |
| 10.59 | 0.79 | 52.81 | 10.38 | N | N | N |
| 7.8 | 0.7 | 12.55 | 11.52 | A | A | A |
| 16.0 | 1.0 | 130.88 | 9.55 | N | N | N |
| 3.1 | 0.3 | -55.27 | 12.07 | N | N | N |
| 7.03 | 0.52 | 1.44 | 10.33 | A | A | A |
| 8.1 | 1.2 | 16.88 | 16.48 | A | A | A |
| 7.60 | 0.94 | 9.67 | 14.32 | A | A | A |
| 6.69 | 0.3 | -3.46 | 8.50 | A | A | A |
| 7.8 | 0.8 | 12.55 | 12.54 | A | A | A |
| 4.1 | 0.3 | -40.84 | 10.28 | N | N | N |
| 7 | 1.4 | 1.01 | 21.26 | A | A | A |
| 4.47 | 0.19 | -35.50 | 8.37 | N | N | N |
| 5.70 | 0.27 | -17.75 | 8.63 | A | A | A |
| 7.00 | 0.29 | 1.01 | 8.32 | A | A | A |
| 5.71 | 0.3 | -17.60 | 8.93 | A | A | A |
| 6.68 | 1.00 | -3.61 | 16.62 | A | A | A |

TABLE 33. Evaluation Results for Cs-134 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

174

178

179

180

181

182

183

184

185

186

188

189

190

193

197

199

200

201

202

204

205

206

209

210

213

214

217

218

221

222

223

224

225

226

227

228

230

231

232

233

234

235

236

237

239

240

241

243

244

245

246

247

248

249

250

251

257

258

259

260

TABLE 33. Evaluation Results for Cs-134 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

| Reported value [Bq/kg ] 5.078 | Reported uncertainty (k = 1) [Bq/kg ] 0.3933 | Relative bias [%] -26.72 | P-Test [%] 10.59 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 6.07 | 0.42 | -12.41 | 10.00 | A | A | A |
| 6.91 | 0.68 | -0.29 | 12.20 | A | A | A |
| 7.69 | 0.54 | 10.97 | 10.07 | A | A | A |
| 7.4 | 0.4 | 6.78 | 9.02 | A | A | A |
| 5.94 | 0.5 | -14.29 | 11.09 | A | A | A |
| 16.4 | 1.0 | 136.65 | 9.45 | N | N | N |
| 6.28 | 0.7 | -9.38 | 13.28 | A | A | A |
| 5.36 | 0.62 | -22.66 | 13.63 | A | A | A |
| 7.85 | 0.472 | 13.28 | 9.39 | A | A | A |
| 7.792 | 0.857 | 12.44 | 13.15 | A | A | A |
| 6.8 | 0.28 | -1.88 | 8.31 | A | A | A |
| 6.692 | 3.414 | -3.43 | 51.52 | A | N | W |
| 6.49 | 1.2 | -6.35 | 19.85 | A | A | A |
| 6.17 | 0.45 | -10.97 | 10.26 | A | A | A |
| 10.2 | 0.8 | 47.19 | 10.66 | N | N | N |
| 4.2 | 0.2 | -39.39 | 8.64 | N | N | N |
| 7.77 | 0.62 | 12.12 | 10.76 | A | A | A |
| 7.17 | 0.43 | 3.46 | 9.38 | A | A | A |
| 7.3 | 0.5 | 5.34 | 9.95 | A | A | A |
| 6.2 | 0.5 | -10.53 | 10.82 | A | A | A |
| 9.237 | 0.705 | 33.29 | 10.50 | N | N | N |
| 7.1 | 0.3 | 2.45 | 8.36 | A | A | A |
| 6.79 | 1.48 | -2.02 | 22.96 | A | A | A |
| 6.1 | 0.6 | -11.98 | 12.20 | A | A | A |
| 5.76 | 0.41 | -16.88 | 10.14 | A | A | A |
| 9.0 | 0.4 | 29.87 | 8.47 | A | N | W |
| 9.43 | 1.03 | 36.08 | 13.09 | N | N | N |
| 5.9 | 1.0 | -14.86 | 18.42 | A | A | A |
| 8.02 | 1.05 | 15.73 | 14.95 | A | A | A |
| 19.8 | 2.5 | 185.71 | 14.54 | N | N | N |
| 7.09 | 0.32 | 2.31 | 8.51 | A | A | A |
| 6.38 | 0.43 | -7.94 | 9.87 | A | A | A |
| 6.8 | 0.5 | -1.88 | 10.30 | A | A | A |
| 8.0 | 0.8 | 15.44 | 12.33 | A | A | A |
| 5.0 | 0.2 | -27.85 | 8.25 | A | N | W |
| 8.2 | 0.4 | 18.33 | 8.71 | A | A | A |
| 11.62 | 1.25 | 67.68 | 12.95 | N | N | N |
| 7.6 | 0.8 | 9.67 | 12.76 | A | A | A |
| 6.36 | 0.62 | -8.23 | 12.13 | A | A | A |
| 6.56 | 0.26 | -5.34 | 8.23 | A | A | A |
| 4.55803 |  | -34.23 |  | N |  | N |
| 4.8 | 1.3 | -30.74 | 28.03 | N | A | N |

Labcode

264

265

268

270

271

272

273

275

277

279

281

282

283

286

292

294

295

296

298

299

300

301

302

303

305

306

307

309

314

317

318

327

328

329

330

333

334

335

337

341

342

343

344

Activity concentration  [ Bq/kg ]

250

300

Reported Results for Cs-137 determination in Sample 4

TABLE 34. Target values

| Parameter | Value |
| --- | --- |
| Activity Concentration [ Bq/kg ] | 270 |
| Uncertainty [ Bq/kg ] | 17 |
| Maximum Acceptable Relative Bias, [%] | 30 |

Accepted 66%

Not Reported 22%

Not Accepted 7%

Warning 5%

350

400

FIGURE 31. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

100

150

200

336

140

249

159

223

15

146

217

239

124

204

112

333

128

292

173

62

183

31

167

126

47

305

105

200

231

275

90

125

299

58

11

81

2

237

79

180

189

64

66

50

199

8

182

168

18

198

51

158

67

97

321

39

25

132

77

104

302

163

263

258

270

74

230

155

222

251

107

69

27

10

48

102

71

134

55

342

1

28

232

244

19

210

141

259

243

245

294

279

148

250

54

272

36

344

17

186

193

26

296

135

142

123

162

273

293

317

7

144

157

118

341

334

214

73

99

96

91

110

154

4

241

318

340

Laboratory code

FIGURE 32. Reported results and evaluation parameters for Cs-137. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

| Reported value [Bq/kg ] 288.2 | Reported uncertainty (k = 1) [Bq/kg ] 15.9 | Relative bias [%] 6.74 | P-Test [%] 8.37 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 259.27 | 9.54 | -3.97 | 7.29 | A | A | A |
| 401.99 | 10.19 | 48.89 | 6.79 | N | N | N |
| 316.37 | 17.05 | 17.17 | 8.29 | A | A | A |
| 250 | 16.8 | -7.41 | 9.21 | A | A | A |
| 319.4 | 19.6 | 18.30 | 8.79 | A | A | A |
| 263.37 | 15.9 | -2.46 | 8.72 | A | A | A |
| 285 | 7 | 5.56 | 6.76 | A | A | A |
| 258 | 13 | -4.44 | 8.06 | A | A | A |
| 280 | 10 | 3.70 | 7.24 | A | A | A |
| 363 | 10.5 | 34.44 | 6.93 | N | N | N |
| 297.1 | 16.9 | 10.04 | 8.49 | A | A | A |
| 222.2 | 7.3 | -17.70 | 7.10 | A | A | A |
| 304.2 | 21.1 | 12.67 | 9.37 | A | A | A |
| 265.66 | 18.21 | -1.61 | 9.31 | A | A | A |
| 293 | 19 | 8.52 | 9.04 | A | A | A |
| 263.3 | 7.9 | -2.48 | 6.97 | A | A | A |
| 268 | 9 | -0.74 | 7.14 | A | A | A |
| 270.2 | 27.0 | 0.07 | 11.81 | A | A | A |
| 310 | 19.5 | 14.81 | 8.90 | A | A | A |
| 283.9 | 7.5 | 5.15 | 6.83 | A | A | A |
| 289 | 24.9 | 7.04 | 10.67 | A | A | A |
| 275 | 30 | 1.85 | 12.60 | A | A | A |
| 298 | 12 | 10.37 | 7.47 | A | A | A |
| 247 | 15 | -8.52 | 8.75 | A | A | A |
| 269.1 | 18.0 | -0.33 | 9.19 | A | A | A |
| 262 | 21.8 | -2.96 | 10.43 | A | A | A |
| 302 | 25 | 11.85 | 10.40 | A | A | A |
| 253 | 23 | -6.30 | 11.06 | A | A | A |
| 270 | 10 | 0.00 | 7.30 | A | A | A |
| 237.5 | 22.57 | -12.04 | 11.40 | A | A | A |
| 257 | 24 | -4.81 | 11.26 | A | A | A |
| 251 | 8 | -7.04 | 7.06 | A | A | A |
| 250 | 8 | -7.41 | 7.06 | A | A | A |
| 285 | 22 | 5.56 | 9.96 | A | A | A |
| 262.8 | 3.16 | -2.67 | 6.41 | A | A | A |
| 262.93 | 3.01 | -2.62 | 6.40 | A | A | A |
| 267.3 | 14.5 | -1.00 | 8.31 | A | A | A |
| 282 | 13 | 4.44 | 7.80 | A | A | A |
| 287.7 | 8.8 | 6.56 | 7.00 | A | A | A |
| 300.2 | 2.5 | 11.19 | 6.35 | A | A | A |
| 287.5 | 14.7 | 6.48 | 8.11 | A | A | A |
| 315 | 19 | 16.67 | 8.72 | A | A | A |
| 258.00 | 41.28 | -4.44 | 17.19 | A | A | A |
| 265.22 | 11.83 | -1.77 | 7.72 | A | A | A |
| 246 | 10.2 | -8.89 | 7.54 | A | A | A |
| 258 | 6 | -4.44 | 6.71 | A | A | A |
| 261.8 | 17.8 | -3.04 | 9.27 | A | A | A |
| 262.14 | 4.9 | -2.91 | 6.57 | A | A | A |
| 269.03 | 7.06 | -0.36 | 6.82 | A | A | A |
| 466 | 18 | 72.59 | 7.39 | N | N | N |
| 282.40 | 5.66 | 4.59 | 6.61 | A | A | A |
| 305.2 | 6.4 | 13.04 | 6.64 | A | A | A |
| 286 | 4 | 5.93 | 6.45 | A | A | A |
| 315.5 | 2.96 | 16.85 | 6.37 | A | N | W |
| 347 | 14 | 28.52 | 7.48 | A | N | W |
| 277.6 | 5.38 | 2.81 | 6.59 | A | A | A |
| 274.81 | 16.55 | 1.78 | 8.71 | A | A | A |
| 200 | 14 | -25.93 | 9.42 | A | N | W |
| 273 | 12 | 1.11 | 7.68 | A | A | A |

TABLE 35. Evaluation Results for Cs-137 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

1

2

4

5

6

7

8

10

11

12

13

14

15

17

18

19

21

23

25

26

27

28

29

30

31

33

34

36

38

39

42

43

46

47

48

49

50

51

52

53

54

55

57

58

60

62

63

64

66

67

68

69

70

71

72

73

74

75

76

77

| Reported value [Bq/kg ] 260 | Reported uncertainty (k = 1) [Bq/kg ] 11 | Relative bias [%] -3.70 | P-Test [%] 7.59 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 300.4 | 3.6 | 11.26 | 6.41 | A | A | A |
| 259 | 7 | -4.07 | 6.85 | A | A | A |
| 323.72 | 11.57 | 19.90 | 7.24 | A | N | W |
| 347 | 30.0 | 28.52 | 10.70 | A | N | W |
| 208 | 42 | -22.96 | 21.15 | A | A | A |
| 293 | 21 | 8.52 | 9.54 | A | A | A |
| 375.16 | 1.08 | 38.95 | 6.30 | N | N | N |
| 384.0 | 15.0 | 42.22 | 7.41 | N | N | N |
| 253.30 | 9.51 | -6.19 | 7.33 | A | A | A |
| 368.71 | 1.73 | 36.56 | 6.31 | N | N | N |
| 608 | 13 | 125.19 | 6.65 | N | N | N |
| 358.98 | 2.15 | 32.96 | 6.32 | N | N | N |
| 269.32 | 1.76 | -0.25 | 6.33 | A | A | A |
| 349.2 | 8.0 | 29.33 | 6.70 | A | N | W |
| 285.7 | 10.7 | 5.81 | 7.33 | A | A | A |
| 273.29 | 14.62 | 1.22 | 8.26 | A | A | A |
| 251.7 | 13.4 | -6.78 | 8.25 | A | A | A |
| 282 | 8 | 4.44 | 6.91 | A | A | A |
| 328 | 4 | 21.48 | 6.41 | A | N | W |
| 376.33 | 5.93 | 39.38 | 6.49 | N | N | N |
| 240 | 5 | -11.11 | 6.63 | A | A | A |
| 329.1 | 19.8 | 21.89 | 8.71 | A | A | A |
| 272 | 4.4 | 0.74 | 6.50 | A | A | A |
| 264.8 | 8.42 | -1.93 | 7.05 | A | A | A |
| 282 | 10 | 4.44 | 7.23 | A | A | A |
| 315 | 13 | 16.67 | 7.53 | A | A | A |
| 235.23 | 2.28 | -12.88 | 6.37 | A | A | A |
| 255.32 | 26.38 | -5.44 | 12.10 | A | A | A |
| 250 | 5 | -7.41 | 6.61 | A | A | A |
| 243.32 | 4.16 | -9.88 | 6.52 | A | A | A |
| 214.89 | 19.00 | -20.41 | 10.85 | A | A | A |
| 285 | 10 | 5.56 | 7.21 | A | A | A |
| 291.18 | 10.87 | 7.84 | 7.32 | A | A | A |
| 271.7 | 24.5 | 0.63 | 11.00 | A | A | A |
| 270 | 14.9 | 0.00 | 8.37 | A | A | A |
| 286 | 7 | 5.93 | 6.76 | A | A | A |
| 311 | 22 | 15.19 | 9.47 | A | A | A |
| 248.37 | 13.33 | -8.01 | 8.27 | A | A | A |
| 26.3 | 7.6 | -90.26 | 29.58 | N | N | N |
| 295.1 | 17.7 | 9.30 | 8.70 | A | A | A |
| 314.421 | 10.222 | 16.45 | 7.09 | A | A | A |
| 322.7 | 1.4 | 19.52 | 6.31 | A | N | W |
| 234 | 14 | -13.33 | 8.69 | A | A | A |
| 227.7 | 9.1 | -15.67 | 7.46 | A | A | A |
| 308.86 | 16.22 | 14.39 | 8.20 | A | A | A |
| 299.4 | 8.7 | 10.89 | 6.93 | A | A | A |
| 294.27 | 14.2515 | 8.99 | 7.94 | A | A | A |
| 287.2 | 15.2 | 6.37 | 8.23 | A | A | A |
| 259.3 | 8.30 | -3.96 | 7.06 | A | A | A |
| 269.330 | 16.193 | -0.25 | 8.71 | A | A | A |
| 391.4 | 9 | 44.96 | 6.70 | N | N | N |
| 280 | 19 | 3.70 | 9.26 | A | A | A |
| 289 | 12 | 7.04 | 7.54 | A | A | A |
| 326 | 10 | 20.74 | 7.00 | A | N | W |
| 268.65 | 10.69 | -0.50 | 7.45 | A | A | A |
| 203.8 | 7.0 | -24.52 | 7.17 | A | N | W |
| 233 | 22 | -13.70 | 11.35 | A | A | A |
| 315 | 18 | 16.67 | 8.50 | A | A | A |
| 274.9 | 11.5 | 1.81 | 7.56 | A | A | A |

TABLE 35. Evaluation Results for Cs-137 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

79

80

81

83

84

85

86

87

89

90

91

93

96

97

99

102

104

105

107

108

110

112

118

119

120

122

123

124

125

126

128

129

130

131

132

133

134

135

137

140

141

142

144

145

146

147

148

149

151

152

153

154

155

156

157

158

159

161

162

163

| Reported value [Bq/kg ] 310.10 | Reported uncertainty (k = 1) [Bq/kg ] 10.65 | Relative bias [%] 14.85 | P-Test [%] 7.17 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 278 | 36 | 2.96 | 14.40 | A | A | A |
| 248.00 | 9.60 | -8.15 | 7.39 | A | A | A |
| 265 | 13 | -1.85 | 7.98 | A | A | A |
| 250 | 21 | -7.41 | 10.50 | A | A | A |
| 227 | 14 | -15.93 | 8.81 | A | A | A |
| 269 | 8.0 | -0.37 | 6.96 | A | A | A |
| 21.1 | 1.30 | -92.19 | 8.81 | N | N | N |
| 245.1 | 12.32 | -9.22 | 8.06 | A | A | A |
| 290 | 18 | 7.41 | 8.84 | A | A | A |
| 346 | 19 | 28.15 | 8.35 | A | N | W |
| 260.2 | 42.4 | -3.63 | 17.47 | A | A | A |
| 315 | 11 | 16.67 | 7.20 | A | A | A |
| 260.43 | 15.85 | -3.54 | 8.76 | A | A | A |
| 285.6 | 14.3 | 5.78 | 8.04 | A | A | A |
| 264.2 | 21.2 | -2.15 | 10.20 | A | A | A |
| 246 | 12 | -8.89 | 7.96 | A | A | A |
| 300 | 30 | 11.11 | 11.82 | A | A | A |
| 278.10 | 36.15 | 3.00 | 14.44 | A | A | A |
| 306 | 4 | 13.33 | 6.43 | A | A | A |
| 300 | 20 | 11.11 | 9.17 | A | A | A |
| 261 | 21 | -3.33 | 10.22 | A | A | A |
| 295.7 | 11.1 | 9.52 | 7.33 | A | A | A |
| 307.37 | 6.62 | 13.84 | 6.65 | A | A | A |
| 276.2339 | 2.2879 | 2.31 | 6.35 | A | A | A |
| 266.1 | 4.4 | -1.44 | 6.51 | A | A | A |
| 263.01 | 2.18 | -2.59 | 6.35 | A | A | A |
| 252.0 | 15.8 | -6.67 | 8.89 | A | A | A |
| 247.9 | 14.9 | -8.19 | 8.70 | A | A | A |
| 314.61 | 15.76 | 16.52 | 8.05 | A | A | A |
| 236 | 13 | -12.59 | 8.37 | A | A | A |
| 296.38 | 5.71 | 9.77 | 6.58 | A | A | A |
| 332.10 | 9.12 | 23.00 | 6.87 | A | N | W |
| 262.97 | 22.01 | -2.60 | 10.47 | A | A | A |
| 294.04 | 17.01 | 8.90 | 8.55 | A | A | A |
| 355 | 28 | 31.48 | 10.09 | N | N | N |
| 346 | 5 | 28.15 | 6.46 | A | N | W |
| 229.6 | 14.1 | -14.96 | 8.80 | A | A | A |
| 299 | 13 | 10.74 | 7.65 | A | A | A |
| 252.70 | 1.30 | -6.41 | 6.32 | A | A | A |
| 281.00 | 16.90 | 4.07 | 8.71 | A | A | A |
| 210.00 | 12.70 | -22.22 | 8.73 | A | A | A |
| 284.6 | 2.9 | 5.41 | 6.38 | A | A | A |
| 252 | 12 | -6.67 | 7.89 | A | A | A |
| 270.06 | 0.7 | 0.02 | 6.30 | A | A | A |
| 245.9 | 0.2 | -8.93 | 6.30 | A | A | A |
| 274 | 18 | 1.48 | 9.10 | A | A | A |
| 278.01 | 5.21 | 2.97 | 6.57 | A | A | A |
| 252.33 | 19.57 | -6.54 | 9.99 | A | A | A |
| 291 | 12 | 7.78 | 7.53 | A | A | A |
| 258.1 | 7.22 | -4.41 | 6.89 | A | A | A |
| 297.67 | 13.39 | 10.25 | 7.74 | A | A | A |
| 281.6 | 14.4 | 4.30 | 8.11 | A | A | A |
| 1075.75 | 24.5 | 298.43 | 6.70 | N | N | N |
| 260 | 4.5 | -3.70 | 6.53 | A | A | A |
| 233.47 | 5.14 | -13.53 | 6.67 | A | A | A |
| 311 | 16 | 15.19 | 8.13 | A | A | A |
| 515 | 37.3 | 90.74 | 9.60 | N | N | N |
| 297 | 26 | 10.00 | 10.78 | A | A | A |
| 292 | 21 | 8.15 | 9.56 | A | A | A |

TABLE 35. Evaluation Results for Cs-137 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

165

166

167

168

169

170

171

172

173

174

176

178

179

180

181

182

183

184

185

186

188

189

190

193

197

198

199

200

201

202

204

205

206

209

210

213

214

217

218

221

222

223

224

225

226

227

228

230

231

232

233

234

235

236

237

239

240

241

243

244

| Reported value [Bq/kg ] 297.5 | Reported uncertainty (k = 1) [Bq/kg ] 16.7 | Relative bias [%] 10.19 | P-Test [%] 8.44 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 304 | 21 | 12.59 | 9.35 | A | A | A |
| 286.0 | 3.8 | 5.93 | 6.43 | A | A | A |
| 271.1 | 15.2 | 0.41 | 8.43 | A | A | A |
| 142 | 9.1 | -47.41 | 8.98 | N | N | N |
| 300 | 60 | 11.11 | 20.97 | A | A | A |
| 281.93 | 7.05 | 4.42 | 6.77 | A | A | A |
| 395 | 22 | 46.30 | 8.41 | N | N | N |
| 228 | 9 | -15.56 | 7.43 | A | A | A |
| 276.20 | 8.89 | 2.30 | 7.07 | A | A | A |
| 296.3 | 9.4 | 9.74 | 7.05 | A | A | A |
| 277.11 | 22.17 | 2.63 | 10.18 | A | A | A |
| 276 | 39 | 2.22 | 15.47 | A | A | A |
| 236 | 21.57 | -12.59 | 11.10 | A | A | A |
| 243.9 | 1.8 | -9.67 | 6.34 | A | A | A |
| 273 | 14 | 1.11 | 8.12 | A | A | A |
| 277 | 14 | 2.59 | 8.07 | A | A | A |
| 261.7 | 13.1 | -3.07 | 8.04 | A | A | A |
| 301 | 15 | 11.48 | 8.03 | A | A | A |
| 315 | 16 | 16.67 | 8.09 | A | A | A |
| 252.9 | 15.4 | -6.33 | 8.76 | A | A | A |
| 246.00 | 1.24 | -8.89 | 6.32 | A | A | A |
| 299 | 20.8 | 10.74 | 9.38 | A | A | A |
| 302.4 | 30.24 | 12.00 | 11.82 | A | A | A |
| 312 | 4.4 | 15.56 | 6.45 | A | A | A |
| 319.6 | 57.43 | 18.37 | 19.04 | A | A | A |
| 255.4 | 15.6 | -5.41 | 8.77 | A | A | A |
| 252 | 15.37 | -6.67 | 8.77 | A | A | A |
| 243.97 | 15.74 | -9.64 | 9.01 | A | A | A |
| 316 | 14 | 17.04 | 7.70 | A | A | A |
| 298 | 17 | 10.37 | 8.50 | A | A | A |
| 247 | 6.4 | -8.52 | 6.81 | A | A | A |
| 310.32 | 8.35 | 14.93 | 6.85 | A | A | A |
| 259 | 8 | -4.07 | 7.01 | A | A | A |
| 256.1 | 15.4 | -5.15 | 8.71 | A | A | A |
| 242.3 | 19 | -10.26 | 10.06 | A | A | A |
| 332.6 | 21.6 | 23.19 | 9.05 | A | A | A |
| 274 | 5 | 1.48 | 6.56 | A | A | A |
| 285.75 | 35.19 | 5.83 | 13.83 | A | A | A |
| 260 | 27 | -3.70 | 12.14 | A | A | A |
| 250 | 10 | -7.41 | 7.46 | A | A | A |
| 276 | 13 | 2.22 | 7.86 | A | A | A |
| 266 | 2.0 | -1.48 | 6.34 | A | A | A |
| 260.55 | 1.55 | -3.50 | 6.32 | A | A | A |
| 267.27 | 13.18 | -1.01 | 8.00 | A | A | A |
| 241.8 | 6.2 | -10.44 | 6.80 | A | A | A |
| 287.6 | 21.0 | 6.52 | 9.64 | A | A | A |
| 307 | 46 | 13.70 | 16.25 | A | A | A |
| 245.0 | 4.3 | -9.26 | 6.54 | A | A | A |
| 316.46 | 2.14 | 17.21 | 6.33 | A | N | W |
| 672 | 60 | 148.89 | 10.93 | N | N | N |
| 269.37 | 9.5 | -0.23 | 7.22 | A | A | A |
| 37.78 | 2.91 | -86.01 | 9.95 | N | N | N |
| 301.9 | 1.9 | 11.81 | 6.33 | A | A | A |
| 264.0 | 10.4 | -2.22 | 7.43 | A | A | A |
| 283.6 | 18.8 | 5.04 | 9.14 | A | A | A |
| 349 | 34 | 29.26 | 11.60 | A | A | A |
| 242.2 | 4.6 | -10.30 | 6.58 | A | A | A |
| 334 | 16 | 23.70 | 7.91 | A | N | W |
| 318.92 | 23.48 | 18.12 | 9.69 | A | A | A |

TABLE 35. Evaluation Results for Cs-137 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

245

246

247

248

249

250

251

252

257

258

259

260

263

264

265

268

270

271

272

273

275

277

279

281

282

283

286

291

292

293

294

295

296

298

299

300

301

302

303

304

305

306

307

308

309

311

314

315

316

317

318

321

322

327

328

329

330

333

334

335

TABLE 35. Evaluation Results for Cs-137 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

| Reported value [Bq/kg ] 4.62 | Reported uncertainty (k = 1) [Bq/kg ] 0.27 | Relative bias [%] -98.29 | P-Test [%] 8.59 | Trueness evaluation N | Precision evaluation N | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 292.4 | 8.2 | 8.30 | 6.89 | A | A | A |
| 3041 | 513 | 1026.30 | 18.01 | N | N | N |
| 332.31 | 19.43 | 23.08 | 8.59 | A | N | W |
| 287.66 | 14.79 | 6.54 | 8.13 | A | A | A |
| 253.942 |  | -5.95 |  | A |  |  |
| 303 | 7.8 | 12.22 | 6.80 | A | A | A |

Labcode

336

337

340

341

342

343

344

Activity concentration  [ Bq/kg ]

12

14

Reported Results for Tl-208 determination in Sample 4

TABLE 36. Target values

| Parameter | Value |
| --- | --- |
| Activity Concentration [ Bq/kg ] | 11.8 |
| Uncertainty [ Bq/kg ] | 0.8 |
| Maximum Acceptable Relative Bias, [%] | 30 |

Not Accepted 8%

Warning 0%

Accepted 25%

Not Reported 67%

16

18

FIGURE 33. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

6

8

10

222

112

333

92

223

244

204

271

183

166

221

185

132

102

268

55

299

25

298

14

34

170

107

155

71

176

12

33

182

245

134

260

90

180

206

277

23

5

199

17

10

54

146

126

174

158

235

130

281

30

2

46

119

248

165

144

156

122

335

141

258

142

228

28

270

69

81

151

22

302

31

218

53

149

217

190

279

118

240

307

72

57

99

87

89

301

323

294

96

230

27

63

58

305

75

173

42

153

140

315

110

243

68

Laboratory code

FIGURE 34. Reported results and evaluation parameters for Tl-208. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

| Reported value [Bq/kg ] 12.03 | Reported uncertainty (k = 1) [Bq/kg ] 0.68 | Relative bias [%] 1.95 | P-Test [%] 8.83 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 11.63 | 0.67 | -1.44 | 8.90 | A | A | A |
| 11.7 | 0.9 | -0.85 | 10.25 | A | A | A |
| 11.2 | 1.0 | -5.08 | 11.21 | A | A | A |
| 10.95 | 0.66 | -7.20 | 9.07 | A | A | A |
| 11.7 | 1.2 | -0.85 | 12.29 | A | A | A |
| 12.74 | 0.66 | 7.97 | 8.53 | A | A | A |
| 11.6 | 0.6 | -1.69 | 8.53 | A | A | A |
| 10.88 | 1.1 | -7.80 | 12.17 | A | A | A |
| 31.3 | 1.4 | 165.25 | 8.12 | N | N | N |
| 12.6 | 1.7 | 6.78 | 15.10 | A | A | A |
| 12.0 | 0.8 | 1.69 | 9.51 | A | A | A |
| 13.2 | 2.6 | 11.86 | 20.83 | A | A | A |
| 11.2 | 0.8 | -5.08 | 9.85 | A | A | A |
| 11 | 1.2 | -6.78 | 12.84 | A | A | A |
| 35.1 | 3.33 | 197.46 | 11.66 | N | N | N |
| 12.1 | 0.4 | 2.54 | 7.54 | A | A | A |
| 13.3 | 0.7 | 12.71 | 8.58 | A | A | A |
| 11.7 | 0.4 | -0.85 | 7.59 | A | A | A |
| 10.8 | 0.7 | -8.47 | 9.38 | A | A | A |
| 14.9 | 1.6 | 26.27 | 12.70 | A | A | A |
| 32.0 | 4.8 | 171.19 | 16.46 | N | N | N |
| 31.5 | 0.9 | 166.95 | 7.36 | N | N | N |
| 49.7 | 2.3 | 321.19 | 8.21 | N | N | N |
| 12.60 | 0.54 | 6.78 | 8.02 | A | A | A |
| 11.1 | 0.2 | -5.93 | 7.02 | A | A | A |
| 14.33 | 0.79 | 21.44 | 8.74 | A | A | A |
| 32.53 | 1.95 | 175.68 | 9.05 | N | N | N |
| 12.7 | 0.4 | 7.63 | 7.48 | A | A | A |
| 15.41 | 0.46 | 30.59 | 7.41 | N | N | N |
| 17.53 | 1.33 | 48.56 | 10.17 | N | N | N |
| 11.31 | 1.23 | -4.15 | 12.82 | A | A | A |
| 8.14 | 1.23 | -31.02 | 16.56 | N | A | N |
| 24.68 | 1.81 | 109.15 | 9.99 | N | N | N |
| 15.2 | 1.6 | 28.81 | 12.52 | A | A | A |
| 10.7 | 0.5 | -9.32 | 8.23 | A | A | A |
| 11.0 | 0.8 | -6.78 | 9.94 | A | A | A |
| 45.91 | 0.76 | 289.07 | 6.98 | N | N | N |
| 6.9 | 0.4 | -41.53 | 8.92 | N | N | N |
| 13.7 | 0.9 | 16.10 | 9.44 | A | A | A |
| 12.1 | 1.2 | 2.54 | 12.01 | A | A | A |
| 12.3 | 0.9 | 4.24 | 9.98 | A | A | A |
| 11.8 | 0.5 | 0.00 | 7.99 | A | A | A |
| 11.9 | 1.0 | 0.85 | 10.80 | A | A | A |
| 10.7 | 1.0 | -9.32 | 11.55 | A | A | A |
| 11.3 | 0.8 | -4.24 | 9.80 | A | A | A |
| 39.7 | 2.4 | 236.44 | 9.08 | N | N | N |
| 12.33 | 0.80 | 4.49 | 9.38 | A | A | A |
| 12.568 | 1.020 | 6.51 | 10.58 | A | A | A |
| 12.26 | 0.67 | 3.90 | 8.71 | A | A | A |
| 11.77 | 0.60 | -0.25 | 8.48 | A | A | A |
| 13.376 | 0.8923 | 13.36 | 9.51 | A | A | A |
| 12.7 | 0.8 | 7.63 | 9.25 | A | A | A |
| 35.282 | 2.145 | 199.00 | 9.11 | N | N | N |
| 11.02 | 0.7 | -6.61 | 9.29 | A | A | A |
| 12.3 | 1.2 | 4.24 | 11.88 | A | A | A |
| 11.88 | 0.51 | 0.68 | 8.02 | A | A | A |
| 12.22 | 0.29 | 3.56 | 7.18 | A | A | A |
| 10.5 | 1.3 | -11.02 | 14.12 | A | A | A |
| 11.0 | 0.7 | -6.78 | 9.30 | A | A | A |

TABLE 37. Evaluation Results for Tl-208 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

2

5

10

12

14

17

22

23

25

27

28

30

31

33

34

42

46

53

54

55

57

58

63

68

69

71

72

75

81

87

89

90

92

96

99

102

107

110

112

118

119

122

126

130

132

134

140

141

142

144

146

149

151

153

155

156

158

165

166

170

TABLE 37. Evaluation Results for Tl-208 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

| Reported value [Bq/kg ] 33.65 | Reported uncertainty (k = 1) [Bq/kg ] 1.34 | Relative bias [%] 185.17 | P-Test [%] 7.86 | Trueness evaluation N | Precision evaluation N | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 11.8 | 0.7 | 0.00 | 9.01 | A | A | A |
| 11.2 | 0.8 | -5.08 | 9.85 | A | A | A |
| 11.35 | 0.81 | -3.81 | 9.84 | A | A | A |
| 11.28 | 1 | -4.41 | 11.16 | A | A | A |
| 10.4 | 0.7 | -11.86 | 9.55 | A | A | A |
| 10.57 | 1.37 | -10.42 | 14.63 | A | A | A |
| 13.6 | 0.9 | 15.25 | 9.47 | A | A | A |
| 11.67 | 0.47 | -1.10 | 7.89 | A | A | A |
| 10.0 | 2.0 | -15.25 | 21.12 | A | A | A |
| 11.44 | 0.59 | -3.05 | 8.52 | A | A | A |
| 13.6 | 1.8 | 15.25 | 14.87 | A | A | A |
| 13.2 | 0.87 | 11.86 | 9.46 | A | A | A |
| 10.51 | 0.31 | -10.93 | 7.39 | A | A | A |
| 2.57 | 0.16 | -78.22 | 9.20 | N | N | N |
| 9.56 | 0.47 | -18.98 | 8.37 | A | A | A |
| 12.6 | 0.9 | 6.78 | 9.85 | A | A | A |
| 29.1 | 1.2 | 146.61 | 7.94 | N | N | N |
| 11.9 | 0.9 | 0.85 | 10.16 | A | A | A |
| 13.9 | 1.3 | 17.80 | 11.55 | A | A | A |
| 49 | 4.9 | 315.25 | 12.08 | N | N | N |
| 9.8 | 0.74 | -16.95 | 10.15 | A | A | A |
| 11.3 | 2.0 | -4.24 | 18.95 | A | A | A |
| 12.2 | 1.1 | 3.39 | 11.28 | A | A | A |
| 12.50 | 0.59 | 5.93 | 8.26 | A | A | A |
| 11.31 | 1.13 | -4.15 | 12.07 | A | A | A |
| 10.7 | 0.73 | -9.32 | 9.62 | A | A | A |
| 12.6 | 0.78 | 6.78 | 9.18 | A | A | A |
| 10.3 | 0.7 | -12.71 | 9.60 | A | A | A |
| 11.60 | 0.47 | -1.69 | 7.90 | A | A | A |
| 13.66 | 1.28 | 15.76 | 11.57 | A | A | A |
| 11.93 | 1.312 | 1.10 | 12.92 | A | A | A |
| 24.5 | 2.8 | 107.63 | 13.29 | N | N | N |
| 10.9 | 0.5 | -7.63 | 8.19 | A | A | A |
| 10.8 | 0.7 | -8.47 | 9.38 | A | A | A |
| 20.61 | 1.42 | 74.66 | 9.67 | N | N | N |
| 12.9 | 0.4 | 9.32 | 7.46 | A | A | A |
| 32 | 2 | 171.19 | 9.22 | N | N | N |
| 13.9 | 1.8 | 17.80 | 14.62 | A | A | A |
| 42.5 | 4.1 | 260.17 | 11.79 | N | N | N |
| 20.68 | 0.01 | 75.25 | 6.78 | N | N | N |
| 7.44 | 0.39 | -36.95 | 8.57 | N | N | N |
| 12.31 | 1.02 | 4.32 | 10.71 | A | A | A |

Labcode

173

174

176

180

182

183

185

190

199

204

206

217

218

221

222

223

228

230

235

240

243

244

245

248

258

260

268

270

271

277

279

281

294

298

299

301

302

305

307

315

323

333

335

Reported Results for Pb-210 determination in Sample 4

TABLE 38. Target values

| Parameter | Value |
| --- | --- |
| Activity Concentration [ Bq/kg ] | 26.6 |
| Uncertainty [ Bq/kg ] | 1.8 |
| Maximum Acceptable Relative Bias, [%] | 30 |

Warning 2%

Not Accepted 4%

Activity concentration  [ Bq/kg ]

25

30

Accepted 18%

Not Reported 76%

35

40

FIGURE 35. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

10

15

20

249

85

145

53

29

112

271

146

218

174

151

23

232

170

107

334

125

10

135

71

305

302

171

122

130

12

293

217

158

156

210

144

142

21

22

66

272

200

224

233

275

131

155

222

259

231

67

69

221

201

299

123

247

205

189

344

206

173

79

54

165

281

180

128

157

234

263

301

176

28

184

118

240

13

307

154

149

294

Laboratory code

FIGURE 36. Reported results and evaluation parameters for Pb-210. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

| Reported value [Bq/kg ] 22.8 | Reported uncertainty (k = 1) [Bq/kg ] 2.5 | Relative bias [%] -14.29 | P-Test [%] 12.88 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 24 | 3 | -9.77 | 14.21 | A | A | A |
| 43.1 | 6.2 | 62.03 | 15.90 | N | N | N |
| 25.1 | 9.5 | -5.64 | 38.45 | A | N | W |
| 25.2 | 4.36 | -5.26 | 18.58 | A | A | A |
| 22.0 | 4.7 | -17.29 | 22.41 | A | A | A |
| 38.1 | 9.7 | 43.23 | 26.34 | N | A | N |
| 20 | 3 | -24.81 | 16.46 | A | A | A |
| 19.3 | 3.2 | -27.44 | 17.91 | A | A | A |
| 31 | 13 | 16.54 | 42.48 | A | N | W |
| 25.48 | 2.5 | -4.21 | 11.92 | A | A | A |
| 27.21 | 4.34 | 2.29 | 17.33 | A | A | A |
| 27.47 | 1.07 | 3.27 | 7.81 | A | A | A |
| 23 | 4 | -13.53 | 18.66 | A | A | A |
| 30.9 | 3.85 | 16.17 | 14.18 | A | A | A |
| 15.2 | 3.0 | -42.86 | 20.86 | N | A | N |
| 22.2 | 4.1 | -16.54 | 19.67 | A | A | A |
| 21 | 5 | -21.05 | 24.75 | A | A | A |
| 42.1 | 10.2 | 58.27 | 25.16 | N | A | N |
| 23.9 | 3.9 | -10.15 | 17.67 | A | A | A |
| 28.8 | 4.2 | 8.27 | 16.08 | A | A | A |
| 22.54 | 4.14 | -15.26 | 19.57 | A | A | A |
| 32.80 | 3.77 | 23.31 | 13.34 | A | A | A |
| 24 | 3 | -9.77 | 14.21 | A | A | A |
| 25.67 | 2.69 | -3.50 | 12.47 | A | A | A |
| 23.0 | 2.8 | -13.53 | 13.93 | A | A | A |
| 24.937 | 2.280 | -6.25 | 11.37 | A | A | A |
| 24.8 | 3.0 | -6.77 | 13.86 | A | A | A |
| 19.2 | 3.6 | -27.82 | 19.93 | A | A | A |
| 21.1 | 2.18 | -20.68 | 12.35 | A | A | A |
| 157.65 | 45.476 | 492.67 | 29.63 | N | N | N |
| 22.0 | 3.3 | -17.29 | 16.46 | A | A | A |
| 147 | 14 | 452.63 | 11.68 | N | N | N |
| 25.7 | 4.5 | -3.38 | 18.77 | A | A | A |
| 24.5 | 3.2 | -7.89 | 14.71 | A | A | A |
| 33.6 | 6.8 | 26.32 | 21.34 | A | A | A |
| 24.43 | 1.45 | -8.16 | 9.00 | A | A | A |
| 31.49 | 1.33 | 18.38 | 7.98 | A | A | A |
| 22.1 | 4.1 | -16.92 | 19.75 | A | A | A |
| 23.4 | 2.4 | -12.03 | 12.29 | A | A | A |
| 30.23 | 4.9 | 13.65 | 17.56 | A | A | A |
| 22 | 4 | -17.29 | 19.40 | A | A | A |
| 38 | 5 | 42.86 | 14.80 | N | N | N |
| 31.59 | 7.31 | 18.76 | 24.11 | A | A | A |
| 39.6 | 4.7 | 48.87 | 13.66 | N | N | N |
| 29.5 | 6.5 | 10.90 | 23.05 | A | A | A |
| 25.5 | 9.3 | -4.14 | 37.09 | A | N | W |
| 28.3 | 4.2 | 6.39 | 16.31 | A | A | A |
| 29.05 | 4.20 | 9.21 | 15.96 | A | A | A |
| 30.23 | 9.31 | 13.65 | 31.53 | A | N | W |
| 24.67 | 2.4 | -7.26 | 11.85 | A | A | A |
| 24.4 | 5.9 | -8.27 | 25.11 | A | A | A |
| 21.8 | 7.6 | -18.05 | 35.51 | A | N | W |
| 28.30 | 2.35 | 6.39 | 10.71 | A | A | A |
| 25.80 | 4.25 | -3.01 | 17.81 | A | A | A |
| 25.52 | 3.8 | -4.06 | 16.36 | A | A | A |
| 27.15 | 4.29 | 2.07 | 17.19 | A | A | A |
| 22 | 3 | -17.29 | 15.22 | A | A | A |
| 25.65 | 3.97 | -3.57 | 16.89 | A | A | A |
| 34.39 | 2.77 | 29.29 | 10.52 | A | N | W |

TABLE 39. Evaluation Results for Pb-210 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

10

12

13

21

22

23

28

29

53

54

66

67

69

71

79

85

107

112

118

122

123

125

128

130

131

135

142

144

145

146

149

151

154

155

156

157

158

165

170

171

173

174

176

180

184

189

200

201

205

206

210

217

218

221

222

224

231

232

233

234

TABLE 39. Evaluation Results for Pb-210 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

| Reported value [Bq/kg ] 42.2 | Reported uncertainty (k = 1) [Bq/kg ] 2.9 | Relative bias [%] 58.65 | P-Test [%] 9.64 | Trueness evaluation N | Precision evaluation N | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 28.82 | 1.1 | 8.35 | 7.77 | A | A | A |
| 11.7 | 2.8 | -56.02 | 24.87 | N | A | N |
| 27.1 | 0.9 | 1.88 | 7.54 | A | A | A |
| 35.6 | 4.4 | 33.83 | 14.09 | N | A | N |
| 21.1 | 0.5 | -20.68 | 7.17 | A | N | W |
| 25.5 | 2.1 | -4.14 | 10.66 | A | A | A |
| 25.67 | 5.87 | -3.50 | 23.85 | A | A | A |
| 31.57 | 3.157 | 18.68 | 12.07 | A | A | A |
| 24.1 | 1.7 | -9.40 | 9.77 | A | A | A |
| 167 | 43 | 527.82 | 26.62 | N | N | N |
| 28.5 | 1.65 | 7.14 | 8.91 | A | A | A |
| 36.49 | 5.14 | 37.18 | 15.63 | N | A | N |
| 23.4 | 3.6 | -12.03 | 16.81 | A | A | A |
| 23 | 5 | -13.53 | 22.77 | A | A | A |
| 49.1 | 3.1 | 84.59 | 9.25 | N | N | N |
| 22.3 | 2.4 | -16.17 | 12.71 | A | A | A |
| 29.5 | 2.6 | 10.90 | 11.11 | A | A | A |

Labcode

240

247

249

259

263

271

272

275

281

293

294

299

301

302

305

307

334

344

Activity concentration  [ Bq/kg ]

30

40

Reported Results for Bi-212 determination in Sample 4

TABLE 40. Target values

| Parameter | Value |
| --- | --- |
| Activity Concentration [ Bq/kg ] | 32.9 |
| Uncertainty [ Bq/kg ] | 1.9 |
| Maximum Acceptable Relative Bias, [%] | 30 |

Warning 2%

Not Accepted 6%

Accepted 13%

Not Reported 79%

50

FIGURE 37. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

20

222

226

125

132

342

205

223

28

299

74

85

75

170

72

13

54

275

33

34

182

183

206

21

69

63

277

209

17

66

180

228

173

144

90

244

270

189

281

30

233

174

218

155

161

258

2

42

240

171

57

268

315

99

118

93

102

294

149

190

31

89

217

68

301

323

Laboratory code

FIGURE 38. Reported results and evaluation parameters for Bi-212. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

| Reported value [Bq/kg ] 37.41 | Reported uncertainty (k = 1) [Bq/kg ] 4.93 | Relative bias [%] 13.71 | P-Test [%] 14.39 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 25.1 | 1 | -23.71 | 7.02 | A | N | W |
| 32.4 | 3.3 | -1.52 | 11.71 | A | A | A |
| 31.7 | 4.7 | -3.65 | 15.91 | A | A | A |
| 20.7 | 5.4 | -37.08 | 26.72 | N | A | N |
| 34.5 | 2.5 | 4.86 | 9.27 | A | A | A |
| 49.4 | 9.9 | 50.15 | 20.86 | N | A | N |
| 30.0 | 3.3 | -8.81 | 12.42 | A | A | A |
| 31.2 | 4.9 | -5.17 | 16.73 | A | A | A |
| 37.9 | 6.69 | 15.20 | 18.57 | A | A | A |
| 27.3 | 3.3 | -17.02 | 13.40 | A | A | A |
| 39 | 11 | 18.54 | 28.79 | A | A | A |
| 32.0 | 2.2 | -2.74 | 8.98 | A | A | A |
| 32.72 | 2.6 | -0.55 | 9.82 | A | A | A |
| 54.1 | 4.8 | 64.44 | 10.59 | N | N | N |
| 31.73 | 2.56 | -3.56 | 9.92 | A | A | A |
| 24.70 | 1.85 | -24.92 | 9.46 | A | N | W |
| 21.12 | 2.62 | -35.81 | 13.68 | N | N | N |
| 23.82 | 2.34 | -27.60 | 11.40 | A | A | A |
| 22.0 | 4.4 | -33.13 | 20.82 | N | A | N |
| 52.3 | 3.0 | 58.97 | 8.14 | N | N | N |
| 33.54 | 12.47 | 1.95 | 37.63 | A | N | W |
| 42 | 8 | 27.66 | 19.90 | A | A | A |
| 41.3 | 13 | 25.53 | 32.00 | A | N | W |
| 42.32 | 2.87 | 28.63 | 8.91 | A | N | W |
| 41.3 | 3.5 | 25.53 | 10.26 | A | A | A |
| 16.68 | 1.82 | -49.30 | 12.35 | N | N | N |
| 18.3 | 1.6 | -44.38 | 10.48 | N | N | N |
| 33.4 | 5.0 | 1.52 | 16.05 | A | A | A |
| 44.06 | 6.7685 | 33.92 | 16.41 | N | A | N |
| 35.8 | 2.5 | 8.81 | 9.06 | A | A | A |
| 36.0 | 6.1 | 9.42 | 17.90 | A | A | A |
| 24.5 | 2.5 | -25.53 | 11.72 | A | A | A |
| 38.9 | 4.16 | 18.24 | 12.15 | A | A | A |
| 33.4 | 3.4 | 1.52 | 11.70 | A | A | A |
| 35 | 4 | 6.38 | 12.80 | A | A | A |
| 32.74 | 6.60 | -0.49 | 20.97 | A | A | A |
| 31.3 | 3 | -4.86 | 11.19 | A | A | A |
| 31.4 | 2.4 | -4.56 | 9.58 | A | A | A |
| 34 | 4.5 | 3.34 | 14.44 | A | A | A |
| 44.5 | 5.2 | 35.26 | 13.03 | N | N | N |
| 19.45 | 2.67 | -40.88 | 14.89 | N | N | N |
| 31.57 | 6.53 | -4.04 | 21.48 | A | A | A |
| 32.23 | 3.17 | -2.04 | 11.41 | A | A | A |
| 53.7 | 6.7 | 63.22 | 13.75 | N | N | N |
| 35.3 | 5.45 | 7.29 | 16.48 | A | A | A |
| 7.17 | 0.57 | -78.21 | 9.83 | N | N | N |
| 19.50 | 2.35 | -40.73 | 13.36 | N | N | N |
| 10.89 | 0.02 | -66.90 | 5.78 | N | N | N |
| 32.9 | 3.1 | 0.00 | 11.05 | A | A | A |
| 34.62 | 4.27 | 5.23 | 13.62 | A | A | A |
| 38.1 | 2.9 | 15.81 | 9.55 | A | A | A |
| 33.8 | 3.28 | 2.74 | 11.29 | A | A | A |
| 37.36 | 2.08 | 13.56 | 8.02 | A | A | A |
| 39.7 | 5.0 | 20.67 | 13.86 | A | A | A |
| 33.8 | 4.5 | 2.74 | 14.51 | A | A | A |
| 28.26 | 6.5 | -14.10 | 23.71 | A | A | A |
| 32.22 | 1.30 | -2.07 | 7.04 | A | A | A |
| 34.38 | 3.438 | 4.50 | 11.55 | A | A | A |
| 44.0 | 5.6 | 33.74 | 13.98 | N | A | N |

TABLE 41. Evaluation Results for Bi-212 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

2

13

17

21

28

30

31

33

34

42

54

57

63

66

68

69

72

74

75

85

89

90

93

99

102

118

125

132

144

149

155

161

170

171

173

174

180

182

183

189

190

205

206

209

217

218

222

223

226

228

233

240

244

258

268

270

275

277

281

294

TABLE 41. Evaluation Results for Bi-212 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

| Reported value [Bq/kg ] 20.7 | Reported uncertainty (k = 1) [Bq/kg ] 1.7 | Relative bias [%] -37.08 | P-Test [%] 10.04 | Trueness evaluation N | Precision evaluation N | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 65.89 | 6.83 | 100.27 | 11.87 | N | N | N |
| 40.0 | 8.5 | 21.58 | 22.02 | A | A | A |
| 87.49 | 0.14 | 165.93 | 5.78 | N | N | N |
| 19.19 | 1.60 | -41.67 | 10.14 | N | N | N |

Labcode

299

301

315

323

342

Activity concentration  [ Bq/kg ]

30

40

Reported Results for Pb-212 determination in Sample 4

TABLE 42. Target values

| Parameter | Value |
| --- | --- |
| Activity Concentration [ Bq/kg ] | 32.9 |
| Uncertainty [ Bq/kg ] | 1.9 |
| Maximum Acceptable Relative Bias, [%] | 30 |

Not Accepted 5%

Warning 3%

Accepted 29%

Not Reported 63%

50

FIGURE 39. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

20

222

335

154

223

21

125

226

233

183

74

294

75

85

112

71

305

170

57

23

230

51

174

27

134

182

271

107

66

63

306

277

209

221

77

17

129

161

260

122

10

130

12

79

314

42

244

190

299

158

199

156

70

186

268

275

33

341

281

2

30

180

69

31

228

90

34

155

258

328

146

81

151

248

185

251

245

144

218

173

29

298

171

166

270

124

53

279

205

54

235

14

240

307

48

259

72

142

206

118

149

217

86

165

13

315

132

99

102

243

67

89

87

68

301

96

323

93

Laboratory code

FIGURE 40. Reported results and evaluation parameters for Pb-212. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

| Reported value [Bq/kg ] 34.50 | Reported uncertainty (k = 1) [Bq/kg ] 1.42 | Relative bias [%] 4.86 | P-Test [%] 7.09 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 33.0 | 1.5 | 0.30 | 7.35 | A | A | A |
| 33 | 2 | 0.30 | 8.37 | A | A | A |
| 42.5 | 1.1 | 29.18 | 6.33 | A | N | W |
| 38.6 | 2.3 | 17.33 | 8.30 | A | A | A |
| 32.4 | 3.3 | -1.52 | 11.71 | A | A | A |
| 25.2 | 7.1 | -23.40 | 28.76 | A | A | A |
| 30.3 | 1.6 | -7.90 | 7.83 | A | A | A |
| 31.1 | 0.8 | -5.47 | 6.32 | A | A | A |
| 36 | 2 | 9.42 | 8.01 | A | A | A |
| 34.5 | 1.6 | 4.86 | 7.41 | A | A | A |
| 34.7 | 3.4 | 5.47 | 11.37 | A | A | A |
| 34.3 | 2.5 | 4.26 | 9.30 | A | A | A |
| 34.9 | 4.8 | 6.08 | 14.92 | A | A | A |
| 33.3 | 3.16 | 1.22 | 11.11 | A | A | A |
| 39.2 | 3.2 | 19.15 | 10.00 | A | A | A |
| 30.7 | 1.7 | -6.69 | 8.00 | A | A | A |
| 37.6 | 1.6 | 14.29 | 7.17 | A | A | A |
| 38.3 | 0.7 | 16.41 | 6.06 | A | N | W |
| 30.2 | 2.5 | -8.21 | 10.09 | A | A | A |
| 31.8 | 0.7 | -3.34 | 6.18 | A | A | A |
| 31.74 | 1.0 | -3.53 | 6.58 | A | A | A |
| 50.43 | 1.48 | 53.28 | 6.48 | N | N | N |
| 55.0 | 2.4 | 67.17 | 7.24 | N | N | N |
| 34.60 | 1.52 | 5.17 | 7.26 | A | A | A |
| 34.1 | 3.1 | 3.65 | 10.77 | A | A | A |
| 30 | 3 | -8.81 | 11.55 | A | A | A |
| 39.41 | 0.73 | 19.79 | 6.06 | A | N | W |
| 28.98 | 0.74 | -11.91 | 6.31 | A | A | A |
| 29.40 | 1.71 | -10.64 | 8.20 | A | A | A |
| 32.3 | 1.3 | -1.82 | 7.04 | A | A | A |
| 33.1 | 1.2 | 0.61 | 6.82 | A | A | A |
| 35.5 | 1.1 | 7.90 | 6.55 | A | A | A |
| 30.0 | 6.0 | -8.81 | 20.82 | A | A | A |
| 41.5 | 3.3 | 26.14 | 9.83 | A | N | W |
| 52.31 | 0.55 | 59.00 | 5.87 | N | N | N |
| 52.3 | 3.0 | 58.97 | 8.14 | N | N | N |
| 34.86 | 1.65 | 5.96 | 7.47 | A | A | A |
| 80 | 5 | 143.16 | 8.51 | N | N | N |
| 62.65 | 2.10 | 90.43 | 6.68 | N | N | N |
| 48.0 | 1.8 | 45.90 | 6.89 | N | N | N |
| 48.14 | 1.58 | 46.32 | 6.64 | N | N | N |
| 31.6 | 1.0 | -3.95 | 6.59 | A | A | A |
| 30 | 3 | -8.81 | 11.55 | A | A | A |
| 40.1 | 3.3 | 21.88 | 10.05 | A | A | A |
| 32.8 | 2.2 | -0.30 | 8.85 | A | A | A |
| 37.45 | 1.55 | 13.83 | 7.11 | A | A | A |
| 26.68 | 2.12 | -18.91 | 9.82 | A | A | A |
| 32.50 | 4.50 | -1.22 | 15.00 | A | A | A |
| 33 | 2 | 0.30 | 8.37 | A | A | A |
| 43.4 | 3.9 | 31.91 | 10.68 | N | N | N |
| 31.2 | 1.2 | -5.17 | 6.94 | A | A | A |
| 39.694 | 2.998 | 20.65 | 9.51 | A | A | A |
| 35.8 | 2.5 | 8.81 | 9.06 | A | A | A |
| 35.43 | 6.2 | 7.69 | 18.43 | A | A | A |
| 40.61 | 2.38835 | 23.43 | 8.24 | A | N | W |
| 35.6 | 1.9 | 8.21 | 7.86 | A | A | A |
| 18.3 | 0.9 | -44.38 | 7.59 | N | N | N |
| 34.9 | 2.6 | 6.08 | 9.43 | A | A | A |
| 33.9 | 3.3 | 3.04 | 11.32 | A | A | A |

TABLE 43. Evaluation Results for Pb-212 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

2

10

12

13

14

17

21

23

27

29

30

31

33

34

42

48

51

53

54

57

63

66

67

68

69

70

71

72

74

75

77

79

81

85

86

87

89

90

93

96

99

102

107

112

118

122

124

125

129

130

132

134

142

144

146

149

151

154

155

156

TABLE 43. Evaluation Results for Pb-212 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

| Reported value [Bq/kg ] 33.83 | Reported uncertainty (k = 1) [Bq/kg ] 1.56 | Relative bias [%] 2.83 | P-Test [%] 7.39 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 32.5 | 3.1 | -1.22 | 11.15 | A | A | A |
| 42.36 | 0.51 | 28.75 | 5.90 | A | N | W |
| 36 | 4.6 | 9.42 | 14.02 | A | A | A |
| 30.2 | 2.4 | -8.21 | 9.82 | A | A | A |
| 36 | 1.21 | 9.42 | 6.68 | A | A | A |
| 35.9 | 1.9 | 9.12 | 7.83 | A | A | A |
| 31 | 3 | -5.78 | 11.27 | A | A | A |
| 34.51 | 3.07 | 4.89 | 10.61 | A | A | A |
| 31.3 | 3 | -4.86 | 11.19 | A | A | A |
| 28.8 | 1.5 | -12.46 | 7.78 | A | A | A |
| 35.70 | 4.64 | 8.51 | 14.22 | A | A | A |
| 34.2 | 1.0 | 3.95 | 6.47 | A | A | A |
| 33.5 | 1.9 | 1.82 | 8.09 | A | A | A |
| 33.84 | 0.84 | 2.86 | 6.29 | A | A | A |
| 38.03 | 1.03 | 15.59 | 6.38 | A | A | A |
| 40.05 | 2.63 | 21.73 | 8.74 | A | A | A |
| 32.23 | 2.89 | -2.04 | 10.67 | A | A | A |
| 40.9 | 7.0 | 24.32 | 18.06 | A | A | A |
| 35.8 | 2.86 | 8.81 | 9.86 | A | A | A |
| 32.30 | 0.51 | -1.82 | 5.99 | A | A | A |
| 7.17 | 0.57 | -78.21 | 9.83 | N | N | N |
| 24.90 | 1.41 | -24.32 | 8.09 | A | N | W |
| 28.17 | 0.02 | -14.38 | 5.78 | A | A | A |
| 34.7 | 2.4 | 5.47 | 9.01 | A | A | A |
| 30.53 | 1.12 | -7.20 | 6.84 | A | A | A |
| 28.57 | 2.46 | -13.16 | 10.37 | A | A | A |
| 38.5 | 2.2 | 17.02 | 8.12 | A | A | A |
| 38.8 | 2.5 | 17.93 | 8.65 | A | A | A |
| 49.7 | 5 | 51.06 | 11.60 | N | N | N |
| 33.4 | 2.45 | 1.52 | 9.34 | A | A | A |
| 35.8 | 2.5 | 8.81 | 9.06 | A | A | A |
| 35.7 | 2.5 | 8.51 | 9.08 | A | A | A |
| 35.73 | 1.56 | 8.60 | 7.24 | A | A | A |
| 34.93 | 1.20 | 6.17 | 6.72 | A | A | A |
| 39.4 | 1.4 | 19.76 | 6.78 | A | N | W |
| 32.66 | 3.27 | -0.73 | 11.56 | A | A | A |
| 34.2 | 2.1 | 3.95 | 8.43 | A | A | A |
| 36.5 | 4.2 | 10.94 | 12.87 | A | A | A |
| 31.4 | 1.7 | -4.56 | 7.92 | A | A | A |
| 34.25 | 7.6 | 4.10 | 22.93 | A | A | A |
| 32.22 | 1.30 | -2.07 | 7.04 | A | A | A |
| 37.78 | 4.6 | 14.83 | 13.48 | A | A | A |
| 34.48 | 3.448 | 4.80 | 11.55 | A | A | A |
| 29.2 | 2.2 | -11.25 | 9.49 | A | A | A |
| 36.0 | 1.3 | 9.42 | 6.81 | A | A | A |
| 33.7 | 1.8 | 2.43 | 7.87 | A | A | A |
| 57.74 | 3.56 | 75.50 | 8.45 | N | N | N |
| 30 | 2 | -8.81 | 8.82 | A | A | A |
| 32.13 | 2.93 | -2.34 | 10.79 | A | A | A |
| 38.8 | 5.0 | 17.93 | 14.12 | A | A | A |
| 33.2 | 2.9 | 0.91 | 10.47 | A | A | A |
| 42.9 | 3.6 | 30.40 | 10.19 | N | N | N |
| 68.33 | 0.04 | 107.69 | 5.78 | N | N | N |
| 35.2 | 1.5 | 6.99 | 7.18 | A | A | A |
| 11.86 | 1.12 | -63.95 | 11.07 | N | N | N |
| 34.47 | 2.98 | 4.77 | 10.40 | A | A | A |

Labcode

158

161

165

166

170

171

173

174

180

182

183

185

186

190

199

205

206

209

217

218

221

222

223

226

228

230

233

235

240

243

244

245

248

251

258

259

260

268

270

271

275

277

279

281

294

298

299

301

305

306

307

314

315

323

328

335

341

Reported Results for Bi-214 determination in Sample 4

TABLE 44. Target values

| Parameter | Value |
| --- | --- |
| Activity Concentration [ Bq/kg ] | 20.8 |
| Uncertainty [ Bq/kg ] | 1.0 |
| Maximum Acceptable Relative Bias, [%] | 30 |

Not Accepted 7%

Warning 2%

Accepted 27%

Activity concentration  [ Bq/kg ]

20

25

Not Reported 64%

30

FIGURE 41. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

10

15

226

294

214

85

343

70

165

74

42

75

170

132

161

205

46

79

328

124

33

63

240

298

223

23

218

183

260

112

5

209

239

245

306

30

323

99

18

14

51

144

54

281

125

233

174

305

275

66

58

171

27

21

142

235

268

342

141

247

151

72

270

248

2

17

28

128

206

146

36

47

182

140

180

251

314

199

52

22

89

34

53

86

48

81

222

186

149

271

190

57

158

315

279

69

13

301

246

307

162

230

77

102

153

110

93

68

87

118

176

184

129

Laboratory code

FIGURE 42. Reported results and evaluation parameters for Bi-214. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

| Reported value [Bq/kg ] 20.04 | Reported uncertainty (k = 1) [Bq/kg ] 1.00 | Relative bias [%] -3.65 | P-Test [%] 6.93 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 18.18 | 1.54 | -12.60 | 9.74 | A | A | A |
| 24.1 | 0.4 | 15.87 | 5.09 | A | N | W |
| 18.7 | 1.1 | -10.10 | 7.60 | A | A | A |
| 20.1 | 1.5 | -3.37 | 8.88 | A | A | A |
| 18.66 | 1.3 | -10.29 | 8.46 | A | A | A |
| 19.4 | 1.1 | -6.73 | 7.43 | A | A | A |
| 21.53 | 0.89 | 3.51 | 6.34 | A | A | A |
| 17.6 | 0.9 | -15.38 | 7.02 | A | A | A |
| 19.4 | 0.72 | -6.73 | 6.07 | A | A | A |
| 20.1 | 3.2 | -3.37 | 16.63 | A | A | A |
| 18.5 | 1.0 | -11.06 | 7.23 | A | A | A |
| 17.2 | 1.3 | -17.31 | 8.96 | A | A | A |
| 21.7 | 2.7 | 4.33 | 13.34 | A | A | A |
| 20.5 | 2.5 | -1.44 | 13.11 | A | A | A |
| 13.60 | 2.02 | -34.62 | 15.61 | N | A | N |
| 17 | 1 | -18.27 | 7.60 | A | A | A |
| 20.5 | 0.6 | -1.44 | 5.63 | A | A | A |
| 22.2 | 1.4 | 6.73 | 7.93 | A | A | A |
| 18.8 | 1.2 | -9.62 | 7.99 | A | A | A |
| 21.5 | 1.7 | 3.37 | 9.25 | A | A | A |
| 21.9 | 1.3 | 5.29 | 7.64 | A | A | A |
| 18.9 | 1.4 | -9.13 | 8.83 | A | A | A |
| 23.7 | 4.2 | 13.94 | 18.36 | A | A | A |
| 19.3 | 2.7 | -7.21 | 14.79 | A | A | A |
| 17.3 | 0.5 | -16.83 | 5.61 | A | N | W |
| 19.01 | 0.57 | -8.61 | 5.67 | A | A | A |
| 29.3 | 1.4 | 40.87 | 6.78 | N | N | N |
| 23.96 | 1.25 | 15.19 | 7.09 | A | A | A |
| 11.1 | 1.4 | -46.63 | 13.50 | N | N | N |
| 19.90 | 0.98 | -4.33 | 6.88 | A | A | A |
| 13.29 | 0.77 | -36.11 | 7.53 | N | N | N |
| 14.42 | 0.80 | -30.67 | 7.34 | N | N | N |
| 26.3 | 1.0 | 26.44 | 6.13 | A | N | W |
| 17 | 1.3 | -18.27 | 9.03 | A | A | A |
| 22.5 | 0.7 | 8.17 | 5.73 | A | A | A |
| 10.7 | 2.1 | -48.56 | 20.21 | N | A | N |
| 22.2 | 1.6 | 6.73 | 8.66 | A | A | A |
| 29.41 | 1.87 | 41.39 | 7.97 | N | N | N |
| 21.63 | 1.84 | 3.99 | 9.77 | A | A | A |
| 29 | 4 | 39.42 | 14.61 | N | N | N |
| 18.6 | 1.8 | -10.58 | 10.81 | A | A | A |
| 27.22 | 2.88 | 30.87 | 11.62 | N | N | N |
| 28.64 | 1.83 | 37.69 | 8.00 | N | N | N |
| 18 | 1 | -13.46 | 7.35 | A | A | A |
| 30.4 | 1.5 | 46.15 | 6.89 | N | N | N |
| 17.14 | 0.56 | -17.60 | 5.81 | A | N | W |
| 19.0 | 1.92 | -8.65 | 11.19 | A | A | A |
| 20.25 | 1.17 | -2.64 | 7.52 | A | A | A |
| 166.80 | 22.70 | 701.92 | 14.43 | N | N | N |
| 16.0 | 1.4 | -23.08 | 9.98 | A | A | A |
| 20.8 | 4.4 | 0.00 | 21.69 | A | A | A |
| 19.64 | 0.75 | -5.58 | 6.14 | A | A | A |
| 19.489 | 3.193 | -6.30 | 17.07 | A | A | A |
| 18.9 | 3.0 | -9.13 | 16.59 | A | A | A |
| 20.42 | 1.24 | -1.83 | 7.75 | A | A | A |
| 22.917 | 1.63135 | 10.18 | 8.59 | A | A | A |
| 19.8 | 1.3 | -4.81 | 8.14 | A | A | A |
| 28.154 | 0.843 | 35.36 | 5.66 | N | N | N |
| 23.85 | 1.05 | 14.66 | 6.52 | A | A | A |

TABLE 45. Evaluation Results for Bi-214 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

2

5

13

14

17

18

21

22

23

27

28

30

33

34

36

42

46

47

48

51

52

53

54

57

58

63

66

68

69

70

72

74

75

77

79

81

85

86

87

89

93

99

102

110

112

118

124

125

128

129

132

140

141

142

144

146

149

151

153

158

TABLE 45. Evaluation Results for Bi-214 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

| Reported value [Bq/kg ] 16.7 | Reported uncertainty (k = 1) [Bq/kg ] 1.7 | Relative bias [%] -19.71 | P-Test [%] 11.26 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 25.5 | 2.0 | 22.60 | 9.20 | A | A | A |
| 12.64 | 0.62 | -39.23 | 6.87 | N | N | N |
| 14.5 | 1.0 | -30.29 | 8.41 | N | N | N |
| 19.3 | 0.82 | -7.21 | 6.42 | A | A | A |
| 19 | 2 | -8.65 | 11.57 | A | A | A |
| 34 | 2 | 63.46 | 7.60 | N | N | N |
| 20.87 | 1.25 | 0.34 | 7.68 | A | A | A |
| 20.6 | 2 | -0.96 | 10.83 | A | A | A |
| 17.7 | 1.1 | -14.90 | 7.86 | A | A | A |
| 35.2 | 4.2 | 69.23 | 12.86 | N | N | N |
| 22.9 | 1.1 | 10.10 | 6.80 | A | A | A |
| 23.3 | 1.3 | 12.02 | 7.37 | A | A | A |
| 21.12 | 1.69 | 1.54 | 9.34 | A | A | A |
| 16.71 | 0.76 | -19.66 | 6.62 | A | N | W |
| 20.28 | 1.34 | -2.50 | 8.17 | A | A | A |
| 18.25 | 1.37 | -12.26 | 8.91 | A | A | A |
| 9.5 | 0.9 | -54.33 | 10.62 | N | N | N |
| 17.6 | 1.25 | -15.38 | 8.58 | A | A | A |
| 22.70 | 1.84 | 9.13 | 9.42 | A | A | A |
| 17.60 | 0.96 | -15.38 | 7.27 | A | A | A |
| 0.46 | 0.02 | -97.79 | 6.48 | N | N | N |
| 25.61 | 1.30 | 23.12 | 6.99 | A | N | W |
| 19.00 | 1.48 | -8.65 | 9.15 | A | A | A |
| 19.5 | 1.3 | -6.25 | 8.22 | A | A | A |
| 18.28 | 0.80 | -12.12 | 6.50 | A | A | A |
| 17.5 | 1.0 | -15.87 | 7.47 | A | A | A |
| 18.3 | 1.6 | -12.02 | 9.98 | A | A | A |
| 24.9 | 2.3 | 19.71 | 10.41 | A | A | A |
| 19.76 | 0.7 | -5.00 | 5.97 | A | A | A |
| 20.0 | 1.5 | -3.85 | 8.91 | A | A | A |
| 20.90 | 1.93 | 0.48 | 10.41 | A | A | A |
| 17.74 | 2.66 | -14.71 | 15.75 | A | A | A |
| 19.5 | 1.4 | -6.25 | 8.64 | A | A | A |
| 19.9 | 2.1 | -4.33 | 11.60 | A | A | A |
| 23.0 | 1.0 | 10.58 | 6.48 | A | A | A |
| 19.01 | 1.62 | -8.61 | 9.78 | A | A | A |
| 23.94 | 1.7 | 15.10 | 8.58 | A | A | A |
| 18.97 | 2.086 | -8.80 | 12.00 | A | A | A |
| 6.4 | 1.5 | -69.23 | 23.93 | N | N | N |
| 17.5 | 0.9 | -15.87 | 7.04 | A | A | A |
| 24.70 | 1.91 | 18.75 | 9.11 | A | A | A |
| 19 | 1 | -8.65 | 7.13 | A | A | A |
| 18.36 | 1.73 | -11.73 | 10.58 | A | A | A |
| 25.2 | 3.3 | 21.15 | 13.95 | A | A | A |
| 21.1 | 2.4 | 1.44 | 12.35 | A | A | A |
| 23.9 | 2.3 | 14.90 | 10.76 | A | A | A |
| 18.55 | 0.03 | -10.82 | 4.81 | A | A | A |
| 17.1 | 0.9 | -17.79 | 7.13 | A | A | A |
| 19.52 | 0.64 | -6.15 | 5.82 | A | A | A |
| 10.7407 |  | -48.36 |  | N |  | N |

Labcode

161

162

165

170

171

174

176

180

182

183

184

186

190

199

205

206

209

214

218

222

223

226

230

233

235

239

240

245

246

247

248

251

260

268

270

271

275

279

281

294

298

301

305

306

307

314

315

323

328

342

343

Reported Results for Pb-214 determination in Sample 4

TABLE 46. Target values

| Parameter | Value |
| --- | --- |
| Activity Concentration [ Bq/kg ] | 20.8 |
| Uncertainty [ Bq/kg ] | 1.0 |
| Maximum Acceptable Relative Bias, [%] | 30 |

Not Accepted 7%

Warning 3%

Accepted 26%

Activity concentration  [ Bq/kg ]

20

25

Not Reported 64%

30

FIGURE 43. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

10

15

226

294

165

57

74

154

42

214

75

170

161

223

183

142

218

63

240

112

23

328

209

30

281

245

124

18

66

85

305

260

233

21

275

58

33

306

27

17

315

125

174

79

180

51

298

268

72

54

247

2

34

47

248

182

270

128

144

199

36

329

146

14

171

99

151

28

46

190

205

259

53

314

186

222

81

271

204

322

69

52

153

132

22

235

158

251

342

48

162

230

307

246

86

206

279

77

301

89

110

149

341

87

118

13

102

176

68

96

184

241

323

93

343

129

Laboratory code

FIGURE 44. Reported results and evaluation parameters for Pb-214. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

| Reported value [Bq/kg ] 20.44 | Reported uncertainty (k = 1) [Bq/kg ] 1.09 | Relative bias [%] -1.73 | P-Test [%] 7.18 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 31.1 | 0.7 | 49.52 | 5.31 | N | N | N |
| 21.4 | 1.4 | 2.88 | 8.12 | A | A | A |
| 19.8 | 1.9 | -4.81 | 10.73 | A | A | A |
| 18.76 | 1.28 | -9.81 | 8.35 | A | A | A |
| 19.2 | 1.1 | -7.69 | 7.48 | A | A | A |
| 23.55 | 1.55 | 13.22 | 8.15 | A | A | A |
| 18.0 | 1.0 | -13.46 | 7.35 | A | A | A |
| 19.8 | 1.1 | -4.81 | 7.35 | A | A | A |
| 21.9 | 2.3 | 5.29 | 11.55 | A | A | A |
| 18.5 | 1.0 | -11.06 | 7.23 | A | A | A |
| 19.6 | 1.6 | -5.77 | 9.47 | A | A | A |
| 20.5 | 4.1 | -1.44 | 20.57 | A | A | A |
| 21.0 | 2.4 | 0.96 | 12.40 | A | A | A |
| 14.50 | 2.59 | -30.29 | 18.50 | N | A | N |
| 22 | 1 | 5.77 | 6.62 | A | A | A |
| 20.5 | 0.6 | -1.44 | 5.63 | A | A | A |
| 24.8 | 1.9 | 19.23 | 9.04 | A | A | A |
| 20.1 | 1.2 | -3.37 | 7.67 | A | A | A |
| 23.3 | 2.3 | 12.02 | 10.98 | A | A | A |
| 22.4 | 1.2 | 7.69 | 7.20 | A | A | A |
| 20.3 | 1.0 | -2.40 | 6.88 | A | A | A |
| 14.1 | 3.9 | -32.21 | 28.07 | N | A | N |
| 19.5 | 2.6 | -6.25 | 14.17 | A | A | A |
| 17.8 | 0.5 | -14.42 | 5.57 | A | N | W |
| 18.81 | 0.57 | -9.57 | 5.68 | A | A | A |
| 33.6 | 1.6 | 61.54 | 6.77 | N | N | N |
| 23.24 | 1.15 | 11.73 | 6.90 | A | A | A |
| 20.26 | 1.34 | -2.60 | 8.18 | A | A | A |
| 14.27 | 0.98 | -31.39 | 8.38 | N | N | N |
| 15.09 | 0.94 | -27.45 | 7.87 | A | N | W |
| 26.3 | 1.0 | 26.44 | 6.13 | A | N | W |
| 20 | 1 | -3.85 | 6.94 | A | A | A |
| 23.0 | 0.7 | 10.58 | 5.69 | A | A | A |
| 18.9 | 3.8 | -9.13 | 20.67 | A | A | A |
| 25.3 | 2.1 | 21.63 | 9.59 | A | A | A |
| 28.62 | 1.02 | 37.60 | 5.98 | N | N | N |
| 26.92 | 1.95 | 29.42 | 8.69 | A | N | W |
| 41 | 4 | 97.12 | 10.88 | N | N | N |
| 34.81 | 3.34 | 67.36 | 10.73 | N | N | N |
| 21.6 | 1.8 | 3.85 | 9.62 | A | A | A |
| 32.1 | 4.0 | 54.33 | 13.36 | N | N | N |
| 26.93 | 0.47 | 29.47 | 5.11 | A | N | W |
| 18 | 1 | -13.46 | 7.35 | A | A | A |
| 30.8 | 2.6 | 48.08 | 9.71 | N | N | N |
| 18.63 | 0.99 | -10.43 | 7.17 | A | A | A |
| 19.91 | 1.62 | -4.28 | 9.45 | A | A | A |
| 20.68 | 1 | -0.58 | 6.82 | A | A | A |
| 100 | 5.50 | 380.77 | 7.31 | N | N | N |
| 23.5 | 2.1 | 12.98 | 10.15 | A | A | A |
| 17.698 | 0.969 | -14.91 | 7.29 | A | A | A |
| 20.7 | 1.1 | -0.48 | 7.17 | A | A | A |
| 21.0 | 1.02 | 0.96 | 6.83 | A | A | A |
| 27.993 | 1.94595 | 34.58 | 8.45 | N | N | N |
| 21.6 | 1.4 | 3.85 | 8.07 | A | A | A |
| 23.490 | 1.454 | 12.93 | 7.84 | A | A | A |
| 14.5 | 1.74 | -30.29 | 12.93 | N | A | N |
| 23.85 | 1.05 | 14.66 | 6.52 | A | A | A |
| 17.3 | 1.6 | -16.83 | 10.42 | A | A | A |
| 25.0 | 2.7 | 20.19 | 11.82 | A | A | A |

TABLE 47. Evaluation Results for Pb-214 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

2

13

14

17

18

21

22

23

27

28

30

33

34

36

42

46

47

48

51

52

53

54

57

58

63

66

68

69

72

74

75

77

79

81

85

86

87

89

93

96

99

102

110

112

118

124

125

128

129

132

142

144

146

149

151

153

154

158

161

162

TABLE 47. Evaluation Results for Pb-214 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

| Reported value [Bq/kg ] 13.89 | Reported uncertainty (k = 1) [Bq/kg ] 0.55 | Relative bias [%] -33.22 | P-Test [%] 6.23 | Trueness evaluation N | Precision evaluation N | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 15.7 | 1.2 | -24.52 | 9.03 | A | N | W |
| 21.4 | 0.9 | 2.88 | 6.39 | A | A | A |
| 20 | 1 | -3.85 | 6.94 | A | A | A |
| 33 | 2 | 58.65 | 7.74 | N | N | N |
| 20.02 | 1.91 | -3.75 | 10.68 | A | A | A |
| 20.6 | 2 | -0.96 | 10.83 | A | A | A |
| 17.5 | 1.2 | -15.87 | 8.37 | A | A | A |
| 35.6 | 4.6 | 71.15 | 13.79 | N | N | N |
| 22.7 | 1.3 | 9.13 | 7.48 | A | A | A |
| 22.0 | 1.0 | 5.77 | 6.62 | A | A | A |
| 20.95 | 0.97 | 0.72 | 6.67 | A | A | A |
| 23.1 | 4.1 | 11.06 | 18.39 | A | A | A |
| 22.20 | 1.27 | 6.73 | 7.47 | A | A | A |
| 25.89 | 1.20 | 24.47 | 6.68 | A | N | W |
| 18.25 | 1.71 | -12.26 | 10.53 | A | A | A |
| 14.9 | 0.9 | -28.37 | 7.72 | A | N | W |
| 17.7 | 1.42 | -14.90 | 9.35 | A | A | A |
| 22.70 | 1.84 | 9.13 | 9.42 | A | A | A |
| 17.40 | 0.96 | -16.35 | 7.32 | A | A | A |
| 1.25 | 0.02 | -93.99 | 5.07 | N | N | N |
| 25.10 | 1.12 | 20.67 | 6.56 | A | N | W |
| 19.12 | 1.65 | -8.08 | 9.88 | A | A | A |
| 23.8 | 1.8 | 14.42 | 8.96 | A | A | A |
| 18.0 | 1.0 | -13.46 | 7.35 | A | A | A |
| 36.8 | 2.8 | 76.92 | 9.00 | N | N | N |
| 18.6 | 1.9 | -10.58 | 11.29 | A | A | A |
| 25.3 | 2.4 | 21.63 | 10.63 | A | A | A |
| 20.4 | 0.4 | -1.92 | 5.19 | A | A | A |
| 20.5 | 1.5 | -1.44 | 8.76 | A | A | A |
| 24.06 | 2.15 | 15.67 | 10.15 | A | A | A |
| 22.3 | 1 | 7.21 | 6.57 | A | A | A |
| 19.05 | 1.91 | -8.41 | 11.12 | A | A | A |
| 20.2 | 1.4 | -2.88 | 8.43 | A | A | A |
| 20.6 | 1.1 | -0.96 | 7.19 | A | A | A |
| 23.0 | 1.0 | 10.58 | 6.48 | A | A | A |
| 19.4 | 1.99 | -6.73 | 11.33 | A | A | A |
| 26 | 3.48 | 25.00 | 14.22 | A | A | A |
| 18.54 | 2.039 | -10.87 | 12.00 | A | A | A |
| 10.2 | 2.9 | -50.96 | 28.83 | N | A | N |
| 20.2 | 1.1 | -2.88 | 7.26 | A | A | A |
| 26.35 | 1.86 | 26.68 | 8.54 | A | N | W |
| 19 | 1 | -8.65 | 7.13 | A | A | A |
| 19.65 | 1.84 | -5.53 | 10.53 | A | A | A |
| 25.2 | 3.3 | 21.15 | 13.95 | A | A | A |
| 22.5 | 2.1 | 8.17 | 10.50 | A | A | A |
| 19.9 | 2.1 | -4.33 | 11.60 | A | A | A |
| 23.16 | 3.57 | 11.35 | 16.15 | A | A | A |
| 37.42 | 0.04 | 79.90 | 4.81 | N | N | N |
| 18.0 | 1.0 | -13.46 | 7.35 | A | A | A |
| 21.0 | 2.7 | 0.96 | 13.73 | A | A | A |
| 28.06 | 2.42 | 34.90 | 9.87 | N | N | N |
| 24.73 | 1.00 | 18.89 | 6.28 | A | N | W |
| 67.3531 |  | 223.81 |  | N |  | N |

Labcode

165

170

171

174

176

180

182

183

184

186

190

199

204

205

206

209

214

218

222

223

226

230

233

235

240

241

245

246

247

248

251

259

260

268

270

271

275

279

281

294

298

301

305

306

307

314

315

322

323

328

329

341

342

343

Activity concentration  [ Bq/kg ]

20

25

Reported Results for Ra-226 determination in Sample 4

TABLE 48. Target values

| Parameter | Value |
| --- | --- |
| Activity Concentration [ Bq/kg ] | 20.8 |
| Uncertainty [ Bq/kg ] | 1.0 |
| Maximum Acceptable Relative Bias, [%] | 30 |

Not Accepted 10%

Warning 3%

Accepted 22%

Not Reported 65%

30

FIGURE 45. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

10

15

172

92

249

165

42

344

75

38

309

197

6

126

183

189

133

240

209

18

281

66

233

11

55

199

210

27

201

17

247

46

73

174

71

43

51

264

128

47

260

234

131

29

302

178

181

135

245

146

217

89

34

202

251

293

173

86

67

231

119

186

48

107

222

10

271

63

232

206

134

227

263

122

156

158

130

123

246

273

168

12

155

307

224

230

162

330

171

200

272

69

112

157

104

176

54

39

184

259

140

147

182

163

102

110

149

145

33

301

205

226

108

236

323

Laboratory code

FIGURE 46. Reported results and evaluation parameters for Ra-226. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

| Reported value [Bq/kg ] 16.8 | Reported uncertainty (k = 1) [Bq/kg ] 0.8 | Relative bias [%] -19.23 | P-Test [%] 6.77 | Trueness evaluation A | Precision evaluation N | Final Score W |
| --- | --- | --- | --- | --- | --- | --- |
| 23 | 3 | 10.58 | 13.90 | A | A | A |
| 19.1 | 1.3 | -8.17 | 8.33 | A | A | A |
| 25 | 3 | 20.19 | 12.93 | A | A | A |
| 19.9 | 1.8 | -4.33 | 10.24 | A | A | A |
| 18.71 | 1.28 | -10.05 | 8.36 | A | A | A |
| 19.6 | 1.0 | -5.77 | 7.01 | A | A | A |
| 21 | 5 | 0.96 | 24.29 | A | A | A |
| 51.3 | 5.3 | 146.63 | 11.40 | N | N | N |
| 21.7 | 4.7 | 4.33 | 22.19 | A | A | A |
| 15 | 1.2 | -27.88 | 9.33 | A | N | W |
| 35.5 | 1.6 | 70.67 | 6.59 | N | N | N |
| 13.4 | 3.44 | -35.58 | 26.12 | N | A | N |
| 20.3 | 1.9 | -2.40 | 10.52 | A | A | A |
| 20 | 1 | -3.85 | 6.94 | A | A | A |
| 20.5 | 0.6 | -1.44 | 5.63 | A | A | A |
| 22.7 | 1.6 | 9.13 | 8.53 | A | A | A |
| 20.3 | 2.5 | -2.40 | 13.22 | A | A | A |
| 34.3 | 5.8 | 64.90 | 17.58 | N | N | N |
| 19.4 | 1.3 | -6.73 | 8.25 | A | A | A |
| 23.0 | 1.8 | 10.58 | 9.18 | A | A | A |
| 18.91 | 0.6 | -9.09 | 5.76 | A | A | A |
| 22.34 | 1.14 | 7.40 | 7.01 | A | A | A |
| 28.46 | 2.01 | 36.83 | 8.54 | N | N | N |
| 20 | 5 | -3.85 | 25.46 | A | A | A |
| 20.0 | 1.5 | -3.85 | 8.91 | A | A | A |
| 14.76 | 0.87 | -29.04 | 7.61 | A | N | W |
| 22.3 | 1.5 | 7.21 | 8.27 | A | A | A |
| 21.63 | 1.84 | 3.99 | 9.77 | A | A | A |
| 6.16 | 0.54 | -70.38 | 10.00 | N | N | N |
| 43.92 | 4.27 | 111.15 | 10.85 | N | N | N |
| 32.20 | 3.38 | 54.81 | 11.55 | N | N | N |
| 22.7 | 2.8 | 9.13 | 13.24 | A | A | A |
| 61.1 | 6.9 | 193.75 | 12.27 | N | N | N |
| 49.58 | 0.89 | 138.37 | 5.13 | N | N | N |
| 30 | 10 | 44.23 | 33.68 | N | N | N |
| 22.5 | 3 | 8.17 | 14.17 | A | A | A |
| 23.6 | 3.8 | 13.46 | 16.80 | A | A | A |
| 24.7 | 1.0 | 18.75 | 6.29 | A | N | W |
| 17.2 | 0.5 | -17.31 | 5.62 | A | N | W |
| 20.46 | 1.08 | -1.63 | 7.14 | A | A | A |
| 24.0 | 3.0 | 15.38 | 13.39 | A | A | A |
| 20.85 | 0.96 | 0.24 | 6.66 | A | A | A |
| 17.25 | 1.15 | -17.07 | 8.22 | A | A | A |
| 23.2 | 3.7 | 11.54 | 16.66 | A | A | A |
| 21.5 | 1.4 | 3.37 | 8.09 | A | A | A |
| 39.9 | 5.7 | 91.83 | 15.07 | N | N | N |
| 51.3 | 6.0 | 146.63 | 12.65 | N | N | N |
| 21.53 | 1.12 | 3.51 | 7.08 | A | A | A |
| 40.83 | 6.67 | 96.30 | 17.03 | N | N | N |
| 51.241 | 11.8895 | 146.35 | 23.70 | N | N | N |
| 25.2 | 3.7 | 21.15 | 15.45 | A | A | A |
| 23.7 | 3.5 | 13.94 | 15.53 | A | A | A |
| 30.5 | 1.8 | 46.63 | 7.61 | N | N | N |
| 23.85 | 1.05 | 14.66 | 6.52 | A | A | A |
| 26.2 | 1.9 | 25.96 | 8.70 | A | N | W |
| 41.3 | 6.1 | 98.56 | 15.53 | N | N | N |
| 13.27 | 0.83 | -36.20 | 7.89 | N | N | N |
| 25 | 3 | 20.19 | 12.93 | A | A | A |
| 27 | 2.8 | 29.81 | 11.43 | A | N | W |

TABLE 49. Evaluation Results for Ra-226 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

6

10

11

12

17

18

27

29

33

34

38

39

42

43

46

47

48

51

54

55

63

66

67

69

71

73

75

86

89

92

102

104

107

108

110

112

119

122

123

126

128

130

131

133

134

135

140

145

146

147

149

155

156

157

158

162

163

165

168

171

TABLE 49. Evaluation Results for Ra-226 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

| Reported value [Bq/kg ] 1.71 | Reported uncertainty (k = 1) [Bq/kg ] 0.328 | Relative bias [%] -91.78 | P-Test [%] 19.77 | Trueness evaluation N | Precision evaluation N | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 22.11 | 2.7 | 6.30 | 13.12 | A | A | A |
| 20 | 2 | -3.85 | 11.10 | A | A | A |
| 34 | 2 | 63.46 | 7.60 | N | N | N |
| 21.1 | 1.3 | 1.44 | 7.81 | A | A | A |
| 21.2 | 1.6 | 1.92 | 8.95 | A | A | A |
| 41.3 | 9 | 98.56 | 22.32 | N | N | N |
| 17.2 | 1.3 | -17.31 | 8.96 | A | A | A |
| 36.0 | 4.3 | 73.08 | 12.88 | N | N | N |
| 22.7 | 1.3 | 9.13 | 7.48 | A | A | A |
| 17.25 | 0.92 | -17.07 | 7.18 | A | A | A |
| 16.7448 | 2.2643 | -19.50 | 14.35 | A | A | A |
| 19.42 | 1.56 | -6.63 | 9.36 | A | A | A |
| 27.07 | 2.76 | 30.14 | 11.27 | N | N | N |
| 19.7 | 1.3 | -5.29 | 8.16 | A | A | A |
| 21.77 | 2.10 | 4.66 | 10.78 | A | A | A |
| 56.91 | 8.10 | 173.61 | 15.02 | N | N | N |
| 23.08 | 1.27 | 10.96 | 7.31 | A | A | A |
| 18.25 | 1.52 | -12.26 | 9.62 | A | A | A |
| 19.48 | 0.85 | -6.35 | 6.49 | A | A | A |
| 21.6 | 5.6 | 3.85 | 26.37 | A | A | A |
| 22.70 | 1.84 | 9.13 | 9.42 | A | A | A |
| 25.5 | 1.8 | 22.60 | 8.54 | A | N | W |
| 60.43 | 0.2 | 190.53 | 4.82 | N | N | N |
| 23.3 | 0.8 | 12.02 | 5.91 | A | A | A |
| 26.02 | 1.21 | 25.10 | 6.69 | A | N | W |
| 22.36 | 1.32 | 7.50 | 7.61 | A | A | A |
| 23 | 3 | 10.58 | 13.90 | A | A | A |
| 19.06 | 4.08 | -8.37 | 21.94 | A | A | A |
| 20.71 | 1.01 | -0.43 | 6.85 | A | A | A |
| 71.4 | 8.51 | 243.27 | 12.85 | N | N | N |
| 17.7 | 1.0 | -14.90 | 7.42 | A | A | A |
| 21.5 | 3.1 | 3.37 | 15.20 | A | A | A |
| 24.7 | 2.0 | 18.75 | 9.42 | A | A | A |
| 19.9 | 0.9 | -4.33 | 6.60 | A | A | A |
| 11 | 0.9 | -47.12 | 9.49 | N | N | N |
| 21.8 | 2.82 | 4.81 | 13.80 | A | A | A |
| 39.2 | 2 | 88.46 | 7.01 | N | N | N |
| 20.50 | 2.05 | -1.44 | 11.10 | A | A | A |
| 23.3 | 1.8 | 12.02 | 9.10 | A | A | A |
| 20.36 | 1.575 | -2.12 | 9.11 | A | A | A |
| 23.0 | 1.0 | 10.58 | 6.48 | A | A | A |
| 27.1 | 1.8 | 30.29 | 8.20 | N | N | N |
| 24.9 | 1.1 | 19.71 | 6.53 | A | N | W |
| 18.76 | 2.251 | -9.81 | 12.93 | A | A | A |
| 22.1 | 1.1 | 6.25 | 6.92 | A | A | A |
| 53.21 | 7.17 | 155.82 | 14.31 | N | N | N |
| 21 | 3 | 0.96 | 15.07 | A | A | A |
| 25.2 | 3.3 | 21.15 | 13.95 | A | A | A |
| 16.72 | 2.31 | -19.62 | 14.63 | A | A | A |
| 211.92 | 24.17 | 918.85 | 12.38 | N | N | N |
| 27.0 | 2.7 | 29.81 | 11.10 | A | N | W |
| 13.6 | 1.8 | -34.62 | 14.08 | N | A | N |

Labcode

172

173

174

176

178

181

182

183

184

186

189

197

199

200

201

202

205

206

209

210

217

222

224

226

227

230

231

232

233

234

236

240

245

246

247

249

251

259

260

263

264

271

272

273

281

293

301

302

307

309

323

330

344

Activity concentration  [ Bq/kg ]

30

40

Reported Results for Ac-228 determination in Sample 4

TABLE 50. Target values

| Parameter | Value |
| --- | --- |
| Activity Concentration [ Bq/kg ] | 32.9 |
| Uncertainty [ Bq/kg ] | 1.9 |
| Maximum Acceptable Relative Bias, [%] | 30 |

Accepted 38%

Not Accepted 3%

Warning 2%

Not Reported 57%

50

FIGURE 47. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

20

335

102

223

305

125

142

25

230

124

85

170

112

183

140

306

199

204

21

75

232

174

133

182

271

239

63

126

299

17

176

46

259

209

5

66

73

58

33

277

342

233

53

279

180

161

218

173

260

54

47

244

130

12

155

18

27

90

77

2

34

22

122

222

30

298

303

42

10

328

52

189

158

281

141

275

23

168

156

151

134

270

74

228

268

307

248

251

235

107

28

205

245

247

314

132

146

185

29

71

302

69

166

190

206

221

144

258

296

19

14

81

186

240

153

200

118

184

96

214

57

294

224

149

13

344

165

87

315

31

301

110

99

89

68

323

241

93

Laboratory code

FIGURE 48. Reported results and evaluation parameters for Ac-228. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

| Reported value [Bq/kg ] 33.45 | Reported uncertainty (k = 1) [Bq/kg ] 2.36 | Relative bias [%] 1.67 | P-Test [%] 9.12 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 32.30 | 1.85 | -1.82 | 8.13 | A | A | A |
| 34 | 1.5 | 3.34 | 7.27 | A | A | A |
| 33 | 2 | 0.30 | 8.37 | A | A | A |
| 41.9 | 0.7 | 27.36 | 6.01 | A | N | W |
| 37.6 | 2.3 | 14.29 | 8.41 | A | A | A |
| 31.8 | 2.4 | -3.34 | 9.50 | A | A | A |
| 33.11 | 2.82 | 0.64 | 10.29 | A | A | A |
| 37.5 | 3.8 | 13.98 | 11.66 | A | A | A |
| 30.8 | 2.2 | -6.38 | 9.19 | A | A | A |
| 33.78 | 1.34 | 2.67 | 7.01 | A | A | A |
| 34.4 | 1.0 | 4.56 | 6.47 | A | A | A |
| 29.09 | 3.0 | -11.58 | 11.82 | A | A | A |
| 33.3 | 1.2 | 1.22 | 6.81 | A | A | A |
| 35.1 | 3.2 | 6.69 | 10.79 | A | A | A |
| 36 | 3 | 9.42 | 10.14 | A | A | A |
| 33.9 | 3.0 | 3.04 | 10.57 | A | A | A |
| 44.8 | 3.1 | 36.17 | 9.01 | N | N | N |
| 32.5 | 2.5 | -1.22 | 9.62 | A | A | A |
| 33.6 | 2.4 | 2.13 | 9.19 | A | A | A |
| 34.0 | 3.23 | 3.34 | 11.12 | A | A | A |
| 32 | 1 | -2.74 | 6.57 | A | A | A |
| 33 | 1.4 | 0.30 | 7.17 | A | A | A |
| 34.1 | 3.1 | 3.65 | 10.77 | A | A | A |
| 32.6 | 1.8 | -0.91 | 7.99 | A | A | A |
| 32.9 | 0.9 | 0.00 | 6.39 | A | A | A |
| 39.7 | 3.9 | 20.67 | 11.40 | A | A | A |
| 32.5 | 4.7 | -1.22 | 15.57 | A | A | A |
| 31.6 | 0.8 | -3.95 | 6.31 | A | A | A |
| 32.41 | 1.42 | -1.49 | 7.25 | A | A | A |
| 55.2 | 3.1 | 67.78 | 8.06 | N | N | N |
| 36.00 | 2.42 | 9.42 | 8.86 | A | A | A |
| 36 | 2 | 9.42 | 8.01 | A | A | A |
| 32.5 | 1.2 | -1.22 | 6.85 | A | A | A |
| 34.68 | 1.11 | 5.41 | 6.60 | A | A | A |
| 30.90 | 1.02 | -6.08 | 6.65 | A | A | A |
| 33.4 | 1.3 | 1.52 | 6.96 | A | A | A |
| 37.6 | 1.6 | 14.29 | 7.17 | A | A | A |
| 29.4 | 5.9 | -10.64 | 20.88 | A | A | A |
| 42.54 | 1.51 | 29.30 | 6.78 | A | N | W |
| 50.39 | 2.5 | 53.16 | 7.61 | N | N | N |
| 33.34 | 2.884 | 1.34 | 10.40 | A | A | A |
| 77 | 5 | 134.04 | 8.69 | N | N | N |
| 39.56 | 3.15 | 20.24 | 9.84 | A | A | A |
| 47.5 | 4.1 | 44.38 | 10.39 | N | N | N |
| 27.19 | 1.31 | -17.36 | 7.52 | A | A | A |
| 35.1 | 1.0 | 6.69 | 6.44 | A | A | A |
| 47.19 | 1.15 | 43.43 | 6.27 | N | N | N |
| 30 | 3 | -8.81 | 11.55 | A | A | A |
| 38.7 | 1.4 | 17.63 | 6.81 | A | N | W |
| 33.8 | 1.5 | 2.74 | 7.28 | A | A | A |
| 29.25 | 1.43 | -11.09 | 7.57 | A | A | A |
| 28.3 | 2.02 | -13.98 | 9.18 | A | A | A |
| 31.7 | 0.9 | -3.65 | 6.44 | A | A | A |
| 33.0 | 2.0 | 0.30 | 8.37 | A | A | A |
| 35.9 | 3.2 | 9.12 | 10.62 | A | A | A |
| 31.00 | 1.93 | -5.78 | 8.49 | A | A | A |
| 34.6 | 1.4 | 5.17 | 7.05 | A | A | A |
| 30.5 | 2.0 | -7.29 | 8.74 | A | A | A |
| 34.26 | 2.21 | 4.13 | 8.66 | A | A | A |

TABLE 51. Evaluation Results for Ac-228 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

2

5

10

12

13

14

17

18

19

21

22

23

25

27

28

29

30

31

33

34

42

46

47

52

53

54

57

58

63

66

68

69

71

73

74

75

77

81

85

87

89

90

93

96

99

102

107

110

112

118

122

124

125

126

130

132

133

134

140

141

| Reported value [Bq/kg ] 28.734 | Reported uncertainty (k = 1) [Bq/kg ] 3.534 | Relative bias [%] -12.66 | P-Test [%] 13.59 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 36.8 | 2.0 | 11.85 | 7.93 | A | A | A |
| 35.9 | 1.7 | 9.12 | 7.47 | A | A | A |
| 41.609 | 3.0023 | 26.47 | 9.24 | A | N | W |
| 34.6 | 2.8 | 5.17 | 9.94 | A | A | A |
| 38.282 | 2.390 | 16.36 | 8.50 | A | A | A |
| 33.1 | 2.4 | 0.61 | 9.27 | A | A | A |
| 34.5 | 3.0 | 4.86 | 10.44 | A | A | A |
| 34.20 | 1.42 | 3.95 | 7.11 | A | A | A |
| 32.7 | 2.9 | -0.61 | 10.58 | A | A | A |
| 42.12 | 1.02 | 28.02 | 6.26 | A | N | W |
| 36 | 4.7 | 9.42 | 14.28 | A | A | A |
| 34.5 | 2.1 | 4.86 | 8.39 | A | A | A |
| 29.5 | 1.6 | -10.33 | 7.92 | A | A | A |
| 32.85 | 1.32 | -0.15 | 7.04 | A | A | A |
| 31 | 2 | -5.78 | 8.66 | A | A | A |
| 32 | 2 | -2.74 | 8.51 | A | A | A |
| 32.65 | 1.85 | -0.76 | 8.09 | A | A | A |
| 31.3 | 3 | -4.86 | 11.19 | A | A | A |
| 30.0 | 1.7 | -8.81 | 8.09 | A | A | A |
| 39.3 | 4.7 | 19.45 | 13.28 | A | A | A |
| 35.90 | 4.67 | 9.12 | 14.23 | A | A | A |
| 38 | 3 | 15.50 | 9.78 | A | A | A |
| 34.1 | 3 | 3.65 | 10.52 | A | A | A |
| 36.2 | 1.4 | 10.03 | 6.95 | A | A | A |
| 30.6 | 1.56 | -6.99 | 7.70 | A | A | A |
| 38.6 | 4.4 | 17.33 | 12.78 | A | A | A |
| 30.6 | 5.8 | -6.99 | 19.81 | A | A | A |
| 35.12 | 2.21 | 6.75 | 8.54 | A | A | A |
| 36.31 | 2.23 | 10.36 | 8.43 | A | A | A |
| 32.23 | 2.30 | -2.04 | 9.18 | A | A | A |
| 39.6 | 2.2 | 20.36 | 8.01 | A | A | A |
| 32.7 | 2.51 | -0.61 | 9.61 | A | A | A |
| 36.70 | 1.30 | 11.55 | 6.77 | A | A | A |
| 33.80 | 2.13 | 2.74 | 8.55 | A | A | A |
| 27.30 | 0.98 | -17.02 | 6.80 | A | A | A |
| 41.3 | 0.9 | 25.53 | 6.17 | A | N | W |
| 34.7 | 1.9 | 5.47 | 7.96 | A | A | A |
| 29.2 | 1.43 | -11.25 | 7.57 | A | A | A |
| 31 | 2 | -5.78 | 8.66 | A | A | A |
| 32.59 | 3.12 | -0.94 | 11.18 | A | A | A |
| 35.1 | 3.2 | 6.69 | 10.79 | A | A | A |
| 31.46 | 1.42 | -4.38 | 7.33 | A | A | A |
| 38.2 | 2.2 | 16.11 | 8.16 | A | A | A |
| 65.3 | 5.2 | 98.48 | 9.84 | N | N | N |
| 33 | 2.42 | 0.30 | 9.33 | A | A | A |
| 35.5 | 3.1 | 7.90 | 10.47 | A | A | A |
| 35.6 | 1.5 | 8.21 | 7.15 | A | A | A |
| 35.0 | 2.0 | 6.38 | 8.12 | A | A | A |
| 35.07 | 2.35 | 6.60 | 8.85 | A | A | A |
| 36.95 | 1.95 | 12.31 | 7.82 | A | A | A |
| 32.1 | 4 | -2.43 | 13.73 | A | A | A |
| 32.89 | 3.29 | -0.03 | 11.55 | A | A | A |
| 34.8 | 2.5 | 5.78 | 9.22 | A | A | A |
| 34.6 | 1.7 | 5.17 | 7.58 | A | A | A |
| 31.4 | 1.7 | -4.56 | 7.92 | A | A | A |
| 34.32 | 2.83 | 4.32 | 10.07 | A | A | A |
| 32.52 | 1.80 | -1.16 | 8.00 | A | A | A |
| 32.62 | 2.36 | -0.85 | 9.26 | A | A | A |
| 34.24 | 3.424 | 4.07 | 11.55 | A | A | A |

TABLE 51. Evaluation Results for Ac-228 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

Labcode

142

144

146

149

151

153

155

156

158

161

165

166

168

170

173

174

176

180

182

183

184

185

186

189

190

199

200

204

205

206

209

214

218

221

222

223

224

228

230

232

233

235

239

240

241

244

245

247

248

251

258

259

260

268

270

271

275

277

279

281

TABLE 51. Evaluation Results for Ac-228 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

| Reported value [Bq/kg ] 40.7 | Reported uncertainty (k = 1) [Bq/kg ] 3.1 | Relative bias [%] 23.71 | P-Test [%] 9.56 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 36.98 | 2.27 | 12.40 | 8.43 | A | A | A |
| 33.9 | 1.7 | 3.04 | 7.65 | A | A | A |
| 31.7 | 1.4 | -3.65 | 7.27 | A | A | A |
| 46.67 | 3.31 | 41.85 | 9.15 | N | N | N |
| 36 | 3 | 9.42 | 10.14 | A | A | A |
| 33.94 | 6.13 | 3.16 | 18.96 | A | A | A |
| 28 | 2 | -14.89 | 9.19 | A | A | A |
| 30.57 | 2.78 | -7.08 | 10.77 | A | A | A |
| 34.8 | 1.8 | 5.78 | 7.75 | A | A | A |
| 35.7 | 3.4 | 8.51 | 11.14 | A | A | A |
| 43.4 | 4.2 | 31.91 | 11.27 | N | N | N |
| 59.91 | 0.04 | 82.10 | 5.78 | N | N | N |
| 34.0 | 1.7 | 3.34 | 7.64 | A | A | A |
| 15.15 | 1.21 | -53.95 | 9.86 | N | N | N |
| 32.53 | 0.77 | -1.12 | 6.24 | A | A | A |
| 41.9 | 2.2 | 27.36 | 7.81 | A | N | W |

Labcode

294

296

298

299

301

302

303

305

306

307

314

315

323

328

335

342

344

Activity concentration  [ Bq/kg ]

25

Reported Results for Th-234 determination in Sample 4

TABLE 52. Target values

| Parameter | Value |
| --- | --- |
| Activity Concentration [ Bq/kg ] | 23.0 |
| Uncertainty [ Bq/kg ] | 1.5 |
| Maximum Acceptable Relative Bias, [%] | 30 |

Warning 1%

Not Accepted 2%

Accepted 9%

Not Reported 88%

30

35

FIGURE 49. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

10

15

20

275

183

189

174

214

171

259

89

29

112

43

158

206

180

307

85

144

23

63

135

133

155

185

166

173

131

233

66

305

102

21

263

301

184

228

298

154

Laboratory code

FIGURE 50. Reported results and evaluation parameters for Th-234. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

TABLE 53. Evaluation Results for Th-234 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

| Reported value [Bq/kg ] 29.8 | Reported uncertainty (k = 1) [Bq/kg ] 2.7 | Relative bias [%] 29.57 | P-Test [%] 11.16 | Trueness evaluation A | Precision evaluation N | Final Score W |
| --- | --- | --- | --- | --- | --- | --- |
| 24.3 | 3.9 | 5.65 | 17.32 | A | A | A |
| 22 | 5 | -4.35 | 23.64 | A | A | A |
| 22.9 | 3.4 | -0.43 | 16.22 | A | A | A |
| 24.7 | 1.5 | 7.39 | 8.91 | A | A | A |
| 26.794 | 1.5 | 16.50 | 8.59 | A | A | A |
| 24.3 | 4.9 | 5.65 | 21.19 | A | A | A |
| 21.63 | 1.84 | -5.96 | 10.72 | A | A | A |
| 28.1 | 8.5 | 22.17 | 30.94 | A | N | W |
| 22 | 5 | -4.35 | 23.64 | A | A | A |
| 26.74 | 2.70 | 16.26 | 12.02 | A | A | A |
| 25.02 | 1.56 | 8.78 | 9.02 | A | A | A |
| 25 | 3 | 8.70 | 13.66 | A | A | A |
| 24.3 | 3.0 | 5.65 | 13.96 | A | A | A |
| 65.5 | 5.2 | 184.78 | 10.27 | N | N | N |
| 25.8 | 2.3 | 12.17 | 11.05 | A | A | A |
| 23.30 | 1.37 | 1.30 | 8.78 | A | A | A |
| 25.8 | 4.8 | 12.17 | 19.71 | A | A | A |
| 20.9 | 5.2 | -9.13 | 25.72 | A | A | A |
| 26.122 | 2.45 | 13.57 | 11.42 | A | A | A |
| 19 | 3 | -17.39 | 17.08 | A | A | A |
| 24.14 | 6.44 | 4.96 | 27.46 | A | A | A |
| 17.1 | 2.0 | -25.65 | 13.39 | A | A | A |
| 34.7 | 4.2 | 50.87 | 13.75 | N | N | N |
| 25.80 | 4.77 | 12.17 | 19.60 | A | A | A |
| 18 | 6 | -21.74 | 33.97 | A | N | W |
| 23.84 | 5.62 | 3.65 | 24.46 | A | A | A |
| 20.0 | 2.5 | -13.04 | 14.10 | A | A | A |
| 36.9 | 3.4 | 60.43 | 11.29 | N | N | N |
| 26.78 | 6.77 | 16.43 | 26.11 | A | A | A |
| 21.4 | 2 | -6.96 | 11.40 | A | A | A |
| 30.1 | 4.4 | 30.87 | 16.01 | N | A | N |
| 17.02 | 4.37 | -26.00 | 26.49 | A | A | A |
| 41.9 | 5.0 | 82.17 | 13.60 | N | N | N |
| 33.06 | 5.69 | 43.74 | 18.41 | N | A | N |
| 28 | 4 | 21.74 | 15.70 | A | A | A |
| 24.2 | 2.9 | 5.22 | 13.64 | A | A | A |

Labcode

21

23

29

43

63

66

85

89

102

112

131

133

135

144

154

155

158

166

171

173

174

180

183

184

185

189

206

214

228

233

259

263

275

298

301

305

307

Activity concentration  [ Bq/kg ]

1.5

Reported Results for U-235 determination in Sample 4

TABLE 54. Target values

| Parameter | Value |
| --- | --- |
| Activity Concentration [ Bq/kg ] | 1.22 |
| Uncertainty [ Bq/kg ] | 0.15 |
| Maximum Acceptable Relative Bias, [%] | 30 |

Warning 1%

Not Accepted 5%

Accepted 2%

Not Reported 92%

2.0

FIGURE 51. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

0.5

1.0

183

140

135

247

89

155

162

240

47

246

123

144

29

46

67

90

85

66

233

68

102

281

275

33

132

241

215

Laboratory code

FIGURE 52. Reported results and evaluation parameters for U-235. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

TABLE 55. Evaluation Results for U-235 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg by gamma

| Reported value [Bq/kg ] 1.5 | Reported uncertainty (k = 1) [Bq/kg ] 0.4 | Relative bias [%] 22.95 | P-Test [%] 29.36 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 2.8 | 1.2 | 129.51 | 44.59 | N | N | N |
| 1.9 | 0.1 | 55.74 | 13.37 | N | N | N |
| 1.4 | 0.1 | 14.75 | 14.22 | A | A | A |
| 2.05 | 0.7 | 68.03 | 36.29 | N | N | N |
| 1.91 | 0.81 | 56.56 | 44.15 | N | N | N |
| 2.28 | 0.38 | 86.89 | 20.71 | N | N | N |
| 2.03 | 0.41 | 66.39 | 23.65 | N | N | N |
| 1.08 | 0.16 | -11.48 | 19.25 | A | A | A |
| 2.01 | 0.38 | 64.75 | 22.55 | N | N | N |
| 2.38 | 0.30 | 95.08 | 17.61 | N | N | N |
| 1.41 | 0.57 | 15.57 | 42.25 | A | N | W |
| 3.6 | 0.3 | 195.08 | 14.85 | N | N | N |
| 0.87 | 0.10 | -28.69 | 16.83 | A | A | A |
| 0.8 | 0.9 | -34.43 | 113.17 | N | N | N |
| 1.43 | 0.60 | 17.21 | 43.72 | A | N | W |
| 1.32 | 0.27 | 8.20 | 23.87 | A | A | A |
| 1.35 | 0.41 | 10.66 | 32.76 | A | N | W |
| 0.71 | 0.10 | -41.80 | 18.70 | N | A | N |
| 7.88 | 0.21 | 545.90 | 12.58 | N | N | N |
| 2.2 | 0.52 | 80.33 | 26.64 | N | N | N |
| 1.4 | 0.2 | 14.75 | 18.85 | A | A | A |
| 4.0 | 0.3 | 227.87 | 14.40 | N | N | N |
| 1.40 | 0.48 | 14.75 | 36.42 | A | N | W |
| 0.96 | 0.1 | -21.31 | 16.11 | A | A | A |
| 2.63 | 0.7 | 115.57 | 29.32 | N | N | N |
| 2.61 | 0.365 | 113.93 | 18.62 | N | N | N |

Labcode

29

33

46

47

66

67

68

85

89

90

102

123

132

135

140

144

155

162

183

215

233

240

241

246

247

275

281

Reported Results for Cs-134 determination in Sample 5

TABLE 56. Target values

104

Massic activity  [ Bq/sample ]

FIGURE 54. Reported results and evaluation parameters for Cs-134. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

2	3	4	5	6	7	8

68

189

84

333

224

285

33

167

251

96

243

260

142

13

211

273

70

306

210

72

154

199

FIGURE 53. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

144

288

294

176

249

263

110

271

283

172

126

197

309

286

54

198

90

82

97

279

183

307

63

25

228

36

1

58

67

64

206

158

316

261

277

Not Accepted 11%

83

119

153

170

Parameter

Activity per sample [ Bq/sample ] Uncertainty [ Bq/sample ]

Maximum Acceptable Relative Bias, [%]

168

135

131

60

5

334

207

118

29

Warning 8%

140

174

38

145

2

18

4

7

17

304

71

186

141

193

262

268

34

218

173

81

205

Laboratory code

106

264

181

10

6

201

75

20

328

342

12

43

165

156

134

105

300

341

281

55

188

169

Accepted 45%

57

98

46

184

180

223

114

179

122

Not Reported 36%

266

79

130

195

239

291

185

37

317

95

240

222

166

125

258

44

26

115

Value 5.40

0.21

30

35

157

19

275

112

148

161

133

233

39

303

247

73

151

248

42

14

155

85

225

259

298

15

331

104

202

178

190

66

241

69

30

47

11

99

123

40

245

270

296

52

159

27

182

235

204

321

319

177

338

330

287

146

308

282

31

16

89

149

230

93

254

322

102

343

| Reported value [Bq/sample ] 4.9 | Reported uncertainty (k = 1) [Bq/sample ] 0.2 | Relative bias [%] -9.26 | P-Test [%] 5.64 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 5.24 | 0.19 | -2.96 | 5.32 | A | A | A |
| 5.24 | 0.38 | -2.96 | 8.23 | A | A | A |
| 5.16 | 0.40 | -4.44 | 8.67 | A | A | A |
| 5.4 | 0.29 | 0.00 | 6.63 | A | A | A |
| 5.25 | 0.08 | -2.78 | 4.18 | A | A | A |
| 5.4 | 0.3 | 0.00 | 6.78 | A | A | A |
| 6.10 | 0.50 | 12.96 | 9.07 | A | A | A |
| 5.44 | 0.50 | 0.74 | 9.98 | A | A | A |
| 2.9 | 0.04 | -46.30 | 4.13 | N | N | N |
| 5.81 | 0.36 | 7.59 | 7.32 | A | A | A |
| 5.9 | 0.1 | 9.26 | 4.24 | A | A | A |
| 12.01 | 1.31 | 122.41 | 11.58 | N | N | N |
| 5.28 | 0.31 | -2.22 | 7.04 | A | A | A |
| 5.24 | 0.35 | -2.96 | 7.73 | A | A | A |
| 5.68 | 0.40 | 5.19 | 8.04 | A | A | A |
| 5.43 | 0.33 | 0.56 | 7.22 | A | A | A |
| 4.87 | 0.6 | -9.81 | 12.92 | A | A | A |
| 5.66 | 0.44 | 4.81 | 8.69 | A | A | A |
| 6.53 | 0.19 | 20.93 | 4.86 | A | N | W |
| 5.2 | 0.3 | -3.70 | 6.96 | A | A | A |
| 6.1 | 0.6 | 12.96 | 10.58 | A | A | A |
| 8.8 | 0.6 | 62.96 | 7.85 | N | N | N |
| 0.3 | 0.02 | -94.44 | 7.72 | N | N | N |
| 5.33 | 0.24 | -1.30 | 5.95 | A | A | A |
| 5.67 | 0.47 | 5.00 | 9.16 | A | A | A |
| 4.9 | 0.4 | -9.26 | 9.04 | A | A | A |
| 5.59 | 0.64 | 3.52 | 12.09 | A | A | A |
| 5.2 | 0.32 | -3.70 | 7.28 | A | A | A |
| 5.73 | 0.30 | 6.11 | 6.52 | A | A | A |
| 6.17 | 0.253 | 14.26 | 5.65 | A | A | A |
| 5.8 | 0.55 | 7.41 | 10.25 | A | A | A |
| 5.45 | 0.60 | 0.93 | 11.68 | A | A | A |
| 5.63 | 0.42 | 4.26 | 8.41 | A | A | A |
| 5.5 | 0.3 | 1.85 | 6.70 | A | A | A |
| 6.1 | 0.2 | 12.96 | 5.09 | A | A | A |
| 6.3 | 1.0 | 16.67 | 16.34 | A | A | A |
| 4.69 | 0.08 | -13.15 | 4.25 | A | N | W |
| 5.5 | 0.2 | 1.85 | 5.32 | A | A | A |
| 5.5 | 0.2 | 1.85 | 5.32 | A | A | A |
| 4.90 | 0.63 | -9.26 | 13.43 | A | A | A |
| 5.13 | 0.22 | -5.00 | 5.79 | A | A | A |
| 4.87 | 0.12 | -9.81 | 4.60 | A | A | A |
| 4.92 | 0.28 | -8.89 | 6.89 | A | A | A |
| 6.0 | 0.1 | 11.11 | 4.23 | A | N | W |
| 4.92 | 0.7 | -8.89 | 14.75 | A | A | A |
| 0.0581 | 0.0023 | -98.92 | 5.55 | N | N | N |
| 6.05 | 0.19 | 12.04 | 5.00 | A | A | A |
| 3.8 | 0.1 | -29.63 | 4.70 | A | N | W |
| 5.3 | 0.4 | -1.85 | 8.49 | A | A | A |
| 3.94 | 0.13 | -27.04 | 5.10 | A | N | W |
| 5.8 | 0.1 | 7.41 | 4.25 | A | A | A |
| 5.43 | 0.13 | 0.56 | 4.57 | A | A | A |
| 5.54 | 0.085 | 2.59 | 4.18 | A | A | A |
| 5.36 | 0.17 | -0.74 | 5.02 | A | A | A |
| 4.75 | 0.1 | -12.04 | 4.42 | A | N | W |
| 5.04 | 0.09 | -6.67 | 4.28 | A | A | A |
| 0.12 | 0.009 | -97.78 | 8.45 | N | N | N |
| 5.85 | 0.44 | 8.33 | 8.47 | A | A | A |
| 12.68 | 0.6 | 134.81 | 6.12 | N | N | N |

TABLE 57. Evaluation Results for Cs-134 in Sample 5, Rep. Value and Rep. Unc. in Bq/sample by gamma

Labcode

1

2

4

5

6

7

10

11

12

13

14

15

16

17

18

19

20

25

26

27

29

30

31

33

34

35

36

37

38

39

40

42

43

44

46

47

52

54

55

57

58

60

63

64

66

67

68

69

70

71

72

73

75

79

81

82

83

84

85

89

| Reported value [Bq/sample ] 4.74 | Reported uncertainty (k = 1) [Bq/sample ] 0.27 | Relative bias [%] -12.22 | P-Test [%] 6.90 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 20.9 | 0.2 | 287.04 | 4.00 | N | N | N |
| 5.60 | 0.41 | 3.70 | 8.29 | A | A | A |
| 1.0 | 0.01 | -81.48 | 4.02 | N | N | N |
| 4.82 | 0.05 | -10.74 | 4.02 | A | N | W |
| 5.50 | 0.28 | 1.85 | 6.41 | A | A | A |
| 6.14 | 0.10 | 13.70 | 4.22 | A | N | W |
| 80.75 | 1.56 | 1395.37 | 4.34 | N | N | N |
| 5.92 | 0.37 | 9.63 | 7.36 | A | A | A |
| 5.48 | 0.29 | 1.48 | 6.57 | A | A | A |
| 5.39 | 0.24 | -0.19 | 5.91 | A | A | A |
| 4.27 | 0.19 | -20.93 | 5.91 | A | N | W |
| 5.7 | 0.2 | 5.56 | 5.24 | A | A | A |
| 5.52 | 0.57 | 2.22 | 11.03 | A | A | A |
| 5.66 | 0.20 | 4.81 | 5.25 | A | A | A |
| 5.2 | 0.2 | -3.70 | 5.47 | A | A | A |
| 5.07 | 0.2 | -6.11 | 5.54 | A | A | A |
| 5.53 | 0.6 | 2.41 | 11.53 | A | A | A |
| 6.17 | 0.18 | 14.26 | 4.86 | A | N | W |
| 5.61 | 0.38 | 3.89 | 7.81 | A | A | A |
| 4.42 | 0.05 | -18.15 | 4.05 | A | N | W |
| 5.54 | 0.50 | 2.59 | 9.83 | A | A | A |
| 5.13 | 0.34 | -5.00 | 7.68 | A | A | A |
| 5.70 | 0.74 | 5.56 | 13.55 | A | A | A |
| 5.47 | 0.81 | 1.30 | 15.31 | A | A | A |
| 5.1 | 0.4 | -5.56 | 8.75 | A | A | A |
| 5.2 | 0.3 | -3.70 | 6.96 | A | A | A |
| 5.305 | 0.235 | -1.76 | 5.89 | A | A | A |
| 2.053 | 0.192 | -61.98 | 10.13 | N | N | N |
| 4.14 | 0.31 | -23.33 | 8.44 | A | N | W |
| 5.23 | 0.39 | -3.15 | 8.41 | A | A | A |
| 7.79 | 0.33 | 44.26 | 5.75 | N | N | N |
| 5.7 | 0.2 | 5.56 | 5.24 | A | A | A |
| 12.7 | 0.29868 | 135.19 | 4.54 | N | N | N |
| 5.8 | 0.3 | 7.41 | 6.47 | A | A | A |
| 5.071 | 0.225 | -6.09 | 5.90 | A | A | A |
| 4.04 | 0.145 | -25.19 | 5.29 | A | N | W |
| 5.82 | 0.31 | 7.78 | 6.60 | A | A | A |
| 5.47 | 0.55 | 1.30 | 10.78 | A | A | A |
| 5.67 | 0.10 | 5.00 | 4.27 | A | A | A |
| 4.95 | 0.50 | -8.33 | 10.82 | A | A | A |
| 6.41 | 0.13 | 18.70 | 4.39 | A | N | W |
| 5.7 | 0.5 | 5.56 | 9.60 | A | A | A |
| 5.45 | 0.64 | 0.93 | 12.37 | A | A | A |
| 5.6 | 0.6 | 3.70 | 11.40 | A | A | A |
| 0.31 | 0.03 | -94.26 | 10.43 | N | N | N |
| 5.1 | 0.3 | -5.56 | 7.05 | A | A | A |
| 5.5 | 0.5 | 1.85 | 9.89 | A | A | A |
| 5.10 | 0.31 | -5.56 | 7.22 | A | A | A |
| 4.40 | 0.28 | -18.52 | 7.46 | A | A | A |
| 5.35 | 0.13 | -0.93 | 4.59 | A | A | A |
| 5.2 | 0.3 | -3.70 | 6.96 | A | A | A |
| 4.2 | 0.3 | -22.22 | 8.13 | A | N | W |
| 7.46 | 0.37 | 38.15 | 6.30 | N | N | N |
| 5.96 | 0.12 | 10.37 | 4.38 | A | A | A |
| 5.53 | 0.21 | 2.41 | 5.44 | A | A | A |
| 5.52 | 0.11 | 2.22 | 4.37 | A | A | A |
| 5.4 | 1.1 | 0.00 | 20.74 | A | A | A |
| 6.6 | 0.7 | 22.22 | 11.30 | A | A | A |
| 4.84 | 0.25 | -10.37 | 6.47 | A | A | A |

TABLE 57. Evaluation Results for Cs-134 in Sample 5, Rep. Value and Rep. Unc. in Bq/sample by gamma

Labcode

90

93

95

96

97

98

99

102

104

105

106

110

112

114

115

118

119

122

123

125

126

130

131

133

134

135

140

141

142

144

145

146

148

149

151

153

154

155

156

157

158

159

161

165

166

167

168

169

170

172

173

174

176

177

178

179

180

181

182

183

| Reported value [Bq/sample ] 5.50 | Reported uncertainty (k = 1) [Bq/sample ] 1.10 | Relative bias [%] 1.85 | P-Test [%] 20.37 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 5.574 | 0.613 | 3.22 | 11.66 | A | A | A |
| 5.3 | 0.2 | -1.85 | 5.42 | A | A | A |
| 5.5 | 0.3 | 1.85 | 6.70 | A | A | A |
| 0.0667 | 0.0033 | -98.76 | 6.29 | N | N | N |
| 6.0 | 0.3 | 11.11 | 6.33 | A | A | A |
| 5.31 | 0.08 | -1.67 | 4.17 | A | A | A |
| 5.54 | 0.63 | 2.59 | 12.02 | A | A | A |
| 4.4348 | 0.13 | -17.87 | 4.87 | A | N | W |
| 4.7 | 0.2 | -12.96 | 5.76 | A | A | A |
| 4.13 | 0.06 | -23.52 | 4.15 | A | N | W |
| 5.4 | 0.2 | 0.00 | 5.37 | A | A | A |
| 5.95 | 0.34 | 10.19 | 6.91 | A | A | A |
| 6.89 | 0.48 | 27.59 | 7.98 | A | N | W |
| 5.38 | 0.11 | -0.37 | 4.39 | A | A | A |
| 4.934 | 0.091 | -8.63 | 4.30 | A | A | A |
| 5.17 | 0.17 | -4.26 | 5.09 | A | A | A |
| 3.92 | 0.2 | -27.41 | 6.42 | A | N | W |
| 3.530 | 0.08 | -34.63 | 4.50 | N | N | N |
| 5.34 | 0.17 | -1.11 | 5.03 | A | A | A |
| 5.60 | 0.16 | 3.70 | 4.83 | A | A | A |
| 5.52 | 0.11 | 2.22 | 4.37 | A | A | A |
| 0.25 | 0.004 | -95.37 | 4.21 | N | N | N |
| 5.85 | 0.22 | 8.33 | 5.41 | A | A | A |
| 4.88 | 0.24 | -9.63 | 6.27 | A | A | A |
| 16.00 | 1.01 | 196.30 | 7.41 | N | N | N |
| 5.71 | 0.36 | 5.74 | 7.41 | A | A | A |
| 6.7 | 0.4 | 24.07 | 7.13 | A | N | W |
| 5.55 | 0.15 | 2.78 | 4.74 | A | A | A |
| 5.6 | 0.3 | 3.70 | 6.62 | A | A | A |
| 6 | 0.4 | 11.11 | 7.72 | A | A | A |
| 1.57 | 0.15 | -70.93 | 10.32 | N | N | N |
| 6.2 | 0.9 | 14.81 | 15.03 | A | A | A |
| 5.79 | 0.1 | 7.22 | 4.26 | A | A | A |
| 5.8 | 0.4 | 7.41 | 7.92 | A | A | A |
| 4.2 | 0.1 | -22.22 | 4.56 | A | N | W |
| 0.325 | 0.008 | -93.98 | 4.60 | N | N | N |
| 35.43 | 0.91 | 556.11 | 4.66 | N | N | N |
| 5.62 | 0.19 | 4.07 | 5.15 | A | A | A |
| 5.87 | 0.2 | 8.70 | 5.17 | A | A | A |
| 1.84 | 0.28 | -65.93 | 15.71 | N | N | N |
| 5.02 | 0.38 | -7.04 | 8.51 | A | A | A |
| 5.32 | 0.217 | -1.48 | 5.64 | A | A | A |
| 4.22 | 0.26 | -21.85 | 7.29 | A | N | W |
| 5.4 | 0.6433 | 0.00 | 12.53 | A | A | A |
| 5.54 | 0.345 | 2.59 | 7.34 | A | A | A |
| 5.33 | 0.27 | -1.30 | 6.39 | A | A | A |
| 6.23 | 0.52 | 15.37 | 9.21 | A | A | A |
| 4.3 | 0.4 | -20.37 | 10.08 | A | A | A |
| 3.57 | 0.11 | -33.89 | 4.96 | N | N | N |
| 5.7 | 0.6 | 5.56 | 11.22 | A | A | A |
| 5.03 | 0.12 | -6.85 | 4.56 | A | A | A |
| 4.83 | 0.2 | -10.56 | 5.68 | A | A | A |
| 5.50 | 0.66 | 1.85 | 12.61 | A | A | A |
| 8.5 | 0.21 | 57.41 | 4.61 | N | N | N |
| 4.301 | 0.6806 | -20.35 | 16.30 | A | A | A |
| 0.25107 | 0.0024160 | -95.35 | 4.01 | N | N | N |
| 4.50 | 0.23 | -16.67 | 6.42 | A | N | W |
| 7.54 | 0.83 | 39.63 | 11.67 | N | N | N |
| 4.18 | 0.33 | -22.59 | 8.80 | A | A | A |

TABLE 57. Evaluation Results for Cs-134 in Sample 5, Rep. Value and Rep. Unc. in Bq/sample by gamma

Labcode

184

185

186

188

189

190

193

195

197

198

199

201

202

204

205

206

207

210

211

218

222

223

224

225

228

230

233

235

239

240

241

243

245

247

248

249

251

254

258

259

260

261

262

263

264

266

268

270

271

273

275

277

279

281

282

283

285

286

287

288

TABLE 57. Evaluation Results for Cs-134 in Sample 5, Rep. Value and Rep. Unc. in Bq/sample by gamma

| Reported value [Bq/sample ] 5.56 | Reported uncertainty (k = 1) [Bq/sample ] 0.42 | Relative bias [%] 2.96 | P-Test [%] 8.50 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 4.2 | 0.2 | -22.22 | 6.15 | A | N | W |
| 6.24 | 0.21 | 15.56 | 5.14 | A | N | W |
| 5.89 | 0.20 | 9.07 | 5.16 | A | A | A |
| 5.48 | 0.4 | 1.48 | 8.27 | A | A | A |
| 5.77 | 0.49 | 6.85 | 9.34 | A | A | A |
| 5.3 | 0.5 | -1.85 | 10.20 | A | A | A |
| 3.82 | 0.28 | -29.26 | 8.30 | A | N | W |
| 4.85 | 0.10 | -10.19 | 4.40 | A | A | A |
| 8.30 | 0.10 | 53.70 | 4.07 | N | N | N |
| 4.45 | 0.55 | -17.59 | 12.96 | A | A | A |
| 5.0 | 0.2 | -7.41 | 5.58 | A | A | A |
| 5.59 | 0.23 | 3.52 | 5.66 | A | A | A |
| 7.417 | 1.855 | 37.35 | 25.31 | N | A | N |
| 7.25 | 0.21 | 34.26 | 4.85 | N | N | N |
| 61.19 | 5.94 | 1033.15 | 10.46 | N | N | N |
| 5.43 | 0.26 | 0.56 | 6.17 | A | A | A |
| 7.5 | 0.8 | 38.89 | 11.35 | N | N | N |
| 5.91 | 0.05 | 9.44 | 3.98 | A | A | A |
| 0.190 | 0.005 | -96.48 | 4.70 | N | N | N |
| 5.17 | 0.27 | -4.26 | 6.51 | A | A | A |
| 7.46 | 0.37 | 38.15 | 6.30 | N | N | N |
| 5.49 | 0.32 | 1.67 | 7.01 | A | A | A |
| 5.44 | 0.2 | 0.74 | 5.35 | A | A | A |
| 6945.84 |  | 128526.67 |  | N |  | N |

Labcode

291

294

296

298

300

303

304

306

307

308

309

316

317

319

321

322

328

330

331

333

334

338

341

342

343

Reported Results for Sr-90 determination in Sample 5

TABLE 58. Target values

| Parameter | Value |
| --- | --- |
| Activity per sample [ Bq/sample ] | 2.25 |
| Uncertainty [ Bq/sample ] | 0.10 |
| Maximum Acceptable Relative Bias, [%] | 30 |

Warning 2%

Not Accepted 17%

Massic activity  [ Bq/sample ]

2.5

Accepted 14%

Not Reported 67%

3.0

3.5

FIGURE 55. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

1.0

1.5

2.0

299

299

189

187

173

97

167

89

182

277

58

184

140

57

198

30

102

307

199

317

34

35

144

195

126

203

300

43

245

40

202

4

248

112

247

125

44

114

17

11

54

25

47

42

268

225

90

328

82

85

298

67

83

19

296

155

105

1

69

120

334

73

218

37

286

270

79

170

96

330

38

342

201

165

149

157

341

188

281

222

261

110

319

104

68

321

313

294

322

219

151

111

197

206

146

154

75

164

6

99

92

310

39

66

233

275

340

215

Laboratory code

FIGURE 56. Reported results and evaluation parameters for Sr-90. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

| Reported value [Bq/sample ] 2.44 | Reported uncertainty (k = 1) [Bq/sample ] 0.41 | Relative bias [%] 8.44 | P-Test [%] 17.38 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 2.09 | 0.88 | -7.11 | 42.34 | A | N | W |
| 9.0 | 0.18 | 300.00 | 4.87 | N | N | N |
| 2.16 | 0.18 | -4.00 | 9.44 | A | A | A |
| 2.16 | 0.11 | -4.00 | 6.76 | A | A | A |
| 2.37 | 0.13 | 5.33 | 7.06 | A | A | A |
| 2.18 | 0.4 | -3.11 | 18.88 | A | A | A |
| 1.7 | 0.3 | -24.44 | 18.20 | A | A | A |
| 1.84 | 0.13 | -18.22 | 8.35 | A | A | A |
| 1.86 | 0.15 | -17.33 | 9.21 | A | A | A |
| 2.60 | 0.37 | 15.56 | 14.91 | A | A | A |
| 3.4 | 0.9 | 51.11 | 26.84 | N | A | N |
| 12.30 | 0.65 | 446.67 | 6.91 | N | N | N |
| 2.05 | 0.11 | -8.89 | 6.97 | A | A | A |
| 2.20 | 0.20 | -2.22 | 10.12 | A | A | A |
| 2.02 | 0.24 | -10.22 | 12.69 | A | A | A |
| 2.14 | 0.13 | -4.89 | 7.53 | A | A | A |
| 2.2 | 0.1 | -2.22 | 6.36 | A | A | A |
| 2.18 | 0.27 | -3.11 | 13.16 | A | A | A |
| 1.4 | 0.1 | -37.78 | 8.41 | N | N | N |
| 1.33 | 0.17 | -40.89 | 13.53 | N | N | N |
| 18.3 | 0.9 | 713.33 | 6.63 | N | N | N |
| 2.35 | 0.16 | 4.44 | 8.13 | A | A | A |
| 4.782 | 0.096 | 112.53 | 4.88 | N | N | N |
| 2.45 | 0.18 | 8.89 | 8.59 | A | A | A |
| 2.5 | 0.2 | 11.11 | 9.15 | A | A | A |
| 8.131 | 0.040 | 261.38 | 4.47 | N | N | N |
| 2.81 | 0.28 | 24.89 | 10.91 | A | A | A |
| 2.34 | 0.13 | 4.00 | 7.11 | A | A | A |
| 2.358 | 0.072 | 4.80 | 5.39 | A | A | A |
| 2.34 | 0.05 | 4.00 | 4.93 | A | A | A |
| 0.68 | 0.07 | -69.78 | 11.21 | N | N | N |
| 2.28 | 0.15 | 1.33 | 7.94 | A | A | A |
| 9.23 | 1.28 | 310.22 | 14.56 | N | N | N |
| 2.97 | 0.1 | 32.00 | 5.58 | N | N | N |
| 0.34 | 0.02 | -84.89 | 7.37 | N | N | N |
| 9.14 | 0.015 | 306.22 | 4.45 | N | N | N |
| 1.725 | 0.606 | -23.33 | 35.41 | A | N | W |
| 4.75 | 2.88 | 111.11 | 60.79 | N | N | N |
| 2.42 | 0.17 | 7.56 | 8.31 | A | A | A |
| 4.71 | 0.02 | 109.33 | 4.46 | N | N | N |
| 7.47 | 0.22 | 232.00 | 5.33 | N | N | N |
| 2.1 | 0.1 | -6.67 | 6.51 | A | A | A |
| 2.15 | 0.11 | -4.44 | 6.78 | A | A | A |
| 2.45 | 0.22 | 8.89 | 10.02 | A | A | A |
| 2.11 | 0.11 | -6.22 | 6.85 | A | A | A |
| 1.960 | 0.014 | -12.89 | 4.50 | A | N | W |
| 1.38 | 0.75 | -38.67 | 54.53 | N | N | N |
| 1.88 | 0.20 | -16.44 | 11.53 | A | A | A |
| 7.64 | 0.32 | 239.56 | 6.11 | N | N | N |
| 4.08 | 0.36 | 81.33 | 9.88 | N | N | N |
| 7.31 | 0.17 | 224.89 | 5.02 | N | N | N |
| 7.81 | 0.7 | 247.11 | 10.00 | N | N | N |
| 2.4 | 1.1 | 6.67 | 46.05 | A | N | W |
| 4.11 | 0.50 | 82.67 | 12.95 | N | N | N |
| 8.3429 | 0.2459 | 270.80 | 5.33 | N | N | N |
| 0.64 | 0.07 | -71.56 | 11.81 | N | N | N |
| 2.85 | 0.63 | 26.67 | 22.55 | A | A | A |
| 0.34 | 0.08 | -84.89 | 23.95 | N | N | N |
| 0.73 | 0.3 | -67.56 | 41.34 | N | N | N |

TABLE 59. Evaluation Results for Sr-90 in Sample 5, Rep. Value and Rep. Unc. in Bq/sample by beta

Labcode

1

4

6

11

17

19

25

30

34

35

37

38

39

40

42

43

44

47

54

57

58

66

67

68

69

73

75

79

82

83

85

89

90

92

96

97

99

102

104

105

110

111

112

114

120

125

126

140

144

146

149

151

154

155

157

164

167

170

173

182

TABLE 59. Evaluation Results for Sr-90 in Sample 5, Rep. Value and Rep. Unc. in Bq/sample by beta

| Reported value [Bq/sample ] 1.37 | Reported uncertainty (k = 1) [Bq/sample ] 0.30 | Relative bias [%] -39.11 | P-Test [%] 22.34 | Trueness evaluation N | Precision evaluation A | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 0.149 | 0.016 | -93.38 | 11.62 | N | N | N |
| 4.3 | 0.4 | 91.11 | 10.31 | N | N | N |
| 0.104 | 0.010 | -95.38 | 10.59 | N | N | N |
| 1.93 | 0.14 | -14.22 | 8.51 | A | A | A |
| 7.5556 | 0.7619 | 235.80 | 11.02 | N | N | N |
| 1.5 | 0.07 | -33.33 | 6.44 | N | N | N |
| 1.80 | 0.10 | -20.00 | 7.11 | A | N | W |
| 4.0 | 0.1 | 77.78 | 5.10 | N | N | N |
| 2.08 | 0.04 | -7.56 | 4.84 | A | A | A |
| 1.97 | 0.12 | -12.44 | 7.54 | A | A | A |
| 7.584 | 0.036 | 237.07 | 4.47 | N | N | N |
| 56 | 6 | 2388.89 | 11.60 | N | N | N |
| 2.54 | 0.12 | 12.89 | 6.49 | A | A | A |
| 7.03 | 0.69 | 212.44 | 10.77 | N | N | N |
| 4.58 | 0.50 | 103.56 | 11.79 | N | N | N |
| 2.26 | 0.28 | 0.44 | 13.16 | A | A | A |
| 18.75 | 2.26 | 733.33 | 12.85 | N | N | N |
| 2.02 | 0.09 | -10.22 | 6.29 | A | A | A |
| 2.1 | 0.05 | -6.67 | 5.04 | A | A | A |
| 2.1 | 0.2 | -6.67 | 10.51 | A | A | A |
| 4.66 | 0.45 | 107.11 | 10.63 | N | N | N |
| 2.25 | 0.14 | 0.00 | 7.65 | A | A | A |
| 2.68 | 0.47 | 19.11 | 18.09 | A | A | A |
| 22.8 | 6.0 | 913.33 | 26.69 | N | N | N |
| 0.935 | 0.051 | -58.44 | 7.04 | N | N | N |
| 4.464 | 0.446 | 98.40 | 10.93 | N | N | N |
| 2.6 | 1.3 | 15.56 | 50.20 | A | N | W |
| 5.9 | 0.6 | 162.22 | 11.10 | N | N | N |
| 2.39 | 0.238 | 6.22 | 10.90 | A | A | A |
| 2.34 | 0.11 | 4.00 | 6.47 | A | A | A |
| 0.06 | 0.02 | -97.33 | 33.63 | N | N | N |
| 2.0 | 0.27 | -11.11 | 14.21 | A | A | A |
| 1.75 | 0.11 | -22.22 | 7.70 | A | N | W |
| 9.89 | 0.038 | 339.56 | 4.46 | N | N | N |
| 5.32 | 0.03 | 136.44 | 4.48 | N | N | N |
| 1.801 | 0.06 | -19.96 | 5.55 | A | N | W |
| 4.732 | 0.725 | 110.31 | 15.95 | N | N | N |
| 4.99 | 0.28 | 121.78 | 7.16 | N | N | N |
| 6.29 | 1.21 | 179.56 | 19.74 | N | N | N |
| 2.30 | 0.11 | 2.22 | 6.53 | A | A | A |
| 3.1 | 0.5 | 37.78 | 16.73 | N | A | N |
| 2.5 | 0.5 | 11.11 | 20.49 | A | A | A |
| 28.4 | 2.6 | 1162.22 | 10.18 | N | N | N |
| 4.15 | 0.89 | 84.44 | 21.90 | N | N | N |
| 3.91 | 0.59 | 73.78 | 15.73 | N | N | N |

Labcode

184

187

188

189

195

197

198

199

201

202

203

206

215

218

219

222

225

233

245

247

248

261

268

270

275

277

281

286

294

296

298

299

300

307

310

313

317

319

321

322

328

330

334

340

341

342

Reported Results for Cs-134 determination in Sample 6

TABLE 60. Target values

112

Massic activity  [ Bq/sample ]

FIGURE 58. Reported results and evaluation parameters for Cs-134. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

4	6	8	10	12

68

189

84

333

285

224

167

33

251

96

243

142

13

117

273

306

211

154

288

72

210

249

FIGURE 57. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

199

144

263

286

283

271

197

54

294

126

260

111

110

70

172

90

62

309

67

183

206

140

228

307

97

207

176

170

83

25

1

304

58

82

279

334

Not Accepted 11%

12

38

4

184

Parameter

Activity per sample [ Bq/sample ] Uncertainty [ Bq/sample ]

Maximum Acceptable Relative Bias, [%]

2

119

198

29

168

135

316

5

131

Warning 9%

145

34

57

277

158

193

223

262

118

46

60

81

75

63

20

64

153

105

174

71

186

Laboratory code

341

141

205

43

18

268

342

134

218

165

261

180

181

130

17

173

156

39

106

98

10

169

Accepted 45%

328

264

55

188

112

300

241

114

291

6

Not Reported 35%

122

281

7

317

44

26

79

222

195

37

104

115

133

247

35

125

73

Value 7.84

0.31

30

151

161

258

344

266

225

303

185

95

157

179

275

240

148

166

19

85

202

40

15

233

331

155

36

190

47

298

14

178

66

42

99

52

248

11

259

69

159

30

245

270

123

182

27

338

177

319

296

235

321

308

287

146

282

31

330

16

89

149

230

239

93

254

322

102

343

| Reported value [Bq/sample ] 7.1 | Reported uncertainty (k = 1) [Bq/sample ] 0.3 | Relative bias [%] -9.44 | P-Test [%] 5.79 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 7.28 | 0.27 | -7.14 | 5.42 | A | A | A |
| 7.22 | 0.39 | -7.91 | 6.69 | A | A | A |
| 7.42 | 0.58 | -5.36 | 8.76 | A | A | A |
| 8.03 | 0.28 | 2.42 | 5.27 | A | A | A |
| 8.09 | 0.11 | 3.19 | 4.18 | A | A | A |
| 7.9 | 0.4 | 0.77 | 6.42 | A | A | A |
| 8.81 | 0.71 | 12.37 | 8.98 | A | A | A |
| 7.2 | 0.7 | -8.16 | 10.50 | A | A | A |
| 4.4 | 0.05 | -43.88 | 4.11 | N | N | N |
| 8.62 | 0.53 | 9.95 | 7.31 | A | A | A |
| 8.4 | 0.1 | 7.14 | 4.13 | A | A | A |
| 17.08 | 1.86 | 117.86 | 11.59 | N | N | N |
| 7.81 | 0.49 | -0.38 | 7.42 | A | A | A |
| 7.74 | 0.52 | -1.28 | 7.80 | A | A | A |
| 8.32 | 0.58 | 6.12 | 8.01 | A | A | A |
| 7.68 | 0.515 | -2.04 | 7.78 | A | A | A |
| 7.07 | 0.8 | -9.82 | 11.99 | A | A | A |
| 8.10 | 0.62 | 3.32 | 8.62 | A | A | A |
| 9.49 | 0.23 | 21.05 | 4.64 | A | N | W |
| 7.3 | 0.5 | -6.89 | 7.91 | A | A | A |
| 9.0 | 0.8 | 14.80 | 9.73 | A | A | A |
| 12.8 | 0.7 | 63.27 | 6.75 | N | N | N |
| 0.44 | 0.03 | -94.39 | 7.88 | N | N | N |
| 7.5 | 0.58 | -4.34 | 8.69 | A | A | A |
| 8.19 | 0.66 | 4.46 | 8.98 | A | A | A |
| 8.5 | 0.7 | 8.42 | 9.14 | A | A | A |
| 8.12 | 0.92 | 3.57 | 12.00 | A | A | A |
| 7.22 | 0.41 | -7.91 | 6.92 | A | A | A |
| 7.83 | 0.40 | -0.13 | 6.46 | A | A | A |
| 8.39 | 0.343 | 7.02 | 5.69 | A | A | A |
| 8.70 | 0.83 | 10.97 | 10.33 | A | A | A |
| 7.74 | 0.85 | -1.28 | 11.67 | A | A | A |
| 8.09 | 0.61 | 3.19 | 8.51 | A | A | A |
| 7.6 | 0.3 | -3.06 | 5.59 | A | A | A |
| 8.6 | 0.3 | 9.69 | 5.27 | A | A | A |
| 8.8 | 1.4 | 12.24 | 16.39 | A | A | A |
| 6.23 | 0.18 | -20.54 | 4.90 | A | N | W |
| 8.0 | 0.2 | 2.04 | 4.68 | A | A | A |
| 7.5 | 0.3 | -4.34 | 5.62 | A | A | A |
| 7.10 | 0.86 | -9.44 | 12.74 | A | A | A |
| 7.60 | 0.32 | -3.06 | 5.78 | A | A | A |
| 6.7 | 0.5 | -14.54 | 8.45 | A | A | A |
| 7.67 | 0.21 | -2.17 | 4.81 | A | A | A |
| 7.68 | 0.42 | -2.04 | 6.75 | A | A | A |
| 8.7 | 0.14 | 10.97 | 4.27 | A | A | A |
| 6.83 | 0.09 | -12.88 | 4.17 | A | N | W |
| 0.0847 | 0.0035 | -98.92 | 5.72 | N | N | N |
| 8.85 | 0.29 | 12.88 | 5.14 | A | A | A |
| 6.4 | 0.1 | -18.37 | 4.25 | A | N | W |
| 7.7 | 0.6 | -1.79 | 8.74 | A | A | A |
| 5.73 | 0.13 | -26.91 | 4.56 | A | N | W |
| 8.2 | 0.2 | 4.59 | 4.65 | A | A | A |
| 7.63 | 0.18 | -2.68 | 4.60 | A | A | A |
| 8.1 | 0.11 | 3.32 | 4.18 | A | A | A |
| 7.61 | 0.23 | -2.93 | 4.98 | A | A | A |
| 7.11 | 0.1 | -9.31 | 4.20 | A | A | A |
| 7.02 | 0.108 | -10.46 | 4.24 | A | A | A |
| 0.18 | 0.12 | -97.70 | 66.78 | N | N | N |
| 8.36 | 0.62 | 6.63 | 8.40 | A | A | A |

TABLE 61. Evaluation Results for Cs-134 in Sample 6, Rep. Value and Rep. Unc. in Bq/sample by gamma

Labcode

1

2

4

5

6

7

10

11

12

13

14

15

16

17

18

19

20

25

26

27

29

30

31

33

34

35

36

37

38

39

40

42

43

44

46

47

52

54

55

57

58

60

62

63

64

66

67

68

69

70

71

72

73

75

79

81

82

83

84

85

| Reported value [Bq/sample ] 17.98 | Reported uncertainty (k = 1) [Bq/sample ] 0.9 | Relative bias [%] 129.34 | P-Test [%] 6.38 | Trueness evaluation N | Precision evaluation N | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 6.67 | 0.32 | -14.92 | 6.22 | A | A | A |
| 30.2 | 0.3 | 285.20 | 4.08 | N | N | N |
| 8.26 | 0.52 | 5.36 | 7.43 | A | A | A |
| 1.0 | 0.01 | -87.24 | 4.08 | N | N | N |
| 6.97 | 0.05 | -11.10 | 4.02 | A | N | W |
| 7.89 | 0.40 | 0.64 | 6.43 | A | A | A |
| 8.75 | 0.13 | 11.61 | 4.22 | A | N | W |
| 119.1 | 2.0 | 1419.13 | 4.30 | N | N | N |
| 8.12 | 0.49 | 3.57 | 7.21 | A | A | A |
| 7.70 | 0.40 | -1.79 | 6.53 | A | A | A |
| 7.89 | 0.33 | 0.64 | 5.76 | A | A | A |
| 6.39 | 0.25 | -18.49 | 5.56 | A | N | W |
| 6.388 | 0.25 | -18.52 | 5.56 | A | N | W |
| 8.0 | 0.3 | 2.04 | 5.45 | A | A | A |
| 8.00 | 0.82 | 2.04 | 10.99 | A | A | A |
| 8.16 | 0.29 | 4.08 | 5.32 | A | A | A |
| 4.93 | 0.24 | -37.12 | 6.27 | N | N | N |
| 7.6 | 0.3 | -3.06 | 5.59 | A | A | A |
| 7.28 | 0.8 | -7.14 | 11.68 | A | A | A |
| 8.05 | 0.87 | 2.68 | 11.51 | A | A | A |
| 9.30 | 0.24 | 18.62 | 4.72 | A | N | W |
| 8.20 | 0.53 | 4.59 | 7.58 | A | A | A |
| 6.32 | 0.68 | -19.39 | 11.46 | A | A | A |
| 7.8 | 0.7 | -0.51 | 9.81 | A | A | A |
| 7.45 | 0.40 | -4.97 | 6.67 | A | A | A |
| 8.18 | 0.98 | 4.34 | 12.62 | A | A | A |
| 7.75 | 1.14 | -1.15 | 15.23 | A | A | A |
| 7.4 | 0.5 | -5.61 | 7.83 | A | A | A |
| 6.9 | 0.4 | -11.99 | 7.02 | A | A | A |
| 7.721 | 0.341 | -1.52 | 5.93 | A | A | A |
| 2.705 | 0.074 | -65.50 | 4.81 | N | N | N |
| 6.04 | 0.45 | -22.96 | 8.43 | A | N | W |
| 7.47 | 0.58 | -4.72 | 8.71 | A | A | A |
| 11.05 | 0.46 | 40.94 | 5.74 | N | N | N |
| 8.3 | 0.2 | 5.87 | 4.63 | A | A | A |
| 19.3 | 0.434 | 146.17 | 4.55 | N | N | N |
| 8.2 | 0.4 | 4.59 | 6.28 | A | A | A |
| 7.693 | 0.340 | -1.88 | 5.93 | A | A | A |
| 5.56 | 0.2 | -29.08 | 5.35 | A | N | W |
| 8.48 | 0.45 | 8.16 | 6.62 | A | A | A |
| 7.83 | 0.80 | -0.13 | 10.96 | A | A | A |
| 8.28 | 0.14 | 5.61 | 4.30 | A | A | A |
| 7.53 | 0.76 | -3.95 | 10.84 | A | A | A |
| 8.89 | 0.16 | 13.39 | 4.34 | A | N | W |
| 8.2 | 0.7 | 4.59 | 9.41 | A | A | A |
| 7.78 | 0.91 | -0.77 | 12.35 | A | A | A |
| 8.3 | 1.1 | 5.87 | 13.83 | A | A | A |
| 0.41 | 0.03 | -94.77 | 8.32 | N | N | N |
| 7.4 | 0.4 | -5.61 | 6.70 | A | A | A |
| 7.9 | 0.7 | 0.77 | 9.70 | A | A | A |
| 7.02 | 0.43 | -10.46 | 7.29 | A | A | A |
| 6.54 | 0.43 | -16.58 | 7.67 | A | A | A |
| 7.81 | 0.18 | -0.38 | 4.58 | A | A | A |
| 7.7 | 0.5 | -1.79 | 7.60 | A | A | A |
| 7.0 | 0.5 | -10.71 | 8.16 | A | A | A |
| 9.85 | 0.32 | 25.64 | 5.12 | A | N | W |
| 8.63 | 0.14 | 10.08 | 4.27 | A | A | A |
| 8.28 | 0.27 | 5.61 | 5.13 | A | A | A |
| 7.80 | 0.15 | -0.51 | 4.40 | A | A | A |

TABLE 61. Evaluation Results for Cs-134 in Sample 6, Rep. Value and Rep. Unc. in Bq/sample by gamma

Labcode

89

90

93

95

96

97

98

99

102

104

105

106

110

111

112

114

115

117

118

119

122

123

125

126

130

131

133

134

135

140

141

142

144

145

146

148

149

151

153

154

155

156

157

158

159

161

165

166

167

168

169

170

172

173

174

176

177

178

179

180

| Reported value [Bq/sample ] 7.8 | Reported uncertainty (k = 1) [Bq/sample ] 2.2 | Relative bias [%] -0.51 | P-Test [%] 28.48 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 9.4 | 1.0 | 19.90 | 11.35 | A | A | A |
| 6.89 | 0.35 | -12.12 | 6.44 | A | A | A |
| 7.25 | 1.45 | -7.53 | 20.39 | A | A | A |
| 8.260 | 0.909 | 5.36 | 11.69 | A | A | A |
| 7.7 | 0.2 | -1.79 | 4.73 | A | A | A |
| 8.0 | 0.4 | 2.04 | 6.37 | A | A | A |
| 0.0944 | 0.0083 | -98.80 | 9.64 | N | N | N |
| 8.6 | 0.5 | 9.69 | 7.03 | A | A | A |
| 7.54 | 0.11 | -3.83 | 4.21 | A | A | A |
| 8.11 | 0.92 | 3.44 | 12.01 | A | A | A |
| 6.2012 | 0.15 | -20.90 | 4.64 | A | N | W |
| 7.3 | 0.2 | -6.89 | 4.81 | A | A | A |
| 6.01 | 0.08 | -23.34 | 4.17 | A | N | W |
| 8.37 | 0.38 | 6.76 | 6.02 | A | A | A |
| 7.726 | 0.151 | -1.45 | 4.41 | A | A | A |
| 6.896 | 0.109 | -12.04 | 4.26 | A | N | W |
| 6.97 | 0.22 | -11.10 | 5.06 | A | A | A |
| 5.84 | 0.2 | -25.51 | 5.23 | A | N | W |
| 5.472 | 0.109 | -30.20 | 4.43 | N | N | N |
| 7.77 | 0.25 | -0.89 | 5.10 | A | A | A |
| 8.10 | 0.24 | 3.32 | 4.94 | A | A | A |
| 7.59 | 0.14 | -3.19 | 4.36 | A | A | A |
| 0.36 | 0.005 | -95.41 | 4.19 | N | N | N |
| 8.25 | 0.28 | 5.23 | 5.21 | A | A | A |
| 6.94 | 0.34 | -11.48 | 6.30 | A | A | A |
| 23.00 | 1.08 | 193.37 | 6.14 | N | N | N |
| 8.46 | 0.56 | 7.91 | 7.71 | A | A | A |
| 10.0 | 0.5 | 27.55 | 6.37 | A | N | W |
| 26.5 | 0.58 | 238.01 | 4.52 | N | N | N |
| 8.3 | 0.5 | 5.87 | 7.21 | A | A | A |
| 8.0 | 0.5 | 2.04 | 7.40 | A | A | A |
| 2.2 | 0.2 | -71.94 | 9.91 | N | N | N |
| 9.0 | 1.3 | 14.80 | 14.98 | A | A | A |
| 8.19 | 0.12 | 4.46 | 4.22 | A | A | A |
| 8.8 | 0.5 | 12.24 | 6.92 | A | A | A |
| 5.9 | 0.2 | -24.74 | 5.21 | A | N | W |
| 0.481 | 0.009 | -93.86 | 4.37 | N | N | N |
| 52.23 | 1.36 | 566.20 | 4.73 | N | N | N |
| 8.20 | 0.28 | 4.59 | 5.22 | A | A | A |
| 8.85 | 0.2 | 12.88 | 4.55 | A | N | W |
| 6.38 | 0.64 | -18.62 | 10.78 | A | A | A |
| 7.79 | 0.61 | -0.64 | 8.77 | A | A | A |
| 7.60 | 0.362 | -3.06 | 6.19 | A | A | A |
| 6.12 | 0.40 | -21.94 | 7.64 | A | N | W |
| 7.938 | 0.9393 | 1.25 | 12.48 | A | A | A |
| 8.22 | 0.505 | 4.85 | 7.31 | A | A | A |
| 7.75 | 0.40 | -1.15 | 6.50 | A | A | A |
| 9.28 | 0.72 | 18.37 | 8.71 | A | A | A |
| 6.2 | 0.4 | -20.92 | 7.57 | A | N | W |
| 5.14 | 0.15 | -34.44 | 4.91 | N | N | N |
| 8.3 | 0.6 | 5.87 | 8.24 | A | A | A |
| 7.50 | 0.07 | -4.34 | 4.06 | A | A | A |
| 7.11 | 0.286 | -9.31 | 5.64 | A | A | A |
| 8.075 | 0.969 | 3.00 | 12.63 | A | A | A |
| 12 | 0.27 | 53.06 | 4.55 | N | N | N |
| 6.156 | 0.9651 | -21.48 | 16.17 | A | A | A |
| 0.32708 | 0.0027546 | -95.83 | 4.04 | N | N | N |
| 6.12 | 0.29 | -21.94 | 6.17 | A | N | W |
| 11.0 | 1.21 | 40.31 | 11.69 | N | N | N |

TABLE 61. Evaluation Results for Cs-134 in Sample 6, Rep. Value and Rep. Unc. in Bq/sample by gamma

Labcode

181

182

183

184

185

186

188

189

190

193

195

197

198

199

202

205

206

207

210

211

218

222

223

224

225

228

230

233

235

239

240

241

243

245

247

248

249

251

254

258

259

260

261

262

263

264

266

268

270

271

273

275

277

279

281

282

283

285

286

287

TABLE 61. Evaluation Results for Cs-134 in Sample 6, Rep. Value and Rep. Unc. in Bq/sample by gamma

| Reported value [Bq/sample ] 5.65 | Reported uncertainty (k = 1) [Bq/sample ] 0.45 | Relative bias [%] -27.93 | P-Test [%] 8.89 | Trueness evaluation A | Precision evaluation N | Final Score W |
| --- | --- | --- | --- | --- | --- | --- |
| 8.03 | 0.59 | 2.42 | 8.34 | A | A | A |
| 6.3 | 0.3 | -19.64 | 6.19 | A | N | W |
| 9.94 | 0.32 | 26.79 | 5.10 | A | N | W |
| 8.61 | 0.28 | 9.82 | 5.12 | A | A | A |
| 8 | 0.5 | 2.04 | 7.40 | A | A | A |
| 8.25 | 0.68 | 5.23 | 9.14 | A | A | A |
| 7.1 | 0.6 | -9.44 | 9.33 | A | A | A |
| 5.32 | 0.49 | -32.14 | 10.02 | N | N | N |
| 6.95 | 0.10 | -11.35 | 4.21 | A | N | W |
| 10.79 | 0.09 | 37.63 | 4.04 | N | N | N |
| 6.71 | 0.57 | -14.41 | 9.37 | A | A | A |
| 7.4 | 0.1 | -5.61 | 4.18 | A | A | A |
| 8.09 | 0.11 | 3.19 | 4.18 | A | A | A |
| 9.859 | 2.465 | 25.75 | 25.31 | A | A | A |
| 10.21 | 0.21 | 30.23 | 4.46 | N | N | N |
| 67.7 | 6.21 | 763.52 | 9.99 | N | N | N |
| 7.92 | 0.33 | 1.02 | 5.74 | A | A | A |
| 13.3 | 1.3 | 69.64 | 10.54 | N | N | N |
| 8.47 | 0.11 | 8.04 | 4.16 | A | A | A |
| 0.316 | 0.006 | -95.97 | 4.39 | N | N | N |
| 7.2 | 0.4 | -8.16 | 6.82 | A | A | A |
| 9.85 | 0.32 | 25.64 | 5.12 | A | N | W |
| 7.72 | 0.41 | -1.53 | 6.62 | A | A | A |
| 7.75 | 0.33 | -1.15 | 5.81 | A | A | A |
| 1109.83 |  | 14055.99 |  | N |  | N |
| 8.2 | 1.3 | 4.59 | 16.34 | A | A | A |

Labcode

288

291

294

296

298

300

303

304

306

307

308

309

316

317

319

321

322

328

330

331

333

334

338

341

342

343

344

Reported Results for Sr-90 determination in Sample 7

TABLE 62. Target values

| Parameter | Value |
| --- | --- |
| Activity per sample [ Bq/sample ] | 4.52 |
| Uncertainty [ Bq/sample ] | 0.22 |
| Maximum Acceptable Relative Bias, [%] | 30 |

Not Accepted 18%

Warning 2%

Accepted 20%

Massic activity  [ Bq/sample ]

5

Not Reported 60%

6

7

FIGURE 59. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

2

3

4

299

189

302

187

196

167

277

15

58

263

102

57

184

96

97

135

126

35

89

195

203

125

199

30

300

38

25

245

44

144

11

173

17

198

279

34

112

330

82

317

40

90

83

166

170

264

185

85

296

62

248

19

202

247

105

225

67

4

218

328

1

307

334

42

104

54

149

73

52

182

280

119

133

342

268

270

37

120

341

69

179

43

319

321

286

110

313

210

201

322

294

281

155

261

16

206

154

99

146

47

188

151

117

111

68

75

157

46

164

181

197

219

79

222

6

344

310

165

140

39

92

233

275

340

215

145

Laboratory code

FIGURE 60. Reported results and evaluation parameters for Sr-90. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores respectively.

| Reported value [Bq/sample ] 4.75 | Reported uncertainty (k = 1) [Bq/sample ] 0.79 | Relative bias [%] 5.09 | P-Test [%] 17.33 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 4.73 | 0.89 | 4.65 | 19.44 | A | A | A |
| 11.6 | 0.23 | 156.64 | 5.26 | N | N | N |
| 4.25 | 0.36 | -5.97 | 9.77 | A | A | A |
| 2.1 | 0.04 | -53.54 | 5.23 | N | N | N |
| 8.69 | 0.28 | 92.26 | 5.84 | N | N | N |
| 4.29 | 0.18 | -5.09 | 6.43 | A | A | A |
| 4.61 | 0.24 | 1.99 | 7.13 | A | A | A |
| 4.10 | 0.7 | -9.29 | 17.75 | A | A | A |
| 4.1 | 0.5 | -9.29 | 13.13 | A | A | A |
| 4.37 | 0.32 | -3.32 | 8.79 | A | A | A |
| 3.73 | 0.30 | -17.48 | 9.40 | A | A | A |
| 5.10 | 0.7 | 12.83 | 14.56 | A | A | A |
| 4.1 | 1.1 | -9.29 | 27.27 | A | A | A |
| 14.60 | 0.77 | 223.01 | 7.18 | N | N | N |
| 4.45 | 0.24 | -1.55 | 7.26 | A | A | A |
| 4.80 | 0.46 | 6.19 | 10.75 | A | A | A |
| 5.75 | 0.12 | 27.21 | 5.30 | A | N | W |
| 4.18 | 0.23 | -7.52 | 7.35 | A | A | A |
| 9.64 | 0.1 | 113.27 | 4.98 | N | N | N |
| 9.17 | 0.1 | 102.88 | 4.99 | N | N | N |
| 4.9 | 1.4 | 8.41 | 28.98 | A | A | A |
| 4.82 | 0.11 | 6.64 | 5.38 | A | A | A |
| 2.74 | 0.15 | -39.38 | 7.33 | N | N | N |
| 2.40 | 0.29 | -46.90 | 13.03 | N | N | N |
| 4.6 | 0.3 | 1.77 | 8.14 | A | A | A |
| 4.72 | 0.06 | 4.42 | 5.03 | A | A | A |
| 9.51 | 0.12 | 110.40 | 5.03 | N | N | N |
| 5.4 | 0.4 | 19.47 | 8.86 | A | A | A |
| 4.9 | 0.4 | 8.41 | 9.50 | A | A | A |
| 9.521 | 0.043 | 110.64 | 4.89 | N | N | N |
| 10.5 | 0.54 | 132.30 | 7.08 | N | N | N |
| 4.43 | 0.24 | -1.99 | 7.28 | A | A | A |
| 4.5 | 0.144 | -0.44 | 5.82 | A | A | A |
| 4.56 | 0.10 | 0.88 | 5.34 | A | A | A |
| 3.85 | 0.57 | -14.82 | 15.58 | A | A | A |
| 4.48 | 0.28 | -0.88 | 7.92 | A | A | A |
| 16.41 | 3.18 | 263.05 | 19.98 | N | N | N |
| 2.96 | 0.09 | -34.51 | 5.74 | N | N | N |
| 3.19 | 0.08 | -29.42 | 5.48 | A | N | W |
| 8.92 | 0.015 | 97.35 | 4.87 | N | N | N |
| 2.716 | 0.728 | -39.91 | 27.24 | N | A | N |
| 4.81 | 1.88 | 6.42 | 39.39 | A | N | W |
| 4.67 | 0.30 | 3.32 | 8.06 | A | A | A |
| 6.56 | 0.09 | 45.13 | 5.06 | N | N | N |
| 9.38 | 0.28 | 107.52 | 5.71 | N | N | N |
| 4.4 | 0.2 | -2.65 | 6.66 | A | A | A |
| 9.37 | 0.45 | 107.30 | 6.84 | N | N | N |
| 4.92 | 0.5 | 8.85 | 11.27 | A | A | A |
| 5.14 | 0.46 | 13.72 | 10.19 | A | A | A |
| 4.02 | 0.21 | -11.06 | 7.14 | A | A | A |
| 3.580 | 0.019 | -20.80 | 4.90 | A | N | W |
| 4.94 | 0.24 | 9.29 | 6.88 | A | A | A |
| 3.3 | 0.6 | -26.99 | 18.82 | A | A | A |
| 13.45 | 0.38 | 197.57 | 5.63 | N | N | N |
| 4.19 | 0.87 | -7.30 | 21.33 | A | A | A |
| 567 | 23 | 12444.25 | 6.34 | N | N | N |
| 8.99 | 0.38 | 98.89 | 6.45 | N | N | N |
| 4.89 | 0.43 | 8.19 | 10.05 | A | A | A |
| 9.34 | 0.22 | 106.64 | 5.41 | N | N | N |

TABLE 63. Evaluation Results for Sr-90 in Sample 7, Rep. Value and Rep. Unc. in Bq/sample by beta

Labcode

1

4

6

11

15

16

17

19

25

30

34

35

37

38

39

40

42

43

44

46

47

52

54

57

58

62

67

68

69

73

75

79

82

83

85

89

90

92

96

97

99

102

104

105

110

111

112

117

119

120

125

126

133

135

140

144

145

146

149

151

| Reported value [Bq/sample ] 8.8 | Reported uncertainty (k = 1) [Bq/sample ] 0.8 | Relative bias [%] 94.69 | P-Test [%] 10.31 | Trueness evaluation N | Precision evaluation N | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 8.4 | 1.1 | 85.84 | 13.97 | N | N | N |
| 9.55 | 1.00 | 111.28 | 11.55 | N | N | N |
| 9.7269 | 2.0544 | 115.20 | 21.67 | N | N | N |
| 12.35 | 0.51 | 173.23 | 6.38 | N | N | N |
| 4.5 | 0.9 | -0.44 | 20.58 | A | A | A |
| 0.75 | 0.08 | -83.41 | 11.72 | N | N | N |
| 4.53 | 0.56 | 0.22 | 13.29 | A | A | A |
| 4.26 | 0.5 | -5.75 | 12.71 | A | A | A |
| 5.55 | 0.11 | 22.79 | 5.26 | A | N | W |
| 9.8 | 1.0 | 116.81 | 11.31 | N | N | N |
| 4.9 | 0.5 | 8.41 | 11.31 | A | A | A |
| 2.95 | 0.60 | -34.73 | 20.91 | N | A | N |
| 4.56 | 0.88 | 0.88 | 19.90 | A | A | A |
| 0.209 | 0.019 | -95.38 | 10.31 | N | N | N |
| 9.2 | 0.5 | 103.54 | 7.30 | N | N | N |
| 0.130 | 0.013 | -97.12 | 11.12 | N | N | N |
| 3.87 | 0.27 | -14.38 | 8.51 | A | A | A |
| 0.674 | 0.118 | -85.09 | 18.17 | N | N | N |
| 9.8365 | 0.9874 | 117.62 | 11.16 | N | N | N |
| 4.3 | 0.07 | -4.87 | 5.13 | A | A | A |
| 4.05 | 0.13 | -10.40 | 5.83 | A | A | A |
| 7.5 | 0.3 | 65.93 | 6.30 | N | N | N |
| 4.63 | 0.06 | 2.43 | 5.04 | A | A | A |
| 4.00 | 0.24 | -11.50 | 7.73 | A | A | A |
| 8.775 | 0.034 | 94.14 | 4.88 | N | N | N |
| 7.310 | 0.449 | 61.73 | 7.84 | N | N | N |
| 92 | 9 | 1935.40 | 10.93 | N | N | N |
| 4.74 | 0.13 | 4.87 | 5.59 | A | A | A |
| 10.17 | 1.02 | 125.00 | 11.15 | N | N | N |
| 11.00 | 0.50 | 143.36 | 6.66 | N | N | N |
| 4.71 | 0.26 | 4.20 | 7.36 | A | A | A |
| 23.57 | 2.82 | 421.46 | 12.92 | N | N | N |
| 4.13 | 0.19 | -8.63 | 6.70 | A | A | A |
| 4.63 | 0.10 | 2.43 | 5.32 | A | A | A |
| 4.6 | 0.4 | 1.77 | 9.97 | A | A | A |
| 8.60 | 0.51 | 90.27 | 7.67 | N | N | N |
| 2.58 | 0.10 | -42.92 | 6.22 | N | N | N |
| 4.55 | 0.439 | 0.66 | 10.81 | A | A | A |
| 5.01 | 0.27 | 10.84 | 7.26 | A | A | A |
| 5.07 | 0.68 | 12.17 | 14.27 | A | A | A |
| 28.2 | 7.0 | 523.89 | 25.30 | N | N | N |
| 1.820 | 0.082 | -59.73 | 6.63 | N | N | N |
| 4.36 | 0.083 | -3.54 | 5.23 | A | A | A |
| 4.901 | 0.196 | 8.43 | 6.30 | A | A | A |
| 8.16 | 0.816 | 80.53 | 11.12 | N | N | N |
| 6.14 | 0.92 | 35.84 | 15.75 | N | A | N |
| 8.0 | 0.5 | 76.99 | 7.92 | N | N | N |
| 4.58 | 0.42 | 1.33 | 10.38 | A | A | A |
| 0.07 | 0.02 | -98.45 | 28.98 | N | N | N |
| 4.1 | 0.43 | -9.29 | 11.56 | A | A | A |
| 0.202 | 0.0036 | -95.53 | 5.18 | N | N | N |
| 4.79 | 0.15 | 5.97 | 5.79 | A | A | A |
| 11.73 | 0.040 | 159.51 | 4.88 | N | N | N |
| 6.61 | 0.04 | 46.24 | 4.90 | N | N | N |
| 4.43 | 0.28 | -1.99 | 7.98 | A | A | A |
| 5.9650 | 0.845 | 31.97 | 14.98 | N | A | N |
| 6.06 | 0.34 | 34.07 | 7.43 | N | N | N |
| 7.76 | 0.83 | 71.68 | 11.75 | N | N | N |
| 4.74 | 0.17 | 4.87 | 6.05 | A | A | A |

TABLE 63. Evaluation Results for Sr-90 in Sample 7, Rep. Value and Rep. Unc. in Bq/sample by beta

Labcode

154

155

157

164

165

166

167

170

173

179

181

182

184

185

187

188

189

195

196

197

198

199

201

202

203

206

210

215

218

219

222

225

233

245

247

248

261

263

264

268

270

275

277

279

280

281

286

294

296

299

300

302

307

310

313

317

319

321

322

328

TABLE 63. Evaluation Results for Sr-90 in Sample 7, Rep. Value and Rep. Unc. in Bq/sample by beta

| Reported value [Bq/sample ] 4.4 | Reported uncertainty (k = 1) [Bq/sample ] 0.6 | Relative bias [%] -2.65 | P-Test [%] 14.48 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 4.8 | 1.0 | 6.19 | 21.39 | A | A | A |
| 29.6 | 2.7 | 554.87 | 10.34 | N | N | N |
| 5.16 | 1.10 | 14.16 | 21.87 | A | A | A |
| 4.96 | 0.74 | 9.73 | 15.69 | A | A | A |
| 11.6 | 1.9 | 156.64 | 17.09 | N | N | N |

Labcode

330

334

340

341

342

344

Activity [ Bq/kg ]

30

40

50

Display of intercomparison values

Reported Results for gross alpha Determination in Sample 1

TABLE 64. Target values

| Parameter | Value |
| --- | --- |
| Robust Mean, Bq/kg | 27.8 |
| Robust SD, Bq/kg | 6.4 |
| N | 240 |

10

20

224

302

310

322

341

91

229

210

68

260

313

345

75

202

295

74

223

47

97

46

82

304

288

152

56

165

96

16

173

99

197

169

87

147

131

333

93

19

319

219

106

70

201

206

127

268

199

83

296

279

123

54

232

321

4

119

65

34

126

158

198

104

17

84

277

130

39

25

85

58

145

144

146

89

195

2

151

312

299

290

49

315

35

112

101

157

154

120

180

44

37

247

129

248

50

86

185

317

175

124

42

32

95

40

293

12

291

263

6

249

22

43

55

204

261

133

149

27

188

135

241

181

52

36

136

334

225

259

3

245

196

78

98

57

31

26

187

155

102

159

218

270

307

207

184

28

160

330

203

29

262

214

11

72

1

235

33

189

90

164

167

14

221

186

45

266

67

115

272

79

53

Labcode

FIGURE 61. Reported results and evaluation parameters for gross alpha in Sample 1.

TABLE 65. IC Evaluation Results for gross alpha in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 1 | 34.2 | 2.1 | 1.00 | A |
| 2 | 27.65 | 1.89 | 0.02 | A |
| 3 | 32.1 | 1.18 | 0.67 | A |
| 4 | 25.66 | 0.42 | 0.33 | A |
| 6 | 30.0 | 1.5 | 0.34 | A |
| 11 | 34.12 | 2.32 | 0.99 | A |
| 12 | 29.5 | 10.3 | 0.27 | A |
| 14 | 37.0 | 2.4 | 1.44 | A |
| 16 | 20.04 | 0.74 | 1.21 | A |
| 17 | 26.2 | 3.2 | 0.25 | A |
| 19 | 22.0 | 3.1 | 0.91 | A |
| 22 | 30.24 | 1.81 | 0.38 | A |
| 25 | 26.65 | 4.0 | 0.18 | A |
| 26 | 32.9 | 2.04 | 0.80 | A |
| 27 | 30.9 | 1.9 | 0.48 | A |
| 28 | 33.4 | 5.0 | 0.87 | A |
| 29 | 34 | 2 | 0.97 | A |
| 31 | 32.8 | 5.8 | 0.78 | A |
| 32 | 28.98 | 0.63 | 0.18 | A |
| 33 | 34.8 | 4.6 | 1.09 | A |
| 34 | 25.9 | 3.5 | 0.30 | A |
| 35 | 27.9 | 3.1 | 0.02 | A |
| 36 | 31.3 | 2.0 | 0.55 | A |
| 37 | 28.6 | 2.6 | 0.13 | A |

TABLE 65. IC Evaluation Results for gross alpha in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 39 | 26.5 | 1.1 | 0.20 | A |
| 40 | 29.2 | 1.75 | 0.22 | A |
| 42 | 28.9 | 7.08 | 0.17 | A |
| 43 | 30.4 | 1.8 | 0.41 | A |
| 44 | 28.3 | 1.5 | 0.08 | A |
| 45 | 37.3 | 2.78 | 1.48 | A |
| 46 | 16.7 | 0.6 | 1.73 | A |
| 47 | 16.0 | 0.6 | 1.84 | A |
| 49 | 27.9 | 9 | 0.02 | A |
| 50 | 28.7 | 9 | 0.14 | A |
| 52 | 31.2 | 2.1 | 0.53 | A |
| 53 | 141.7 | 0.11 | 17.80 | N |
| 54 | 25.5 | 2.4 | 0.36 | A |
| 55 | 30.5 | 2.2 | 0.42 | A |
| 56 | 19.5 | 1.6 | 1.30 | A |
| 57 | 32.4 | 0.8 | 0.72 | A |
| 58 | 26.8 | 4.3 | 0.16 | A |
| 65 | 25.82 | 15.49 | 0.31 | A |
| 67 | 37.84 | 2 | 1.57 | A |
| 68 | 10.83 | 0.26 | 2.65 | W |
| 70 | 23.4 | 2.2 | 0.69 | A |
| 72 | 34.2 | 3.2 | 1.00 | A |
| 74 | 15.46 | 0.22 | 1.93 | A |
| 75 | 14.81 | 0.23 | 2.03 | W |
| 78 | 32.22 | 4.46 | 0.69 | A |
| 79 | 45.742 | 1.351 | 2.80 | W |
| 82 | 17.17 | 1.88 | 1.66 | A |
| 83 | 24.78 | 2.72 | 0.47 | A |
| 84 | 26.3 | 3.0 | 0.23 | A |
| 85 | 26.74 | 5.35 | 0.17 | A |
| 86 | 28.7 | 1.5 | 0.14 | A |
| 87 | 21.51 | 1.60 | 0.98 | A |
| 89 | 27.4 | 3.5 | 0.06 | A |
| 90 | 36.03 | 8.03 | 1.29 | A |
| 91 | 8.45 | 0.26 | 3.02 | N |
| 93 | 21.98 | 1.24 | 0.91 | A |
| 95 | 29.0 | 2.4 | 0.19 | A |
| 96 | 19.701 | 0.323 | 1.27 | A |
| 97 | 16.24 | 0.63 | 1.81 | A |
| 98 | 32.37 | 3.9 | 0.71 | A |
| 99 | 20.4 | 1.5 | 1.16 | A |
| 101 | 28.0 | 1.8 | 0.03 | A |
| 102 | 33.14 | 3.38 | 0.83 | A |
| 104 | 26.17 | 1.11 | 0.25 | A |
| 106 | 22.55 | 4.67 | 0.82 | A |
| 112 | 28 | 4 | 0.03 | A |
| 115 | 40.37 | 3.54 | 1.96 | A |
| 119 | 25.7 | 0.7 | 0.33 | A |
| 120 | 28.23 | 6.66 | 0.07 | A |
| 123 | 25.3 | 1.4 | 0.39 | A |
| 124 | 28.90 | 0.93 | 0.17 | A |
| 126 | 26.0 | 2.1 | 0.28 | A |
| 127 | 24 | 1.87 | 0.59 | A |
| 129 | 28.67 | 0.25 | 0.14 | A |
| 130 | 26.5 | 4.6 | 0.20 | A |
| 131 | 21.82 | 8.21 | 0.93 | A |
| 133 | 30.7 | 1.9 | 0.45 | A |
| 135 | 31.0 | 6.2 | 0.50 | A |
| 136 | 31.8 | 0.5 | 0.62 | A |
| 144 | 27.1 | 2.0 | 0.11 | A |
| 145 | 27.0 | 2.1 | 0.13 | A |

TABLE 65. IC Evaluation Results for gross alpha in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 146 | 27.2 | 1.6 | 0.09 | A |
| 147 | 21.55 | 1.57 | 0.98 | A |
| 149 | 30.78 | 0.11 | 0.47 | A |
| 151 | 27.73 | 2.25 | 0.01 | A |
| 152 | 18.7 |  | 1.42 | A |
| 154 | 28.18 | 2.8 | 0.06 | A |
| 155 | 33.1 | 0.7 | 0.83 | A |
| 157 | 28.1 | 4.4 | 0.05 | A |
| 158 | 26.04 | 3.91 | 0.28 | A |
| 159 | 33.19 | 0.73 | 0.84 | A |
| 160 | 33.5 | 2.1 | 0.89 | A |
| 164 | 36.59 | 18.32 | 1.37 | A |
| 165 | 19.67 | 1.29 | 1.27 | A |
| 167 | 36.80 | 5.85 | 1.41 | A |
| 169 | 21.22 | 7.34 | 1.03 | A |
| 173 | 20.3 | 3.65 | 1.17 | A |
| 175 | 28.8 | 2.1 | 0.16 | A |
| 180 | 28.27 | 2.63 | 0.07 | A |
| 181 | 31.1 | 3.1 | 0.52 | A |
| 184 | 33.3 | 6.0 | 0.86 | A |
| 185 | 28.74 | 4.17 | 0.15 | A |
| 186 | 37.0 | 1.1 | 1.44 | A |
| 187 | 32.9 | 1.7 | 0.80 | A |
| 188 | 31.0 | 2.9 | 0.50 | A |
| 189 | 35.3 | 4.3 | 1.17 | A |
| 195 | 27.6 | 3.5 | 0.03 | A |
| 196 | 32.12 | 3.6 | 0.67 | A |
| 197 | 20.8557 | 2.2239 | 1.09 | A |
| 198 | 26.1 | 0.3 | 0.27 | A |
| 199 | 24.64 | 0.49 | 0.49 | A |
| 201 | 23.4 | 2.3 | 0.69 | A |
| 202 | 15.02 | 1.03 | 2.00 | A |
| 203 | 33.96 | 1.31 | 0.96 | A |
| 204 | 30.5 | 2.5 | 0.42 | A |
| 206 | 23.82 | 0.25 | 0.62 | A |
| 207 | 33.30 | 1.45 | 0.86 | A |
| 210 | 10.1 | 0.5 | 2.77 | W |
| 214 | 34 | 2 | 0.97 | A |
| 218 | 33.2 | 4.35 | 0.84 | A |
| 219 | 22.47 | 2.42 | 0.83 | A |
| 221 | 37.00 | 2.50 | 1.44 | A |
| 223 | 15.68 | 0.84 | 1.89 | A |
| 224 | 0.07 | 0.02 | 4.33 | N |
| 225 | 32.1 | 1.5 | 0.67 | A |
| 229 | 9.82 | 3.0 | 2.81 | W |
| 232 | 25.5 | 2.9 | 0.36 | A |
| 235 | 34.3 | 5.8 | 1.02 | A |
| 241 | 31.02 | 0.28 | 0.50 | A |
| 245 | 32.1 | 2.8 | 0.67 | A |
| 247 | 28.6 | 1.5 | 0.13 | A |
| 248 | 28.68 | 2.36 | 0.14 | A |
| 249 | 30.1 | 0.4 | 0.36 | A |
| 259 | 32.1 | 2 | 0.67 | A |
| 260 | 11.10 | 1.42 | 2.61 | W |
| 261 | 30.6 | 3.8 | 0.44 | A |
| 262 | 34.0 | 3.52 | 0.97 | A |
| 263 | 29.9 | 0.3 | 0.33 | A |
| 266 | 37.4 | 7.79 | 1.50 | A |
| 268 | 24.6 | 1.3 | 0.50 | A |
| 270 | 33.2 | 2.7 | 0.84 | A |
| 272 | 41.606 | 2.244 | 2.16 | W |

TABLE 65. IC Evaluation Results for gross alpha in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 277 | 26.30 | 1.32 | 0.23 | A |
| 279 | 25.1 | 1.88 | 0.42 | A |
| 288 | 18.40 | 0.15 | 1.47 | A |
| 290 | 27.8 | 1.23 | 0.00 | A |
| 291 | 29.8 | 2.5 | 0.31 | A |
| 293 | 29.3 | 2.9 | 0.23 | A |
| 295 | 15.4 | 0.6 | 1.94 | A |
| 296 | 24.82 | 1.38 | 0.47 | A |
| 299 | 27.8 | 1.03 | 0.00 | A |
| 302 | 0.08 | 0.002 | 4.33 | N |
| 304 | 17.6 | 1.3 | 1.59 | A |
| 307 | 33.26 | 1.66 | 0.85 | A |
| 310 | 2.188 | 0.0342 | 4.00 | N |
| 312 | 27.78 | 3.417 | 0.00 | A |
| 313 | 11.55 | 0.18 | 2.54 | W |
| 315 | 27.9 | 1.0 | 0.02 | A |
| 317 | 28.786 | 0.45 | 0.15 | A |
| 319 | 22.125 | 2.873 | 0.89 | A |
| 321 | 25.51 | 2.05 | 0.36 | A |
| 322 | 6.93 | 0.83 | 3.26 | N |
| 330 | 33.5 | 4.5 | 0.89 | A |
| 333 | 21.9 | 0.8 | 0.92 | A |
| 334 | 32 | 7 | 0.66 | A |
| 341 | 6.94 | 1.86 | 3.26 | N |
| 345 | 11.6052 | 1.87 | 2.53 | W |

Activity [ Bq/kg ]

15

20

25

30

Reported Results for gross beta Determination in Sample 1

TABLE 66. Target values

| Parameter | Value |
| --- | --- |
| Robust Mean, Bq/kg | 17.3 |
| Robust SD, Bq/kg | 5.8 |
| N | 240 |

5

10

302

341

224

322

313

259

151

68

319

131

277

232

65

229

89

185

91

12

99

149

123

11

175

321

270

43

93

16

164

49

47

199

201

56

46

345

130

87

96

50

101

136

95

83

310

291

126

133

127

206

45

97

36

290

39

32

235

157

124

159

58

245

84

112

247

315

55

5

223

262

203

141

180

188

173

90

52

1

72

181

198

120

27

241

193

260

279

296

186

210

29

189

17

34

119

82

197

248

160

219

75

79

330

42

104

37

14

33

25

344

57

137

94

54

31

3

70

86

337

78

67

44

195

288

85

145

35

202

196

146

19

218

304

167

299

158

184

187

169

22

6

102

115

165

26

155

28

144

263

298

266

214

293

333

147

204

2

129

317

307

66

4

154

261

268

334

249

225

272

53

221

Labcode

FIGURE 62. Reported results and evaluation parameters for gross beta in Sample 1.

TABLE 67. IC Evaluation Results for gross beta in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 1 | 16.1 | 1.2 | 0.21 | A |
| 2 | 31.90 | 2.61 | 2.52 | W |
| 3 | 19.5 | 1.98 | 0.38 | A |
| 4 | 38.20 | 0.90 | 3.60 | N |
| 5 | 15.39 | 0.80 | 0.33 | A |
| 6 | 23.0 | 4.5 | 0.98 | A |
| 11 | 11.06 | 0.72 | 1.08 | A |
| 12 | 10.5 | 3.8 | 1.17 | A |
| 14 | 18.3 | 0.3 | 0.17 | A |
| 16 | 11.71 | 0.45 | 0.96 | A |
| 17 | 17.1 | 1.9 | 0.03 | A |
| 19 | 21.1 | 3.0 | 0.66 | A |
| 22 | 22.96 | 1.38 | 0.98 | A |
| 25 | 18.90 | 2.9 | 0.28 | A |
| 26 | 24.5 | 8.11 | 1.24 | A |
| 27 | 16.5 | 3.2 | 0.14 | A |
| 28 | 25.3 | 3.9 | 1.38 | A |
| 29 | 17 | 1 | 0.05 | A |
| 31 | 19.5 | 6.1 | 0.38 | A |
| 32 | 14.56 | 0.31 | 0.47 | A |
| 33 | 18.5 | 2.5 | 0.21 | A |
| 34 | 17.1 | 1.8 | 0.03 | A |
| 35 | 20.9 | 3.1 | 0.62 | A |
| 36 | 14.5 | 1.0 | 0.48 | A |
| 37 | 18.3 | 3.4 | 0.17 | A |

TABLE 67. IC Evaluation Results for gross beta in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 39 | 14.5 | 0.8 | 0.48 | A |
| 42 | 18.1 | 4.43 | 0.14 | A |
| 43 | 11.6 | 0.6 | 0.98 | A |
| 44 | 20.5 | 0.8 | 0.55 | A |
| 45 | 14.2 | 1.05 | 0.53 | A |
| 46 | 12.4 | 0.5 | 0.84 | A |
| 47 | 12.0 | 0.5 | 0.91 | A |
| 49 | 12 | 2.8 | 0.91 | A |
| 50 | 12.8 | 3 | 0.78 | A |
| 52 | 16.0 | 1.6 | 0.22 | A |
| 53 | 146.4 | 0.13 | 22.26 | N |
| 54 | 19.4 | 1.6 | 0.36 | A |
| 55 | 15.2 | 1.2 | 0.36 | A |
| 56 | 12.3 | 0.9 | 0.86 | A |
| 57 | 19.3 | 0.9 | 0.34 | A |
| 58 | 14.8 | 2.4 | 0.43 | A |
| 65 | 9.05 | 5.43 | 1.42 | A |
| 66 | 37.4 | 2.22 | 3.47 | N |
| 67 | 20.44 | 0.94 | 0.54 | A |
| 68 | 7.89 | 0.24 | 1.62 | A |
| 70 | 19.65 | 1.8 | 0.41 | A |
| 72 | 16.3 | 1.2 | 0.17 | A |
| 75 | 17.729 | 0.184 | 0.07 | A |
| 78 | 20.04 | 2.82 | 0.47 | A |
| 79 | 17.98 | 1.80 | 0.12 | A |
| 82 | 17.16 | 2.31 | 0.02 | A |
| 83 | 13.63 | 1.5 | 0.63 | A |
| 84 | 14.9 | 0.88 | 0.41 | A |
| 85 | 20.67 | 4.96 | 0.58 | A |
| 86 | 19.8 | 1.1 | 0.43 | A |
| 87 | 12.51 | 0.72 | 0.83 | A |
| 89 | 10.0 | 1.2 | 1.26 | A |
| 90 | 15.91 | 2.06 | 0.24 | A |
| 91 | 10.49 | 1.08 | 1.17 | A |
| 93 | 11.67 | 1.00 | 0.97 | A |
| 94 | 19.39 | 1.96 | 0.36 | A |
| 95 | 13.0 | 1.1 | 0.74 | A |
| 96 | 12.647 | 0.319 | 0.80 | A |
| 97 | 14.25 | 2.25 | 0.53 | A |
| 99 | 10.50 | 0.27 | 1.17 | A |
| 101 | 12.8 | 0.4 | 0.78 | A |
| 102 | 23.36 | 3.71 | 1.04 | A |
| 104 | 18.21 | 0.65 | 0.16 | A |
| 112 | 15 | 2 | 0.40 | A |
| 115 | 23.63 | 1.19 | 1.09 | A |
| 119 | 17.1 | 0.5 | 0.03 | A |
| 120 | 16.48 | 3.86 | 0.14 | A |
| 123 | 10.8 | 1.1 | 1.12 | A |
| 124 | 14.72 | 0.20 | 0.44 | A |
| 126 | 13.9 | 0.14 | 0.59 | A |
| 127 | 14.10 | 0.23 | 0.55 | A |
| 129 | 32.29 | 0.35 | 2.58 | W |
| 130 | 12.5 | 1.9 | 0.83 | A |
| 131 | 8.21 | 0.61 | 1.57 | A |
| 133 | 14.1 | 0.8 | 0.55 | A |
| 136 | 12.8 | 0.3 | 0.78 | A |
| 137 | 19.39 | 1.96 | 0.36 | A |
| 141 | 15.61 | 1.17 | 0.29 | A |
| 144 | 25.4 | 1.9 | 1.40 | A |
| 145 | 20.7 | 1.8 | 0.59 | A |
| 146 | 21.0 | 1.17 | 0.64 | A |

TABLE 67. IC Evaluation Results for gross beta in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 147 | 30.99 | 2.24 | 2.36 | W |
| 149 | 10.66 | 0.01 | 1.14 | A |
| 151 | 7.63 | 0.33 | 1.67 | A |
| 154 | 38.5 | 3.85 | 3.66 | N |
| 155 | 24.7 | 2.0 | 1.28 | A |
| 157 | 14.7 | 3.0 | 0.45 | A |
| 158 | 22.17 | 3.33 | 0.84 | A |
| 159 | 14.73 | 0.78 | 0.44 | A |
| 160 | 17.4 | 1.3 | 0.02 | A |
| 164 | 11.92 | 5.96 | 0.93 | A |
| 165 | 23.72 | 4.78 | 1.11 | A |
| 167 | 22.10 | 1.43 | 0.83 | A |
| 169 | 22.88 | 5.57 | 0.96 | A |
| 173 | 15.8 | 2.6 | 0.26 | A |
| 175 | 11.2 | 0.8 | 1.05 | A |
| 180 | 15.78 | 0.93 | 0.26 | A |
| 181 | 16.3 | 1.6 | 0.17 | A |
| 184 | 22.6 | 4.1 | 0.91 | A |
| 185 | 10.21 | 1.84 | 1.22 | A |
| 186 | 16.9 | 0.4 | 0.07 | A |
| 187 | 22.8 | 1.3 | 0.95 | A |
| 188 | 15.8 | 1.5 | 0.26 | A |
| 189 | 17 | 2 | 0.05 | A |
| 193 | 16.79 | 0.84 | 0.09 | A |
| 195 | 20.6 | 2.7 | 0.57 | A |
| 196 | 20.95 | 2.3 | 0.63 | A |
| 197 | 17.2595 | 3.1245 | 0.01 | A |
| 198 | 16.3 | 0.4 | 0.17 | A |
| 199 | 12.03 | 0.86 | 0.91 | A |
| 201 | 12.1 | 1.0 | 0.90 | A |
| 202 | 20.91 | 4.97 | 0.62 | A |
| 203 | 15.46 | 0.60 | 0.32 | A |
| 204 | 31.2 | 3.9 | 2.40 | W |
| 206 | 14.20 | 0.15 | 0.53 | A |
| 210 | 16.9 | 1.3 | 0.07 | A |
| 214 | 28 | 3 | 1.84 | A |
| 218 | 21.5 | 3.61 | 0.72 | A |
| 219 | 17.48 | 1.8 | 0.03 | A |
| 221 | 273.42 | 0.05 | 44.16 | N |
| 223 | 15.39 | 0.84 | 0.33 | A |
| 224 | 1.165 | 0.075 | 2.78 | W |
| 225 | 55.5 | 5.7 | 6.59 | N |
| 229 | 9.60 | 2.85 | 1.33 | A |
| 232 | 9 | 1 | 1.43 | A |
| 235 | 14.6 | 1.7 | 0.47 | A |
| 241 | 16.67 | 0.22 | 0.11 | A |
| 245 | 14.8 | 1.3 | 0.43 | A |
| 247 | 15.0 | 3.9 | 0.40 | A |
| 248 | 17.34 | 1.45 | 0.01 | A |
| 249 | 54.4 | 0.9 | 6.40 | N |
| 259 | 6.9 | 0.9 | 1.79 | A |
| 260 | 16.81 | 2.02 | 0.08 | A |
| 261 | 40.3 | 5.5 | 3.97 | N |
| 262 | 15.4 | 1.19 | 0.33 | A |
| 263 | 25.9 | 0.3 | 1.48 | A |
| 266 | 27.9 | 8.21 | 1.83 | A |
| 268 | 41.2 | 2.5 | 4.12 | N |
| 270 | 11.51 | 0.73 | 1.00 | A |
| 272 | 66.77 | 3.07 | 8.53 | N |
| 277 | 8.57 | 0.29 | 1.51 | A |
| 279 | 16.85 | 0.8425 | 0.08 | A |

TABLE 67. IC Evaluation Results for gross beta in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 288 | 20.6 | 0.18 | 0.57 | A |
| 290 | 14.5 | 0.65 | 0.48 | A |
| 291 | 13.8 | 1.7 | 0.60 | A |
| 293 | 28.4 | 2.8 | 1.91 | A |
| 296 | 16.89 | 1.01 | 0.07 | A |
| 298 | 27.2 | 1.7 | 1.71 | A |
| 299 | 22.1 | 0.7 | 0.83 | A |
| 302 | 0.3 | 0.012 | 2.93 | W |
| 304 | 21.6 | 0.9 | 0.74 | A |
| 307 | 36.44 | 3.78 | 3.30 | N |
| 310 | 13.751 | 0.1104 | 0.61 | A |
| 313 | 5.88 | 0.07 | 1.97 | A |
| 315 | 15.1 | 0.5 | 0.38 | A |
| 317 | 35.96 | 0.49 | 3.22 | N |
| 319 | 8.195 | 1.065 | 1.57 | A |
| 321 | 11.27 | 0.65 | 1.04 | A |
| 322 | 4.47 | 0.09 | 2.21 | W |
| 330 | 18.0 | 3.6 | 0.12 | A |
| 333 | 28.9 | 0.9 | 2.00 | A |
| 334 | 48 | 10 | 5.29 | N |
| 337 | 19.94 | 1.5 | 0.46 | A |
| 341 | 0.74 | 0.2 | 2.86 | W |
| 344 | 19.2 | 6.8 | 0.33 | A |
| 345 | 12.4457 | 0.43 | 0.84 | A |

Activity [ Bq/kg ]

15

20

Reported Results for gross alpha Determination in Sample 2

TABLE 68. Target values

| Parameter | Value |
| --- | --- |
| Robust Mean, Bq/kg | 13.2 |
| Robust SD, Bq/kg | 6.1 |
| N | 226 |

5

10

224

302

310

91

210

345

260

97

229

202

102

1

154

333

68

322

295

70

53

223

152

304

93

74

16

89

197

75

219

99

277

96

180

319

87

206

82

42

19

54

247

101

199

321

198

27

201

268

84

225

334

17

290

127

85

146

279

56

39

119

266

32

67

34

159

288

43

133

189

218

49

221

272

296

45

158

235

131

149

112

40

50

58

157

313

52

312

262

181

26

106

167

291

57

72

44

115

28

129

173

4

164

270

35

136

165

196

22

79

195

3

293

160

120

317

185

248

33

203

188

145

232

307

37

147

98

299

86

186

124

12

104

123

25

151

245

214

261

184

130

175

6

31

155

14

330

78

144

47

29

2

207

187

90

126

204

65

46

241

249

263

Labcode

FIGURE 63. Reported results and evaluation parameters for gross alpha in Sample 2.

TABLE 69. IC Evaluation Results for gross alpha in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 1 | 5.0 | 0.4 | 1.34 | A |
| 2 | 25.60 | 1.73 | 2.03 | W |
| 3 | 15.4 | 0.92 | 0.36 | A |
| 4 | 14.40 | 0.32 | 0.20 | A |
| 6 | 21.0 | 0.4 | 1.28 | A |
| 12 | 18.7 | 6.5 | 0.90 | A |
| 14 | 23.1 | 1.5 | 1.62 | A |
| 16 | 6.72 | 0.31 | 1.06 | A |
| 17 | 10.2 | 1.5 | 0.49 | A |
| 19 | 8.40 | 1.20 | 0.79 | A |
| 22 | 14.84 | 0.89 | 0.27 | A |
| 25 | 19.4 | 3.0 | 1.02 | A |
| 26 | 12.7 | 1.00 | 0.08 | A |
| 27 | 9.6 | 1.1 | 0.59 | A |
| 28 | 14.1 | 2.1 | 0.15 | A |
| 29 | 25 | 2 | 1.93 | A |
| 31 | 21.9 | 3.9 | 1.43 | A |
| 32 | 10.90 | 0.39 | 0.38 | A |
| 33 | 16.3 | 4.4 | 0.51 | A |
| 34 | 11 | 1.5 | 0.36 | A |
| 35 | 14.7 | 1.5 | 0.25 | A |
| 37 | 17.8 | 2.9 | 0.75 | A |
| 39 | 10.5 | 0.5 | 0.44 | A |
| 40 | 12.1 | 0.726 | 0.18 | A |
| 42 | 8.2 | 2.00 | 0.82 | A |

TABLE 69. IC Evaluation Results for gross alpha in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 43 | 11.4 | 0.8 | 0.30 | A |
| 44 | 13.6 | 0.8 | 0.07 | A |
| 45 | 11.8 | 0.93 | 0.23 | A |
| 46 | 29.8 | 0.9 | 2.72 | W |
| 47 | 24.9 | 0.8 | 1.92 | A |
| 49 | 11.6 | 3.7 | 0.26 | A |
| 50 | 12.1 | 4 | 0.18 | A |
| 52 | 12.5 | 0.9 | 0.11 | A |
| 53 | 5.7 | 0.08 | 1.23 | A |
| 54 | 8.93 | 0.89 | 0.70 | A |
| 56 | 10.5 | 1.4 | 0.44 | A |
| 57 | 13.2 | 0.4 | 0.00 | A |
| 58 | 12.20 | 1.96 | 0.16 | A |
| 65 | 29.78 | 17.87 | 2.72 | W |
| 67 | 10.9 | 0.96 | 0.38 | A |
| 68 | 5.43 | 0.13 | 1.27 | A |
| 70 | 5.55 | 0.46 | 1.25 | A |
| 72 | 13.31 | 1.39 | 0.02 | A |
| 74 | 6.69 | 0.13 | 1.07 | A |
| 75 | 6.82 | 0.14 | 1.05 | A |
| 78 | 24.51 | 3.43 | 1.85 | A |
| 79 | 14.902 | 0.583 | 0.28 | A |
| 82 | 8.11 | 1.37 | 0.83 | A |
| 84 | 9.87 | 1.15 | 0.55 | A |
| 85 | 10.39 | 2.08 | 0.46 | A |
| 86 | 18.1 | 1.2 | 0.80 | A |
| 87 | 7.89 | 0.98 | 0.87 | A |
| 89 | 6.72 | 0.90 | 1.06 | A |
| 90 | 26.96 | 6.74 | 2.26 | W |
| 91 | 2.85 | 0.15 | 1.70 | A |
| 93 | 6.49 | 0.39 | 1.10 | A |
| 96 | 7.319 | 0.129 | 0.96 | A |
| 97 | 4.27 | 0.24 | 1.46 | A |
| 98 | 17.83 | 2.14 | 0.76 | A |
| 99 | 6.95 | 0.42 | 1.02 | A |
| 101 | 9.0 | 0.7 | 0.69 | A |
| 102 | 4.99 | 0.49 | 1.35 | A |
| 104 | 18.8 | 1.03 | 0.92 | A |
| 106 | 12.9 | 2.8 | 0.05 | A |
| 112 | 12 | 2 | 0.20 | A |
| 115 | 13.95 | 2.17 | 0.12 | A |
| 119 | 10.7 | 0.4 | 0.41 | A |
| 120 | 15.63 | 3.69 | 0.40 | A |
| 123 | 19.1 | 1.1 | 0.97 | A |
| 124 | 18.60 | 0.59 | 0.89 | A |
| 126 | 27.0 | 1.1 | 2.26 | W |
| 127 | 10.3 | 0.82 | 0.48 | A |
| 129 | 14.23 | 0.91 | 0.17 | A |
| 130 | 20.9 | 3.7 | 1.26 | A |
| 131 | 11.96 | 1.07 | 0.20 | A |
| 133 | 11.5 | 0.8 | 0.28 | A |
| 136 | 14.7 | 0.3 | 0.25 | A |
| 144 | 24.7 | 1.8 | 1.89 | A |
| 145 | 16.9 | 1.2 | 0.61 | A |
| 146 | 10.4 | 0.85 | 0.46 | A |
| 147 | 17.80 | 1.47 | 0.75 | A |
| 149 | 12.00 | 0.08 | 0.20 | A |
| 151 | 19.60 | 1.7 | 1.05 | A |
| 152 | 6.07 |  | 1.17 | A |
| 154 | 5.1 | 0.5 | 1.33 | A |
| 155 | 23.0 | 0.6 | 1.61 | A |

TABLE 69. IC Evaluation Results for gross alpha in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 157 | 12.3 | 2.0 | 0.15 | A |
| 158 | 11.82 | 1.77 | 0.23 | A |
| 159 | 11.09 | 0.43 | 0.35 | A |
| 160 | 15.5 | 1.1 | 0.38 | A |
| 164 | 14.6959 | 7.3578 | 0.25 | A |
| 165 | 14.72 | 1.1 | 0.25 | A |
| 167 | 13.00 | 2.07 | 0.03 | A |
| 173 | 14.25 | 1.9 | 0.17 | A |
| 175 | 20.9 | 1.5 | 1.26 | A |
| 180 | 7.64 | 0.78 | 0.91 | A |
| 181 | 12.6 | 1.3 | 0.10 | A |
| 184 | 20.4 | 4.7 | 1.18 | A |
| 185 | 16.08 | 2.25 | 0.47 | A |
| 186 | 18.5 | 0.6 | 0.87 | A |
| 187 | 26.3 | 1.3 | 2.15 | W |
| 188 | 16.6 | 2.6 | 0.56 | A |
| 189 | 11.6 | 1.5 | 0.26 | A |
| 195 | 15.1 | 1.6 | 0.31 | A |
| 196 | 14.73 | 1.5 | 0.25 | A |
| 197 | 6.7358 | 0.9581 | 1.06 | A |
| 198 | 9.5 | 0.1 | 0.61 | A |
| 199 | 9.25 | 0.41 | 0.65 | A |
| 201 | 9.7 | 0.9 | 0.57 | A |
| 202 | 4.70 | 0.68 | 1.39 | A |
| 203 | 16.41 | 1.06 | 0.53 | A |
| 204 | 27.4 | 2.3 | 2.33 | W |
| 206 | 7.95 | 0.14 | 0.86 | A |
| 207 | 26.07 | 1.44 | 2.11 | W |
| 210 | 3.45 | 0.2 | 1.60 | A |
| 214 | 20 | 2 | 1.11 | A |
| 218 | 11.6 | 1.51 | 0.26 | A |
| 219 | 6.85 | 0.67 | 1.04 | A |
| 221 | 11.61 | 0.80 | 0.26 | A |
| 223 | 5.97 | 0.53 | 1.19 | A |
| 224 | 0.04 | 0.01 | 2.16 | W |
| 225 | 9.9 | 1.0 | 0.54 | A |
| 229 | 4.35 | 1.00 | 1.45 | A |
| 232 | 17.1 | 2.1 | 0.64 | A |
| 235 | 11.9 | 2.2 | 0.21 | A |
| 241 | 31.42 | 0.28 | 2.99 | W |
| 245 | 19.6 | 1.7 | 1.05 | A |
| 247 | 8.96 | 0.90 | 0.70 | A |
| 248 | 16.19 | 0.81 | 0.49 | A |
| 249 | 32.5 | 0.4 | 3.16 | N |
| 260 | 3.73 | 1.18 | 1.55 | A |
| 261 | 20.3 | 2.5 | 1.16 | A |
| 262 | 12.6 | 1.37 | 0.10 | A |
| 263 | 160 | 1 | 24.07 | N |
| 266 | 10.8 | 2.76 | 0.39 | A |
| 268 | 9.78 | 0.53 | 0.56 | A |
| 270 | 14.7 | 1.2 | 0.25 | A |
| 272 | 11.626 | 1.311 | 0.26 | A |
| 277 | 7.22 | 0.33 | 0.98 | A |
| 279 | 10.4 | 0.98 | 0.46 | A |
| 288 | 11.2 | 0.95 | 0.33 | A |
| 290 | 10.3 | 0.91 | 0.48 | A |
| 291 | 13.1 | 1.5 | 0.02 | A |
| 293 | 15.4 | 1.5 | 0.36 | A |
| 295 | 5.5 | 0.2 | 1.26 | A |
| 296 | 11.64 | 0.81 | 0.26 | A |
| 299 | 17.9 | 0.7 | 0.77 | A |

TABLE 69. IC Evaluation Results for gross alpha in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 302 | 0.28 | 0.043 | 2.12 | W |
| 304 | 6.1 | 0.4 | 1.16 | A |
| 307 | 17.26 | 1.26 | 0.67 | A |
| 310 | 0.8039 | 0.0164 | 2.03 | W |
| 312 | 12.54 | 2.229 | 0.11 | A |
| 313 | 12.34 | 0.20 | 0.14 | A |
| 317 | 16.057 | 0.33 | 0.47 | A |
| 319 | 7.816 | 0.940 | 0.88 | A |
| 321 | 9.44 | 0.68 | 0.62 | A |
| 322 | 5.49 | 1.21 | 1.26 | A |
| 330 | 23.4 | 4.5 | 1.67 | A |
| 333 | 5.1 | 0.2 | 1.33 | A |
| 334 | 10 | 3 | 0.52 | A |
| 345 | 3.5676 | 0.41 | 1.58 | A |

Activity [ Bq/kg ]

150

200

250

Reported Results for gross beta Determination in Sample 2

TABLE 70. Target values

| Parameter | Value |
| --- | --- |
| Robust Mean, Bq/kg | 135 |
| Robust SD, Bq/kg | 42 |
| N | 231 |

50

100

302

224

263

70

310

322

229

68

319

321

313

47

97

151

175

131

87

202

157

164

46

115

247

96

91

89

99

210

65

93

197

56

188

290

201

198

219

185

12

232

95

199

293

67

50

43

270

149

75

49

159

345

130

304

58

260

154

184

136

180

145

223

291

193

181

42

5

133

123

241

112

25

173

141

124

195

32

299

158

126

277

167

119

279

35

16

206

101

53

203

84

44

127

120

146

186

72

39

86

54

298

85

1

235

337

104

330

137

94

262

196

218

82

268

37

17

19

3

344

333

189

248

160

296

90

14

33

165

29

102

57

78

34

4

187

26

45

79

334

288

245

31

66

155

307

52

266

147

28

6

214

317

27

22

129

249

261

144

272

2

204

225

221

Labcode

FIGURE 64. Reported results and evaluation parameters for gross beta in Sample 2.

TABLE 71. IC Evaluation Results for gross beta in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 1 | 144.2 | 9.4 | 0.22 | A |
| 2 | 222.71 | 18.03 | 2.09 | W |
| 3 | 156.9 | 4.03 | 0.52 | A |
| 4 | 174.56 | 1.79 | 0.94 | A |
| 5 | 127.27 | 3.27 | 0.18 | A |
| 6 | 200.0 | 13.0 | 1.55 | A |
| 12 | 109 | 38 | 0.62 | A |
| 14 | 161.7 | 2.1 | 0.64 | A |
| 16 | 135.62 | 4.44 | 0.01 | A |
| 17 | 154.4 | 13.0 | 0.46 | A |
| 19 | 155 | 22 | 0.48 | A |
| 22 | 207.49 | 12.45 | 1.73 | A |
| 25 | 131 | 20.0 | 0.10 | A |
| 26 | 176 | 14.8 | 0.98 | A |
| 27 | 203 | 8 | 1.62 | A |
| 28 | 199.7 | 30.0 | 1.54 | A |
| 29 | 166 | 5 | 0.74 | A |
| 31 | 187 | 54 | 1.24 | A |
| 32 | 133.08 | 0.87 | 0.05 | A |
| 33 | 162.8 | 21.8 | 0.66 | A |
| 34 | 173 | 17 | 0.90 | A |
| 35 | 135 | 18 | 0.00 | A |
| 37 | 154 | 21 | 0.45 | A |
| 39 | 142.3 | 5.0 | 0.17 | A |
| 42 | 126.8 | 31.07 | 0.20 | A |

TABLE 71. IC Evaluation Results for gross beta in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 43 | 116 | 5 | 0.45 | A |
| 44 | 138 | 4 | 0.07 | A |
| 45 | 176 | 12.8 | 0.98 | A |
| 46 | 83.7 | 1.6 | 1.22 | A |
| 47 | 74.3 | 1.5 | 1.45 | A |
| 49 | 118 | 27 | 0.40 | A |
| 50 | 115 | 26.5 | 0.48 | A |
| 52 | 193 | 18 | 1.38 | A |
| 53 | 137.2 | 0.01 | 0.05 | A |
| 54 | 142.5 | 4.6 | 0.18 | A |
| 56 | 99.5 | 7.1 | 0.85 | A |
| 57 | 171.3 | 3.6 | 0.86 | A |
| 58 | 120.10 | 19.22 | 0.35 | A |
| 65 | 98.02 | 58.81 | 0.88 | A |
| 66 | 189.0 | 5.63 | 1.29 | A |
| 67 | 114.66 | 1.91 | 0.48 | A |
| 68 | 62.6 | 3.4 | 1.72 | A |
| 70 | 21.1 | 1.4 | 2.71 | W |
| 72 | 142.3 | 10.6 | 0.17 | A |
| 75 | 117.222 | 0.460 | 0.42 | A |
| 78 | 172.31 | 22.16 | 0.89 | A |
| 79 | 176 | 17.6 | 0.98 | A |
| 82 | 151.29 | 9.15 | 0.39 | A |
| 84 | 138 | 7.8 | 0.07 | A |
| 85 | 144.16 | 34.61 | 0.22 | A |
| 86 | 142.4 | 4.8 | 0.18 | A |
| 87 | 79.07 | 1.56 | 1.33 | A |
| 89 | 90.90 | 10.5 | 1.05 | A |
| 90 | 161.24 | 30.63 | 0.62 | A |
| 91 | 90.52 | 1.51 | 1.06 | A |
| 93 | 98.26 | 8.36 | 0.87 | A |
| 94 | 146.55 | 14.69 | 0.28 | A |
| 95 | 111 | 9 | 0.57 | A |
| 96 | 87.902 | 2.202 | 1.12 | A |
| 97 | 74.34 | 2.45 | 1.44 | A |
| 99 | 91.5 | 1.7 | 1.04 | A |
| 101 | 136.7 | 4.0 | 0.04 | A |
| 102 | 168.1 | 13.9 | 0.79 | A |
| 104 | 144.98 | 2.05 | 0.24 | A |
| 112 | 130 | 10 | 0.12 | A |
| 115 | 86.73 | 3.73 | 1.15 | A |
| 119 | 135 | 1.5 | 0.00 | A |
| 120 | 141.8 | 33.39 | 0.16 | A |
| 123 | 129 | 13 | 0.14 | A |
| 124 | 131.78 | 0.85 | 0.08 | A |
| 126 | 134 | 0.9 | 0.02 | A |
| 127 | 141 | 0.235 | 0.14 | A |
| 129 | 207.99 | 3.69 | 1.74 | A |
| 130 | 119 | 18 | 0.38 | A |
| 131 | 77.74 | 5.25 | 1.36 | A |
| 133 | 129 | 7 | 0.14 | A |
| 136 | 122.8 | 1.7 | 0.29 | A |
| 137 | 146.55 | 14.69 | 0.28 | A |
| 141 | 131.72 | 5.98 | 0.08 | A |
| 144 | 216 | 16 | 1.93 | A |
| 145 | 124 | 7 | 0.26 | A |
| 146 | 142 | 6 | 0.17 | A |
| 147 | 199.16 | 14.21 | 1.53 | A |
| 149 | 116.84 | 0.45 | 0.43 | A |
| 151 | 75.51 | 3.30 | 1.42 | A |
| 154 | 121.7 | 12 | 0.32 | A |

TABLE 71. IC Evaluation Results for gross beta in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 155 | 189.0 | 13.3 | 1.29 | A |
| 157 | 83.1 | 9.7 | 1.24 | A |
| 158 | 133.9 | 20.09 | 0.03 | A |
| 159 | 118.0 | 2.2 | 0.40 | A |
| 160 | 160.9 | 9.5 | 0.62 | A |
| 164 | 83.4528 | 41.7728 | 1.23 | A |
| 165 | 165.49 | 73.75 | 0.73 | A |
| 167 | 135.00 | 7.70 | 0.00 | A |
| 173 | 131 | 13.6 | 0.10 | A |
| 175 | 77 | 6 | 1.38 | A |
| 180 | 123.07 | 6.52 | 0.28 | A |
| 181 | 126.5 | 12.7 | 0.20 | A |
| 184 | 122 | 24 | 0.31 | A |
| 185 | 106.97 | 19.26 | 0.67 | A |
| 186 | 142 | 3 | 0.17 | A |
| 187 | 174.7 | 9.7 | 0.95 | A |
| 188 | 101.6 | 6.1 | 0.80 | A |
| 189 | 160 | 18 | 0.60 | A |
| 193 | 125.75 | 6.29 | 0.22 | A |
| 195 | 133 | 17 | 0.05 | A |
| 196 | 148.12 | 13.6 | 0.31 | A |
| 197 | 98.3126 | 7.977 | 0.87 | A |
| 198 | 104.1 | 1.3 | 0.74 | A |
| 199 | 112.20 | 1.41 | 0.54 | A |
| 201 | 102.1 | 7.0 | 0.78 | A |
| 202 | 81.15 | 5.73 | 1.28 | A |
| 203 | 137.95 | 3.52 | 0.07 | A |
| 204 | 224 | 28 | 2.12 | W |
| 206 | 135.65 | 0.38 | 0.02 | A |
| 210 | 91.7 | 5.7 | 1.03 | A |
| 214 | 201 | 9 | 1.57 | A |
| 218 | 148.7 | 10.7 | 0.33 | A |
| 219 | 104.28 | 10.5 | 0.73 | A |
| 221 | 2727.00 | 0.30 | 61.71 | N |
| 223 | 124.32 | 2.10 | 0.25 | A |
| 224 | 1.565 | 0.008 | 3.18 | N |
| 225 | 242 | 12 | 2.55 | W |
| 229 | 56.98 | 9.88 | 1.86 | A |
| 232 | 110 | 10 | 0.60 | A |
| 235 | 144.6 | 19.9 | 0.23 | A |
| 241 | 129.08 | 0.66 | 0.14 | A |
| 245 | 182.5 | 14.4 | 1.13 | A |
| 247 | 87.6 | 4.7 | 1.13 | A |
| 248 | 160.43 | 7.59 | 0.61 | A |
| 249 | 212.2 | 1.1 | 1.84 | A |
| 260 | 120.64 | 14.02 | 0.34 | A |
| 261 | 215 | 29 | 1.90 | A |
| 262 | 148 | 7.53 | 0.31 | A |
| 263 | 19.5 | 0.3 | 2.75 | W |
| 266 | 198 | 57.8 | 1.50 | A |
| 268 | 152 | 8.1 | 0.40 | A |
| 270 | 116.8 | 6.4 | 0.43 | A |
| 272 | 220.94 | 10.02 | 2.05 | W |
| 277 | 134.00 | 4.42 | 0.02 | A |
| 279 | 135 | 3.375 | 0.00 | A |
| 288 | 182 | 0.17 | 1.12 | A |
| 290 | 102 | 3.66 | 0.79 | A |
| 291 | 125.5 | 8.9 | 0.23 | A |
| 293 | 113 | 11 | 0.52 | A |
| 296 | 161.07 | 8.97 | 0.62 | A |
| 298 | 143 | 9 | 0.19 | A |

TABLE 71. IC Evaluation Results for gross beta in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 299 | 133.5 | 4.3 | 0.04 | A |
| 302 | 1.53 | 0.16 | 3.18 | N |
| 304 | 119.9 | 4.8 | 0.36 | A |
| 307 | 189 | 5.2 | 1.29 | A |
| 310 | 34.158 | 0.110 | 2.40 | W |
| 313 | 72.97 | 0.18 | 1.48 | A |
| 317 | 202.32 | 1.15 | 1.60 | A |
| 319 | 64.653 | 8.405 | 1.67 | A |
| 321 | 68.93 | 3.95 | 1.57 | A |
| 322 | 50.22 | 5.27 | 2.02 | W |
| 330 | 145 | 28 | 0.24 | A |
| 333 | 159.4 | 5.0 | 0.58 | A |
| 334 | 180 | 40 | 1.07 | A |
| 337 | 144.61 | 10.4 | 0.23 | A |
| 344 | 159 | 64 | 0.57 | A |
| 345 | 118.9038 | 3.71 | 0.38 | A |

Activity [ Bq/kg ]

30

40

Reported Results for Th-232 Determination in Sample 4

TABLE 72. Target values

Parameter	Value

50

Target value [Bq/kg]	32.9 Uncertainty [Bq/kg]	2.0

20

172

336

308

202

38

215

108

46

174

231

11

209

43

51

47

293

27

281

311

179

70

55

227

97

86

296

330

48

1

210

186

162

123

246

334

165

89

241

318

Labcode

FIGURE 65. Reported results and evaluation parameters for Th-232 in Sample 4.

TABLE 73. IC Evaluation Results for Th-232 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 1 | 37.4 | 2.4 | 2.25 | W |
| 11 | 32.2 | 1.7 | 0.35 | A |
| 27 | 33.3 | 1.2 | 0.20 | A |
| 38 | 25.1 | 1.6 | 3.90 | N |
| 43 | 32.3 | 3.0 | 0.30 | A |
| 46 | 32 | 1 | 0.45 | A |
| 47 | 33 | 1.4 | 0.05 | A |
| 48 | 37.3 | 3.3 | 2.20 | W |
| 51 | 32.6 | 2.0 | 0.15 | A |
| 55 | 35.6 | 2.3 | 1.35 | A |
| 70 | 35.1 | 3.7 | 1.10 | A |
| 86 | 36.7 | 3.2 | 1.90 | A |
| 89 | 50.39 | 2.5 | 8.75 | N |
| 97 | 35.99 | 1.42 | 1.55 | A |
| 108 | 31 | 2.3 | 0.95 | A |
| 123 | 39.8 | 1.3 | 3.45 | N |
| 162 | 39.7 | 2.3 | 3.40 | N |
| 165 | 42.12 | 1.02 | 4.61 | N |
| 172 | 3.24 | 0.357 | 14.83 | N |
| 174 | 32 | 4 | 0.45 | A |
| 179 | 34.5 | 1.0 | 0.80 | A |
| 186 | 38 | 3 | 2.55 | W |
| 202 | 15.13 | 1.45 | 8.88 | N |
| 209 | 32.23 | 2.63 | 0.34 | A |
| 210 | 37.77 | 2.85 | 2.44 | W |
| 215 | 29.0 | 0.9 | 1.95 | A |

TABLE 73. IC Evaluation Results for Th-232 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 227 | 35.6 | 1.0 | 1.35 | A |
| 231 | 32.09 | 0.35 | 0.40 | A |
| 241 | 65.3 | 5.2 | 16.20 | N |
| 246 | 40.4 | 2.5 | 3.75 | N |
| 281 | 34.06 | 4.087 | 0.58 | A |
| 293 | 33.2 | 1.5 | 0.15 | A |
| 296 | 36.98 | 2.27 | 2.04 | W |
| 308 | 9.12 | 0.11 | 11.89 | N |
| 311 | 34.1 | 2.4 | 0.60 | A |
| 318 | 70.2 | 5.1 | 18.65 | N |
| 330 | 37 | 3.7 | 2.05 | W |
| 334 | 41.1 | 2.0 | 4.10 | N |
| 336 | 4.62 | 0.27 | 14.14 | N |

Activity [ Bq/kg ]

25

30

Reported Results for U-234 Determination in Sample 4

TABLE 74. Target values

Parameter	Value

35

Target value [Bq/kg]	23.0 Uncertainty [Bq/kg]	1.5

10

15

20

135

89

307

215

Labcode

FIGURE 66. Reported results and evaluation parameters for U-234 in Sample 4.

TABLE 75. IC Evaluation Results for U-234 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 89 | 21.63 | 1.84 | 0.91 | A |
| 135 | 20 | 3 | 2.00 | A |
| 215 | 170 | 6 | 98.00 | N |
| 307 | 24.2 | 2.9 | 0.80 | A |

Activity [ Bq/kg ]

25

30

Reported Results for U-238 Determination in Sample 4

TABLE 76. Target values

Parameter	Value

35

Target value [Bq/kg]	23.0 Uncertainty [Bq/kg]	1.5

10

15

20

336

125

202

97

135

174

240

11

67

247

134

89

179

107

122

232

130

73

201

10

12

307

156

123

231

128

302

38

71

293

47

157

162

1

246

176

241

273

46

272

281

215

Labcode

FIGURE 67. Reported results and evaluation parameters for U-238 in Sample 4.

TABLE 77. IC Evaluation Results for U-238 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 1 | 29.9 | 7.2 | 4.60 | N |
| 10 | 24 | 2 | 0.67 | A |
| 11 | 21.2 | 2.8 | 1.20 | A |
| 12 | 24 | 2 | 0.67 | A |
| 38 | 25.4 | 5.5 | 1.60 | A |
| 46 | 42 | 4 | 12.67 | N |
| 47 | 28.1 | 5.0 | 3.40 | N |
| 67 | 21.25 | 2.78 | 1.17 | A |
| 71 | 26 | 5 | 2.00 | A |
| 73 | 23.3 | 4.0 | 0.20 | A |
| 89 | 21.63 | 1.84 | 0.91 | A |
| 97 | 17.98 | 0.63 | 3.35 | N |
| 107 | 21.9 | 3.4 | 0.73 | A |
| 122 | 21.9 | 3.0 | 0.73 | A |
| 123 | 24.8 | 3.1 | 1.20 | A |
| 125 | 9.10 | 2.12 | 9.27 | N |
| 128 | 24.94 | 3.33 | 1.29 | A |
| 130 | 23.0 | 2.0 | 0.00 | A |
| 134 | 21.5 | 2.9 | 1.00 | A |
| 135 | 18.6 | 1.7 | 2.93 | W |
| 156 | 24.5 | 3.0 | 1.00 | A |
| 157 | 28.3 | 6.7 | 3.53 | N |
| 162 | 29.8 | 5.2 | 4.53 | N |
| 174 | 19 | 3 | 2.67 | W |
| 176 | 36 | 5 | 8.67 | N |
| 179 | 21.75 | 0.82 | 0.83 | A |

TABLE 77. IC Evaluation Results for U-238 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 201 | 23.5 | 2.8 | 0.33 | A |
| 202 | 17.1 | 1.8 | 3.93 | N |
| 215 | 169 | 8 | 97.33 | N |
| 231 | 24.80 | 5.07 | 1.20 | A |
| 232 | 22 | 3 | 0.67 | A |
| 240 | 20.5 | 3.6 | 1.67 | A |
| 241 | 36.8 | 2.8 | 9.20 | N |
| 246 | 30.4 | 6.1 | 4.93 | N |
| 247 | 21.50 | 0.8 | 1.00 | A |
| 272 | 54.6 | 3.9 | 21.07 | N |
| 273 | 38 | 5 | 10.00 | N |
| 281 | 55.53 | 6.664 | 21.69 | N |
| 293 | 27.0 | 1.4 | 2.67 | W |
| 302 | 25 | 3 | 1.33 | A |
| 307 | 24.2 | 2.9 | 0.80 | A |
| 336 | 2.28 | 0.18 | 13.81 | N |

- References References

[1]  International Organization for Standardization (ISO). (2010). Conformity assessment - General requirements for proficiency

testing, ISO/IEC 17043:2010. Geneva: Switzerland.

[2] International Organization for Standardization (ISO). (2022). Statistical methods for use in proficiency testing by interlabo- ratory comparison, ISO 13528:2022. Geneva: Switzerland.

# Contributors to evaluation and report

S. Patterson, B. Seslak, A. Trinkl: Terrestrial Environmental Radiochemistry Laboratory, Division of Physical and Chemical Sci- ences, Department of Nuclear Sciences and Applications

# Disclaimer

This report has been generated automatically and is for your personal information only. If you find, that any information pro- vided on this document might be incorrect please contact us as soon as possible.