IAEA-TERC-2025-01 Worldwide Proficiency Test Exercise, Pie Charts, S-Shapes and Reported Results with Scores

Terrestrial Environmental Radiochemistry Laboratory, International Atomic Energy Agency

January 7, 2026

Contents

- Description of Samples	1

  - Sample 1, Drinking Water (spiked)	1

  - Sample 2, Natural mineral water (non-spiked)	2

  - Sample 3, Soil	2

  - Samples 4, Vegetation	2

  - Samples 5, Simulated contaminated surface sample (white envelope)	2

  - Reference date for decay correction	3

  - Codes for analytical techniques in evaluation tables	3

- Evaluation of Results	3

  - Evaluation of parameters with independently assigned target values	3

  - Evaluation of parameters without independently assigned target value (“intercomparison” approach)	4

- Data Evaluation	5

  - Evaluation of activity concentrations of radionuclides and mass fractions by ICP-MS	5

  - Display of intercomparison values	158

- Contributors to Evaluation and Report	176

- Disclaimer	176

Abstract

The IAEA Terrestrial Environmental Radiochemistry (TERC) Laboratory provided the Worldwide Proficiency Test Exercise on the determination of anthropogenic and natural radionuclides in water, soil, vegetation and simulated contaminated surface samples.

# 1 Description of Samples

- Sample 1: Water spiked with naturally occurring and anthropogenic radionuclides

- Sample 2: Water sample containing environmentally occurring radionuclides

- Sample 3: Soil containing environmentally occurring radionuclides

- Sample 4: Vegetation containing environmentally occurring radionuclides

- Sample 5: Printed contaminated surface containing a mixture of Ba-133 and Sr-90

## Sample 1, Drinking Water (spiked)

Drinking water was gravimetrically spiked with a known amount of a standard solution, prepared from individual solutions with a certified activity concentration of radionuclides. The batch of water was acidified to pH < 2 (0.05M HNO3) for stabilization, and thoroughly mixed prior to bottling. Each bottle contains approximately 500 g of water.

The target values for activity concentration and corresponding uncertainties were assigned by formulation, and control mea- surements were performed by gamma-ray spectrometry, alpha-particle spectrometry and liquid scintillation counting.

## Sample 2, Natural mineral water (non-spiked)

Natural mineral water was acidified to pH < 2 (0.05M HNO3) for stabilization, and thoroughly mixed prior to bottling. Each bottle contains approximately 500 g of water.

The water contains naturally occurring radionuclides at low levels. Specifically, uranium is present at a mass fraction between

  - ng/g and 15 ng/g.

The water contains elevated mass concentrations of certain inorganic ions compared to typical drinking water, including:

    - Calcium (Ca): between 100 and 500 mg/L

    - Magnesium (Mg): between 20 and 200 mg/L

    - Sulphate (SO4): between 100 and 500 mg/L

    - Sodium (Na): between 10 and 100 mg/L

The determination of target values and associated uncertainties of the radionuclides has been carried-out using alpha-particle spectrometry and ICP-MS.

## Sample 3, Soil

Approximately 200 g of soil is provided for gamma-ray spectrometry measurement. The bottled samples were sterilized (25 kGy gamma dose) and the related irradiation certificate is available on the exercises website. It contains environmentally occurring radionuclides.

The determination of activity concentration target values and associated uncertainties of the radionuclides has been carried-out using high precision semiconductor gamma-ray spectrometry measurements in compliance with ISO/IEC 17025. The sample was characterized for gamma-emitting radionuclides only.

## Samples 4, Vegetation

Approximately 100 g of vegetation sample is provided for gamma-ray spectrometry measurement. The bottled samples were sterilized (25 kGy gamma dose) and the related certificate is available on the PT website. It contains environmentally occurring radionuclides.

The determination of target values and associated uncertainties of the radionuclides has been carried-out using high precision semiconductor gamma-ray spectrometry measurements in compliance with ISO/IEC 17025. The samples are characterized for gamma-emitting radionuclides only.

## Samples 5, Simulated contaminated surface sample (white envelope)

A set of printed canvas circles for contamination surface monitoring have been included in this proficiency test (Figure 1). They were prepared using an in-house printing technique. One blank sample of the same size, outlined with black ink (non-spiked) is also added for correction (Blank).

The simulated contaminated surface sample (purple) is printed with a combination of magenta ink (Sr-90) and blue ink (Ba- 133), and the activity of both radionuclides should be reported (Sample 05).

The simulated contaminated surface sample (blue) is a QC sample (QC-1) and was spiked with a known activity of Ba-133 (see Table 1).The simulated contaminated surface sample (magenta) is a QC sample (QC-2) and was spiked with a known activity of Sr-90 (see Table 1).

The provided QC samples are intended to support the validation and/or calibration of measurement instruments to facilitate the determination of an unknown Sample 05.

Figure 1. Simulated contaminated surface samples

Table 1. Radionuclide content in simulated contaminated surface QC samples.

| Sample | Sr-90, Bq/sample | Ba-133, Bq/sample |
| --- | --- | --- |
| QC – 1 QC – 2 | - 4.39 ± 0.08* | 7.36 ± 0.25* - |

* Uncertainty is expressed as a combined standard uncertainty with coverage factor k=1

## Reference date for decay correction

The reference date for decay correction in all samples is 2025-01-01.

## Codes for analytical techniques in evaluation tables

Due to the limited space available in the evaluation tables, the analytical techniques are abbreviated in the tables as shown in the following list:

    - alpha: Alpha Spectrometry

    - beta: Beta Counting

    - gamma: Gamma-ray Spectrometry

    - ICP-MS: Inductively Coupled Plasma Mass Spectrometry.

# Evaluation of Results

The reported results were evaluated according to current IAEA procedures for proficiency testing. Reported results were com- pared to a target value. Criteria to evaluate each participant’s performance take into consideration the relative bias of reported values (trueness) as well as the reported uncertainty of measurement (precision).

Target values were assigned independently for most parameters, either by formulation or characterization measurements. In case of results for gross-alpha and gross-beta activity concentrations, a consensus value from reported results was used.

## Evaluation of parameters with independently assigned target values

Each reported result was evaluated according to the following steps:

The relative bias between the reported and the target value (the best estimate of the true value) is expressed by the following equation:

Bi asrelative =

V a l u e r e p o r t e d − V a l u e t a r g e t V aluetar get

∗ 100%

The relative bias was compared to the Maximum Acceptable Relative Bias (MARB) which is determined for each analyte, con- sidering the radioanalytical methods, the level of radioactivity and the complexity of the analysis.

If |Bi asrelative | ≤ M ARB , the result will be rated "Accepted (A)" for trueness.

Based on fit-for-purpose and good laboratory practice principles, the relative expanded uncertainty should cover the relative bias:

The P statistic is calculated using the below equation:

「

， utar get ‘2

P =

， urepor ted ‘2

+

∗ 100%

I Atar get

Arepor ted

The relative bias is then compared to the P value:

|Bi asrelative | ≤ k ∗ P

where k is the coverage factor, k = 2.58 corresponding to a confidence level of 99%.

When the above criterion is fulfilled, the reported result is not significantly different from the target values considering the uncertainties associated with both values. The reported uncertainty of measurement is large enough to cover the bias of the result.

In addition, the P statistic is compared to the MARB.

P ≤ M ARB

When this criterion is fulfilled, the measurement uncertainty is not overestimated and fit-for-purpose in relation to the MARB criterion of this PT exercise.

When both criteria related to the measurement uncertainty are fulfilled, the reported result is rated "accepted (A)” for precision (measurement uncertainty). The result is rated "Not Accepted (N)" for precision if either of the two conditions are not fulfilled. The final score is assigned according to the detailed evaluation described above. The possible scores are listed below:

- "Accepted (A)" when both, trueness and precision were rated "Accepted"

- "Not Accepted (N)" when the trueness rating is "Not Accepted"

- "Warning (W)" when the trueness rating is “Accepted” but the precision rating is "Not Accepted"

A z score is provided in addition to above mentioned parameters due to requests by several participants.

V a l u e r e p o r t e d − V a l u e t a r g e t

z =

σPT

Where σPT is the standard deviation for proficiency assessment. The standard deviation for proficiency assessment is deter- mined for each analyte, considering the radioanalytical methods, the level of radioactivity and the complexity of the analysis. In case of independently assigned target values, it is connected to the MARB as follows:

σPT  =

M ARB.V aluetar get

3

## Evaluation of parameters without independently assigned target value (“intercomparison” approach)

When no independently assigned target value is available for a parameter, the consensus from all reported results from PT participants is applied as the target value. Robust statistical methods are applied to derive the target value and the robust standard deviation.

This approach was applied to gross alpha/beta measurement results.

The evaluation includes calculation of relative bias, z scores and ζ (zeta) scores according to the following steps: The relative bias is calculated as described in Section 2.1.

A z score is calculated using the following general equation:

V a l u e r e p o r t e d − V a l u e t a r g e t

z =

σPT

Where σPT is the standard deviation for proficiency assessment.

In this case, the standard deviation for proficiency assessment is the robust standard deviation of the population.

Here, the z score provides information on the deviation of the reported value from the consensus value derived from partici- pants’ results, expressed in multiples of the (robust) standard deviation of the population.

It should be emphasized that a z score derived this way is a parameter which depends on the population of reported results, i.e. on the performance of participants.

The following criteria apply for performance ratings based on obtained z scores:

    - |z| ≤ 2 ... accepted (A)

    - 2 < |z| < 3 ... questionable (Q)

    - |z| ≥ 3 ... not accepted (N)

A ζ (zeta) score is calculated using the following general equation:

V a l u e r e p o r t e d − V a l u e t a r g e t

I

ζ =

(urepor ted )2 + (utar get )2

The ζ score allows a combined assessment of the reported value and the reported uncertainty of measurement. The following criteria apply for performance ratings based on obtained ζ scores:

    - |ζ| ≤ 2 ... accepted (A)

    - 2 < |ζ| < 3 ... questionable (Q)

    - |ζ| ≥ 3 ... not accepted (N)

# Data Evaluation

A total of 399 laboratories successfully reported data in the IAEA-TERC-2025-01 proficiency test exercise.

For each of the analytes evaluated, the target value, uncertainty (k = 1) and the MARB are listed. A pie chart showing the dis- tribution of accepted, warning, not acceptable and not reported results, the graph showing all reported results, and a detailed table listing all evaluation results for the analyte (sorted by the laboratory code) are provided. The data points are coloured in blue, yellow and red, corresponding to the final score of "Accepted", "Warning", and "Not Accepted", respectively. The hori- zontal solid blue line indicates the target value. The dotted blue lines indicate the target value +/- its combined uncertainty. The horizontal solid red lines indicate the upper and lower limit of the accepted value range, defined by the target value +/- the maximum acceptable bias.

## Evaluation of activity concentrations of radionuclides and mass fractions by ICP-MS

H-3 determination of activity concentration in Sample 1

| Table 2. Target values |  |
| --- | --- |
| Parameter | Value |
| Activity concentration [ Bq/kg ] | 46.8 |
| Uncertainty [ Bq/kg ] | 2.4 |
| Maximum Acceptable Relative Bias [%] | 30 |

Accepted 39%	Not Accepted 5%	Warning 2%	Not Reported 54%

Figure 2. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

Activity concentration  [ Bq/kg ]

50

60

70

20

30

40

71

254

491

21

355

276

58

114

70

340

155

306

66

312

382

453

417

283

348

348

404

238

232

285

338

239

364

313

386

430

63

219

172

165

122

28

290

7

32

553

369

466

153

244

319

396

525

90

163

304

420

140

268

395

6

15

145

439

215

262

363

437

311

5

33

241

527

242

68

488

49

205

250

496

267

69

213

516

264

514

192

523

65

187

119

401

469

156

166

288

477

381

75

433

91

277

384

508

197

446

303

392

314

9

501

81

42

97

208

287

374

302

532

83

227

539

265

539

198

53

191

72

19

341

100

152

161

170

230

79

194

25

325

325

297

371

429

379

368

73

443

78

329

358

432

402

291

52

295

257

240

179

529

445

175

134

528

317

315

159

350

332

195

555

200

347

207

181

393

464

136

322

11

38

209

55

440

204

346

333

441

483

255

405

54

31

Laboratory code

Figure 3. Reported results and evaluation parameters for H-3. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 3. Evaluation results for H-3 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Reported value [Bq/kg ] 47.8 | Reported uncertainty (k = 1) [Bq/kg ] 1.1 | Relative bias [%] 2.1 | P-Test [%] 5.62 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 47.5 | 1.21 | 1.5 | 5.73 | A | A | A |
| 46.6 | 1.07 | -0.4 | 5.62 | A | A | A |
| 49 | 4.64 | 4.7 | 10.77 | A | A | A |
| 66.87 | 4.02 | 42.9 | 7.90 | N | N | N |
| 47.5 | 1.8 | 1.5 | 6.38 | A | A | A |
| 50.4 | 2.4 | 7.7 | 7.00 | A | A | A |
| 32 | 8 | -31.6 | 25.52 | N | A | N |
| 51.2 | 2.2 | 9.4 | 6.69 | A | A | A |
| 46.51 | 4.06 | -0.6 | 10.12 | A | A | A |
| 266 | 12.1 | 468.4 | 6.85 | N | N | N |
| 46.7 | 2.7 | -0.2 | 7.73 | A | A | A |
| 47.8 | 2.2 | 2.1 | 6.89 | A | A | A |
| 72.18 | 3.58 | 54.2 | 7.13 | N | N | N |
| 49.2 | 1.5 | 5.1 | 5.97 | A | A | A |
| 48 | 5 | 2.6 | 11.61 | A | A | A |
| 52.3 | 4.02 | 11.8 | 9.24 | A | A | A |
| 50 | 5 | 6.8 | 11.24 | A | A | A |
| 237.2 | 7.5 | 406.8 | 6.02 | N | N | N |
| 77.7 | 2.45 | 66.0 | 6.02 | N | N | N |
| 38 | 2 | -18.8 | 7.35 | A | A | A |
| 45.9 | 0.78 | -1.9 | 5.40 | A | A | A |
| 48.4 | 3.6 | 3.4 | 9.03 | A | A | A |
| 41.1 | 1.5185 | -12.2 | 6.32 | A | A | A |
| 47.9 | 3.43 | 2.4 | 8.81 | A | A | A |
| 48.1 | 2.8 | 2.8 | 7.76 | A | A | A |
| 39.9 | 1.3 | -14.7 | 6.08 | A | A | A |
| 14.38 | 0.72 | -69.3 | 7.17 | N | N | N |
| 50.2 | 1.94 | 7.3 | 6.42 | A | A | A |
| 51.7 | 4 | 10.5 | 9.28 | A | A | A |
| 48.67 | 1.36 | 4.0 | 5.84 | A | A | A |
| 51.8 | 2.65 | 10.7 | 7.24 | A | A | A |
| 51.1 | 1.2 | 9.2 | 5.64 | A | A | A |
| 49.14 | 3.33 | 5.0 | 8.50 | A | A | A |

Labcode

5

6

7

9

11

15

19

21

25

28

31

32

33

38

42

49

52

53

54

55

58

63

65

66

68

69

70

71

72

73

75

78

79

81

| Reported value [Bq/kg ] 49.6 | Reported uncertainty (k = 1) [Bq/kg ] 7.2 | Relative bias [%] 6.0 | P-Test [%] 15.40 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 47.1 | 3.3 | 0.6 | 8.68 | A | A | A |
| 48.7 | 3.4 | 4.1 | 8.66 | A | A | A |
| 49.3 | 4.2 | 5.3 | 9.94 | A | A | A |
| 50.7 | 0.9 | 8.3 | 5.43 | A | A | A |
| 39.5 | 2 | -15.6 | 7.21 | A | A | A |
| 48.5 | 5.3 | 3.6 | 12.07 | A | A | A |
| 46.5 | 6.5 | -0.6 | 14.89 | A | A | A |
| 55.1 | 3.9 | 17.7 | 8.74 | A | A | A |
| 63.06 | 7.66 | 34.7 | 13.19 | N | N | N |
| 47.3 | 3.9 | 1.1 | 9.71 | A | A | A |
| 47.5 | 9.5 | 1.5 | 20.65 | A | A | A |
| 50.7 | 2.9 | 8.3 | 7.68 | A | A | A |
| 47 | 4 | 0.4 | 9.94 | A | A | A |
| 40.4 | 4.6 | -13.7 | 12.49 | A | A | A |
| 48.51 | 1.72 | 3.6 | 6.23 | A | A | A |
| 57 | 7.125 | 21.8 | 13.51 | A | A | A |
| 50.8 | 7 | 8.6 | 14.70 | A | A | A |
| 47.1 | 4 | 0.6 | 9.92 | A | A | A |
| 46.3775 | 1.5977 | -0.9 | 6.18 | A | A | A |
| 48.6 | 3.6 | 3.9 | 9.01 | A | A | A |
| 50.8 | 1.8 | 8.6 | 6.23 | A | A | A |
| 46 | 3.2 | -1.7 | 8.64 | A | A | A |
| 54.2 | 1.2 | 15.8 | 5.59 | A | N | W |
| 53.4 | 1.7 | 14.1 | 6.04 | A | A | A |
| 60.52 | 5.11 | 29.3 | 9.88 | A | N | W |
| 48.4 | 3.6 | 3.4 | 9.03 | A | A | A |
| 50 | 2.8 | 6.8 | 7.59 | A | A | A |
| 48.3 | 2.58 | 3.2 | 7.40 | A | A | A |
| 51.1 | 1.25 | 9.2 | 5.68 | A | A | A |
| 58.7 | 2.9 | 25.4 | 7.12 | A | N | W |
| 48.72 | 4.38 | 4.1 | 10.35 | A | A | A |
| 49.8 | 4 | 6.4 | 9.53 | A | A | A |
| 58.9 | 4.1 | 25.9 | 8.65 | A | N | W |
| 83 | 2 | 77.3 | 5.67 | N | N | N |
| 48 | 2.52 | 2.6 | 7.34 | A | A | A |
| 60 | 10 | 28.2 | 17.44 | A | A | A |
| 49.3 | 1.79 | 5.3 | 6.28 | A | A | A |
| 75 | 8 | 60.3 | 11.84 | N | N | N |
| 48.1 | 1.52 | 2.8 | 6.02 | A | A | A |
| 47.6 | 6.7 | 1.7 | 14.98 | A | A | A |
| 45.93 | 2.97 | -1.9 | 8.25 | A | A | A |
| 49.7 | 0.8 | 6.2 | 5.37 | A | A | A |
| 50.8 | 2.6 | 8.6 | 7.25 | A | A | A |
| 44.7 | 4.8 | -4.5 | 11.90 | A | A | A |
| 44.5 | 2.07 | -4.9 | 6.92 | A | A | A |
| 45.3 | 5.2 | -3.2 | 12.57 | A | A | A |
| 53 | 3.6 | 13.2 | 8.51 | A | A | A |
| 47.8 | 1.4 | 2.1 | 5.91 | A | A | A |
| 47.88 | 4.8 | 2.3 | 11.26 | A | A | A |
| 47 | 3.5 | 0.4 | 9.04 | A | A | A |
| 48 | 1.6 | 2.6 | 6.12 | A | A | A |
| 30.15 | 1 | -35.6 | 6.11 | N | N | N |
| 103.64 | 5.3 | 121.4 | 7.24 | N | N | N |
| 52.48 | 2.11 | 12.1 | 6.52 | A | A | A |
| 47.62 | 1.69 | 1.8 | 6.24 | A | A | A |
| 48.2 | 2.1 | 3.0 | 6.73 | A | A | A |
| 49.8 | 8 | 6.4 | 16.86 | A | A | A |
| 48.03 | 3.84 | 2.6 | 9.50 | A | A | A |
| 47.4 | 2.4 | 1.3 | 7.21 | A | A | A |

Labcode

83

90

91

97

100

114

119

122

134

136

140

145

152

153

155

156

159

161

163

165

166

170

172

175

179

181

187

191

192

194

195

197

198

200

204

205

207

208

209

213

215

219

227

230

232

238

239

240

241

242

244

250

254

255

257

262

264

265

267

268

| Reported value [Bq/kg ] 37.1 | Reported uncertainty (k = 1) [Bq/kg ] 1.75 | Relative bias [%] -20.7 | P-Test [%] 6.97 | Trueness evaluation A | Precision evaluation N | Final Score W |
| --- | --- | --- | --- | --- | --- | --- |
| 48.7 | 6.4 | 4.1 | 14.11 | A | A | A |
| 43.4 | 3.54 | -7.3 | 9.63 | A | A | A |
| 44.7 | 2.5 | -4.5 | 7.59 | A | A | A |
| 49.4 | 2 | 5.6 | 6.53 | A | A | A |
| 48.6 | 5.74 | 3.9 | 12.88 | A | A | A |
| 46.57 | 3.4 | -0.5 | 8.92 | A | A | A |
| 52.2 | 9.396 | 11.5 | 18.72 | A | A | A |
| 52.3 | 4.54 | 11.8 | 10.08 | A | A | A |
| 51.3 | 4.2 | 9.6 | 9.66 | A | A | A |
| 49.5 | 3.9 | 5.8 | 9.40 | A | A | A |
| 48.9 | 2.4 | 4.5 | 7.10 | A | A | A |
| 47.1 | 1.4 | 0.6 | 5.93 | A | A | A |
| 40.4 | 6.1 | -13.7 | 15.95 | A | A | A |
| 47.75 | 1.97 | 2.0 | 6.58 | A | A | A |
| 42 | 3 | -10.3 | 8.79 | A | A | A |
| 45.7 | 1.97 | -2.4 | 6.70 | A | A | A |
| 48.97 | 4.42 | 4.6 | 10.38 | A | A | A |
| 56.982 | 5.458 | 21.8 | 10.86 | A | A | A |
| 56.1 | 5.1 | 19.9 | 10.44 | A | A | A |
| 47 | 3.29 | 0.4 | 8.68 | A | A | A |
| 64.7 | 3.8 | 38.2 | 7.80 | N | N | N |
| 51.2 | 4.6 | 9.4 | 10.34 | A | A | A |
| 51.2 | 4.6 | 9.4 | 10.34 | A | A | A |
| 51.9 | 9.2 | 10.9 | 18.45 | A | A | A |
| 58 | 2.9 | 23.9 | 7.16 | A | N | W |
| 89.63 | 4.3 | 91.5 | 7.02 | N | N | N |
| 45.13 | 4.96 | -3.6 | 12.13 | A | A | A |
| 40.28 | 4.93 | -13.9 | 13.27 | A | A | A |
| 50.4 | 2 | 7.7 | 6.48 | A | A | A |
| 88 | 6.6 | 88.0 | 9.09 | N | N | N |
| 59.5 | 7.9 | 27.1 | 14.23 | A | A | A |
| 44 | 3 | -6.0 | 8.53 | A | A | A |
| 44 | 3 | -6.0 | 8.53 | A | A | A |
| 57.6 | 0.8 | 23.1 | 5.31 | A | N | W |
| 35.6 | 1.85 | -23.9 | 7.30 | A | N | W |
| 52 | 0.05 | 11.1 | 5.13 | A | A | A |
| 47.7 | 2.39 | 1.9 | 7.17 | A | A | A |
| 45.6 | 1.3 | -2.6 | 5.87 | A | A | A |
| 51.6 | 5.5 | 10.3 | 11.83 | A | A | A |
| 46.9 | 4.17 | 0.2 | 10.26 | A | A | A |
| 51.3 | 3.5 | 9.6 | 8.54 | A | A | A |
| 49.4 | 2.7 | 5.6 | 7.49 | A | A | A |
| 51.5 | 7.7 | 10.0 | 15.81 | A | A | A |
| 48.62 | 1.63 | 3.9 | 6.13 | A | A | A |
| 42 | 2.4 | -10.3 | 7.68 | A | A | A |
| 48.7 | 7 | 4.1 | 15.26 | A | A | A |
| 45.72 | 4.97 | -2.3 | 12.02 | A | A | A |
| 48.9 | 2.6 | 4.5 | 7.39 | A | A | A |
| 60.57 | 5.15 | 29.4 | 9.93 | A | N | W |
| 47.48 | 3.79 | 1.4 | 9.49 | A | A | A |
| 47 | 1.55 | 0.4 | 6.10 | A | A | A |
| 48.5 | 12.1 | 3.6 | 25.47 | A | A | A |
| 52.1 | 1.7 | 11.3 | 6.08 | A | A | A |
| 44.4 | 1.4 | -5.1 | 6.02 | A | A | A |
| 128.4 | 4.3 | 174.4 | 6.12 | N | N | N |
| 43.3 | 1.9 | -7.5 | 6.75 | A | A | A |
| 47.1 | 3.3 | 0.6 | 8.68 | A | A | A |
| 51.3 | 2.89 | 9.6 | 7.62 | A | A | A |
| 45.8 | 4.2 | -2.1 | 10.51 | A | A | A |

Labcode

276

277

283

285

287

288

290

291

295

297

302

303

304

306

311

312

313

314

315

317

319

322

325

325

329

332

333

338

340

341

346

347

348

348

350

355

358

363

364

368

369

371

374

379

381

382

384

386

392

393

395

396

401

402

404

405

417

420

429

430

| Reported value [Bq/kg ] 52 | Reported uncertainty (k = 1) [Bq/kg ] 4.43 | Relative bias [%] 11.1 | P-Test [%] 9.94 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 48.69 | 3.44 | 4.0 | 8.73 | A | A | A |
| 47.7 | 2.4 | 1.9 | 7.18 | A | A | A |
| 47.5 | 1.9 | 1.5 | 6.50 | A | A | A |
| 82.628 | 1.947 | 76.6 | 5.64 | N | N | N |
| 92.4 | 5 | 97.4 | 7.46 | N | N | N |
| 51.7 | 2.5 | 10.5 | 7.05 | A | A | A |
| 53.9 | 2.3 | 15.2 | 6.67 | A | A | A |
| 48.76 | 2.33 | 4.2 | 7.01 | A | A | A |
| 42.69 | 6.18 | -8.8 | 15.36 | A | A | A |
| 61.85 | 9.28 | 32.2 | 15.86 | N | A | N |
| 46.9 | 0.958 | 0.2 | 5.52 | A | A | A |
| 48.5 | 1.8 | 3.6 | 6.33 | A | A | A |
| 48.6 | 1.2 | 3.9 | 5.69 | A | A | A |
| 92.4 | 5 | 97.4 | 7.46 | N | N | N |
| 47.9 | 1.2 | 2.4 | 5.71 | A | A | A |
| 31.502 | 2.909 | -32.7 | 10.56 | N | N | N |
| 48 | 0.81 | 2.6 | 5.40 | A | A | A |
| 49.1 | 1.6 | 4.9 | 6.08 | A | A | A |
| 48.7 | 3.2 | 4.1 | 8.34 | A | A | A |
| 48.2 | 2.7 | 3.0 | 7.59 | A | A | A |
| 48.1 | 1.5 | 2.8 | 6.00 | A | A | A |
| 48.3 | 4.8 | 3.2 | 11.18 | A | A | A |
| 47 | 2.8 | 0.4 | 7.86 | A | A | A |
| 47.8 | 1.5 | 2.1 | 6.01 | A | A | A |
| 55.5 | 4.4 | 18.6 | 9.44 | A | A | A |
| 53.5 | 4.6 | 14.3 | 10.01 | A | A | A |
| 49.5 | 2 | 5.8 | 6.53 | A | A | A |
| 49.7 | 1.62 | 6.2 | 6.08 | A | A | A |
| 49.8 | 3.51 | 6.4 | 8.72 | A | A | A |
| 46.74 | 3.03 | -0.1 | 8.27 | A | A | A |
| 58.7 | 8.2 | 25.4 | 14.88 | A | A | A |

Labcode

432

433

437

439

440

441

443

445

446

453

464

466

469

477

483

488

491

496

501

508

514

516

523

525

527

528

529

532

539

539

553

555

Na-22 determination of activity concentration in Sample 1

Table 4. Target values

| Parameter | Value |
| --- | --- |
| Activity concentration [ Bq/kg ] | 19.9 |
| Uncertainty [ Bq/kg ] | 1.2 |
| Maximum Acceptable Relative Bias [%] | 20 |

Accepted 70%	Not Accepted 6%	Warning 1%	Not Reported 23%

25

Figure 4. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

15

411

529

252

495

81

71

180

34

491

300

434

27

435

143

128

392

228

281

371

331

272

184

24

134

85

249

306

155

185

207

388

202

402

477

483

68

265

257

535

150

2

364

31

268

522

284

330

395

166

398

133

374

439

412

55

9

145

329

471

53

357

262

213

341

109

218

94

358

295

384

355

433

288

223

293

401

469

19

45

283

528

452

107

553

151

72

78

303

404

437

291

307

348

267

156

15

168

304

417

484

63

121

253

316

450

403

114

170

242

49

57

104

214

271

299

379

40

459

75

260

270

478

17

90

235

319

344

111

428

132

445

516

508

238

514

280

206

382

544

479

127

130

466

217

315

463

13

381

555

152

254

38

323

423

97

Laboratory code

Figure 5. Reported results and evaluation parameters for Na-22. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 5. Evaluation results for Na-22 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

Labcode

Reported value

[Bq/kg ]

Reported uncertainty

(k = 1) [Bq/kg ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

Activity concentration  [ Bq/kg ]

20

2	17.9	1.6	-10.1	10.78	A	A	A

8	21.77	1.31	9.4	8.52	A	A	A

9	18.5	2.2	-7.0	13.33	A	A	A

| Reported value [Bq/kg ] 19.7 | Reported uncertainty (k = 1) [Bq/kg ] 0.93 | Relative bias [%] -1.0 | P-Test [%] 7.66 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 21.5 | 1.1 | 8.0 | 7.91 | A | A | A |
| 19.7 | 0.9 | -1.0 | 7.57 | A | A | A |
| 20.11 | 0.82 | 1.1 | 7.28 | A | A | A |
| 19.25 | 1.2 | -3.3 | 8.67 | A | A | A |
| 17.2 | 0.72 | -13.6 | 7.34 | A | A | A |
| 20 | 0.8 | 0.5 | 7.24 | A | A | A |
| 15.85 | 0.5 | -20.4 | 6.81 | N | N | N |
| 19.4 | 1.683 | -2.5 | 10.57 | A | A | A |
| 18 | 0.61 | -9.6 | 6.92 | A | A | A |
| 17.54 | 0.42 | -11.9 | 6.49 | A | A | A |
| 15.1 | 1.1 | -24.1 | 9.46 | N | A | N |
| 22.21 | 1.9 | 11.6 | 10.47 | A | A | A |
| 20.02 | 1.33 | 0.6 | 8.97 | A | A | A |
| 19.5 | 0.925 | -2.0 | 7.67 | A | A | A |
| 19.9 | 1.2 | 0.0 | 8.53 | A | A | A |
| 19.27 | 0.77 | -3.2 | 7.23 | A | A | A |
| 20 | 2 | 0.5 | 11.68 | A | A | A |
| 20 | 0.9 | 0.5 | 7.52 | A | A | A |
| 20.73 | 0.88 | 4.2 | 7.37 | A | A | A |
| 21.1 | 0.92 | 6.0 | 7.44 | A | A | A |
| 18.53 | 1.11 | -6.9 | 8.50 | A | A | A |
| 18.43 | 0.3 | -7.4 | 6.25 | A | A | A |
| 20 | 2 | 0.5 | 11.68 | A | A | A |
| 18 | 1 | -9.6 | 8.20 | A | A | A |
| 17.2 | 1.5 | -13.6 | 10.60 | A | A | A |
| 19.7 | 1.5 | -1.0 | 9.71 | A | A | A |
| 19.79 | 2 | -0.6 | 11.77 | A | A | A |
| 20 | 1.8 | 0.5 | 10.83 | A | A | A |
| 19.26 | 1.0175 | -3.2 | 8.02 | A | A | A |
| 17.6 | 0.67 | -11.6 | 7.13 | A | A | A |
| 19.20319 | 0.8 | -3.5 | 7.33 | A | A | A |
| 18.5 | 0.9 | -7.0 | 7.75 | A | A | A |
| 14.38 | 0.72 | -27.7 | 7.84 | N | N | N |
| 19.5 | 1.51 | -2.0 | 9.81 | A | A | A |
| 19.5 | 1 | -2.0 | 7.92 | A | A | A |
| 19.3 | 1.94 | -3.0 | 11.72 | A | A | A |
| 20.1 | 4.17 | 1.0 | 21.60 | A | N | W |
| 19.5 | 0.73 | -2.0 | 7.10 | A | A | A |
| 14.4 | 0.3 | -27.6 | 6.38 | N | N | N |
| 14.2 | 0.6 | -28.6 | 7.36 | N | N | N |
| 17.38 | 0.35 | -12.7 | 6.36 | A | A | A |
| 18.2 | 2.6 | -8.5 | 15.51 | A | A | A |
| 17.243 | 2.267 | -13.3 | 14.46 | A | A | A |
| 20.2 | 1.1 | 1.5 | 8.13 | A | A | A |
| 18.89 | 0.87 | -5.1 | 7.59 | A | A | A |
| 95.5 | 5.6 | 379.9 | 8.41 | N | N | N |
| 19.11 | 1.06 | -4.0 | 8.19 | A | A | A |
| 11.33 | 0.32 | -43.1 | 6.66 | N | N | N |
| 20 | 0.9 | 0.5 | 7.52 | A | A | A |
| 19.8 | 0.9 | -0.5 | 7.55 | A | A | A |
| 19.4 | 1.1 | -2.5 | 8.28 | A | A | A |
| 18.8 | 0.2 | -5.5 | 6.12 | A | A | A |
| 18.774407 | 1.05853 | -5.7 | 8.26 | A | A | A |
| 20.26 | 1.74 | 1.8 | 10.49 | A | A | A |
| 19.07 | 2.86 | -4.2 | 16.16 | A | A | A |
| 19.9 | 1.3 | 0.0 | 8.89 | A | A | A |
| 18.68 | 0.92 | -6.1 | 7.79 | A | A | A |
| 18.89 | 1.9 | -5.1 | 11.73 | A | A | A |
| 19.8 | 1.4 | -0.5 | 9.29 | A | A | A |

Labcode

11

13

15

17

19

24

25

27

28

31

32

34

38

40

41

42

45

49

50

51

52

53

55

57

58

60

62

63

65

66

68

69

70

71

72

73

74

75

78

79

81

82

83

85

90

94

97

100

102

104

105

107

109

110

111

113

114

116

119

121

| Reported value [Bq/kg ] 19.9 | Reported uncertainty (k = 1) [Bq/kg ] 1.1 | Relative bias [%] 0.0 | P-Test [%] 8.18 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 20.811281 | 1.875054 | 4.6 | 10.84 | A | A | A |
| 16.2 | 0.82 | -18.6 | 7.87 | A | A | A |
| 20.91 | 0.83 | 5.1 | 7.22 | A | A | A |
| 20.4 | 0.9 | 2.5 | 7.47 | A | A | A |
| 18.3 | 0.5 | -8.0 | 6.62 | A | A | A |
| 17.2 | 0.2 | -13.6 | 6.14 | A | A | A |
| 17.6 | 0.712 | -11.6 | 7.26 | A | A | A |
| 20.1 | 1.3 | 1.0 | 8.84 | A | A | A |
| 16.01079 | 1.205633 | -19.5 | 9.65 | A | A | A |
| 18.5 | 2.96 | -7.0 | 17.10 | A | A | A |
| 14.2341 | 0.7278 | -28.5 | 7.91 | N | N | N |
| 17.8 | 0.9 | -10.6 | 7.87 | A | A | A |
| 19.44 | 1.94 | -2.3 | 11.66 | A | A | A |
| 21.9 | 1.3 | 10.1 | 8.46 | A | A | A |
| 20.6 | 1.2 | 3.5 | 8.38 | A | A | A |
| 17.33 | 0.96 | -12.9 | 8.19 | A | A | A |
| 19.68 | 0.61 | -1.1 | 6.78 | A | A | A |
| 16.8 | 2.2 | -15.6 | 14.42 | A | A | A |
| 20.3 | 1 | 2.0 | 7.79 | A | A | A |
| 19.6 | 0.85 | -1.5 | 7.43 | A | A | A |
| 19.57 | 0.8 | -1.7 | 7.29 | A | A | A |
| 18.2 | 0.61 | -8.5 | 6.90 | A | A | A |
| 19.7 | 1.6 | -1.0 | 10.12 | A | A | A |
| 19.9 | 1.8 | 0.0 | 10.87 | A | A | A |
| 20.11 | 0.91 | 1.1 | 7.54 | A | A | A |
| 20.9 | 0.5 | 5.0 | 6.49 | A | A | A |
| 20.2 | 1.3 | 1.5 | 8.82 | A | A | A |
| 15.12 | 0.67 | -24.0 | 7.48 | N | N | N |
| 14.5 | 0.7 | -27.1 | 7.72 | N | N | N |
| 16.51 | 1.65 | -17.0 | 11.67 | A | A | A |
| 17.23 | 0.85 | -13.4 | 7.79 | A | A | A |
| 17.12 | 2.22 | -14.0 | 14.30 | A | A | A |
| 17.39 | 0.53 | -12.6 | 6.76 | A | A | A |
| 20 | 1.8 | 0.5 | 10.83 | A | A | A |
| 17.25 | 3.45 | -13.3 | 20.89 | A | N | W |
| 20.5 | 1.8 | 3.0 | 10.65 | A | A | A |
| 14.8 | 0.72 | -25.6 | 7.75 | N | N | N |
| 18.23 | 1.46 | -8.4 | 10.03 | A | A | A |
| 20.5 | 1.7 | 3.0 | 10.25 | A | A | A |
| 19.1 | 0.6 | -4.0 | 6.80 | A | A | A |
| 17.48 | 0.26 | -12.2 | 6.21 | A | A | A |
| 17.1 | 1.6 | -14.1 | 11.13 | A | A | A |
| 19 | 0.5 | -4.5 | 6.58 | A | A | A |
| 20.6 | 2.3 | 3.5 | 12.69 | A | A | A |
| 17.4 | 1.1 | -12.6 | 8.74 | A | A | A |
| 20.4 | 1.01 | 2.5 | 7.80 | A | A | A |
| 17 | 1.2 | -14.6 | 9.28 | A | A | A |
| 15.74 | 0.72 | -20.9 | 7.57 | N | N | N |
| 18.7 | 1.44 | -6.0 | 9.78 | A | A | A |
| 20 | 0.2 | 0.5 | 6.11 | A | A | A |
| 21.1 | 2.5 | 6.0 | 13.29 | A | A | A |
| 18.8183 | 0.67876 | -5.4 | 7.03 | A | A | A |
| 19.2 | 1.8 | -3.5 | 11.15 | A | A | A |
| 18.5 | 1.7 | -7.0 | 10.99 | A | A | A |
| 17.95 | 0.77 | -9.8 | 7.40 | A | A | A |
| 19.4 | 0.6 | -2.5 | 6.78 | A | A | A |
| 16.7 | 0.8 | -16.1 | 7.70 | A | A | A |
| 18.6 | 0.5 | -6.5 | 6.60 | A | A | A |
| 19.2 | 0.8 | -3.5 | 7.33 | A | A | A |

Labcode

124

127

128

130

132

133

134

139

140

143

145

147

150

151

152

154

155

156

158

159

161

165

166

168

170

175

177

178

179

180

181

183

184

185

187

189

191

194

197

198

200

202

203

204

206

207

208

209

211

213

214

217

218

223

224

226

227

228

230

232

| Reported value [Bq/kg ] 20.8 | Reported uncertainty (k = 1) [Bq/kg ] 1.6 | Relative bias [%] 4.5 | P-Test [%] 9.77 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 20.2 | 2 | 1.5 | 11.59 | A | A | A |
| 20.5 | 0.941 | 3.0 | 7.58 | A | A | A |
| 17.9 | 0.7 | -10.1 | 7.19 | A | A | A |
| 20 | 0.9 | 0.5 | 7.52 | A | A | A |
| 19.92 | 0.52 | 0.1 | 6.57 | A | A | A |
| 16.151 | 1.108 | -18.8 | 9.13 | A | A | A |
| 17.260534 | 0.757484 | -13.3 | 7.46 | A | A | A |
| 19.8 | 1.4 | -0.5 | 9.29 | A | A | A |
| 12.6 | 0.74 | -36.7 | 8.42 | N | N | N |
| 19.8 | 1.4 | -0.5 | 9.29 | A | A | A |
| 22.16 | 4.21 | 11.4 | 19.93 | A | A | A |
| 19.67 | 0.95 | -1.2 | 7.73 | A | A | A |
| 17.62 | 0.6 | -11.5 | 6.93 | A | A | A |
| 19.8 | 0.9 | -0.5 | 7.55 | A | A | A |
| 20.1 | 1 | 1.0 | 7.82 | A | A | A |
| 17.5 | 0.6 | -12.1 | 6.94 | A | A | A |
| 18.63 | 0.49 | -6.4 | 6.58 | A | A | A |
| 20.1 | 1.6 | 1.0 | 9.99 | A | A | A |
| 17.6 | 2.1 | -11.6 | 13.37 | A | A | A |
| 19.61 | 1.9 | -1.5 | 11.41 | A | A | A |
| 18 | 0.9 | -9.6 | 7.83 | A | A | A |
| 20.1 | 0.8 | 1.0 | 7.23 | A | A | A |
| 20 | 1.2 | 0.5 | 8.51 | A | A | A |
| 17.1 | 1.1 | -14.1 | 8.82 | A | A | A |
| 17.6 | 0.903554 | -11.6 | 7.92 | A | A | A |
| 19.7 | 1.1 | -1.0 | 8.22 | A | A | A |
| 17.1 | 1.1 | -14.1 | 8.82 | A | A | A |
| 20.1 | 2.1 | 1.0 | 12.06 | A | A | A |
| 20.2 | 0.7 | 1.5 | 6.95 | A | A | A |
| 20.56 | 1.1 | 3.3 | 8.06 | A | A | A |
| 16.85 | 0.44 | -15.3 | 6.57 | A | A | A |
| 19.3 | 0.54 | -3.0 | 6.65 | A | A | A |
| 18.09 | 0.38 | -9.1 | 6.39 | A | A | A |
| 19.5 | 1.1 | -2.0 | 8.26 | A | A | A |
| 17.7 | 0.7 | -11.1 | 7.21 | A | A | A |
| 17.3 | 0.4 | -13.1 | 6.46 | A | A | A |
| 19.1 | 0.55 | -4.0 | 6.68 | A | A | A |
| 20 | 0.8 | 0.5 | 7.24 | A | A | A |
| 19.59 | 2.057 | -1.6 | 12.11 | A | A | A |
| 18.1 | 0.6 | -9.1 | 6.88 | A | A | A |
| 19.2 | 1.9 | -3.5 | 11.59 | A | A | A |
| 18.7 | 2 | -6.0 | 12.28 | A | A | A |
| 18.94 | 0.73 | -4.8 | 7.16 | A | A | A |
| 18.33 | 1 | -7.9 | 8.13 | A | A | A |
| 20 | 2 | 0.5 | 11.68 | A | A | A |
| 15.72 | 0.44 | -21.0 | 6.65 | N | N | N |
| 19.5 | 2 | -2.0 | 11.90 | A | A | A |
| 19.7 | 1.8 | -1.0 | 10.95 | A | A | A |
| 17.3 | 2.6 | -13.1 | 16.19 | A | A | A |
| 19.6 | 0.7 | -1.5 | 7.01 | A | A | A |
| 22.3 | 1.6 | 12.1 | 9.37 | A | A | A |
| 17.8 | 0.815 | -10.6 | 7.57 | A | A | A |
| 20 | 1 | 0.5 | 7.83 | A | A | A |
| 17.4 | 1.05 | -12.6 | 8.53 | A | A | A |
| 18.51 | 0.56 | -7.0 | 6.75 | A | A | A |
| 21.233 | 1.035 | 6.7 | 7.75 | A | A | A |
| 19.8 | 2.3 | -0.5 | 13.09 | A | A | A |
| 20.03 | 0.36 | 0.7 | 6.29 | A | A | A |
| 20.2 | 0.883 | 1.5 | 7.45 | A | A | A |

Labcode

233

235

238

239

241

242

247

249

250

252

253

254

256

257

258

260

261

262

264

265

267

268

270

271

272

274

276

277

278

279

280

281

283

284

285

286

287

288

290

291

292

293

294

295

297

299

300

303

304

306

307

310

311

312

313

314

315

316

318

319

| Reported value [Bq/kg ] 19.6 | Reported uncertainty (k = 1) [Bq/kg ] 0.5 | Relative bias [%] -1.5 | P-Test [%] 6.55 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 22.3 | 2.1 | 12.1 | 11.18 | A | A | A |
| 19.7 | 1.04 | -1.0 | 8.01 | A | A | A |
| 20.2 | 1.7 | 1.5 | 10.35 | A | A | A |
| 18.8 | 1.6 | -5.5 | 10.43 | A | A | A |
| 18.5 | 0.7 | -7.0 | 7.12 | A | A | A |
| 18.13 | 0.92 | -8.9 | 7.88 | A | A | A |
| 17 | 0.6 | -14.6 | 6.99 | A | A | A |
| 20.5 | 1.8 | 3.0 | 10.65 | A | A | A |
| 20.93 | 1.44 | 5.2 | 9.15 | A | A | A |
| 21.5 | 0.8 | 8.0 | 7.09 | A | A | A |
| 18.7 | 0.7 | -6.0 | 7.10 | A | A | A |
| 20.2 | 1.7 | 1.5 | 10.35 | A | A | A |
| 18.13 | 0.78 | -8.9 | 7.41 | A | A | A |
| 19.5 | 1.7 | -2.0 | 10.60 | A | A | A |
| 19.6 | 1 | -1.5 | 7.90 | A | A | A |
| 18.3 | 0.4 | -8.0 | 6.41 | A | A | A |
| 19.2 | 1 | -3.5 | 7.97 | A | A | A |
| 20.7 | 0.9 | 4.0 | 7.43 | A | A | A |
| 19.02 | 0.32 | -4.4 | 6.26 | A | A | A |
| 20.6 | 0.6 | 3.5 | 6.70 | A | A | A |
| 18.6 | 0.9 | -6.5 | 7.73 | A | A | A |
| 18.9 | 0.9 | -5.0 | 7.68 | A | A | A |
| 19.8 | 0.7 | -0.5 | 6.99 | A | A | A |
| 17.9 | 1.8 | -10.1 | 11.73 | A | A | A |
| 18.2 | 1.29 | -8.5 | 9.31 | A | A | A |
| 18.5 | 0.31 | -7.0 | 6.26 | A | A | A |
| 16.93 | 0.85 | -14.9 | 7.85 | A | A | A |
| 18.3 | 1 | -8.0 | 8.14 | A | A | A |
| 21.4 | 0.86 | 7.5 | 7.25 | A | A | A |
| 20 | 2.3 | 0.5 | 12.99 | A | A | A |
| 21.54 | 0.77 | 8.2 | 7.01 | A | A | A |
| 20.69 | 1.17 | 4.0 | 8.27 | A | A | A |
| 19 | 2.1 | -4.5 | 12.59 | A | A | A |
| 18.41 | 0.82 | -7.5 | 7.50 | A | A | A |
| 16.85 | 2.43 | -15.3 | 15.63 | A | A | A |
| 17.3 | 0.6 | -13.1 | 6.96 | A | A | A |
| 17.4 | 0.6 | -12.6 | 6.95 | A | A | A |
| 18.43 | 0.86 | -7.4 | 7.62 | A | A | A |
| 16.6 | 0.76 | -16.6 | 7.57 | A | A | A |
| 19.3 | 1.5 | -3.0 | 9.84 | A | A | A |
| 18.19 | 1.8 | -8.6 | 11.59 | A | A | A |
| 20.37 | 1.07 | 2.4 | 8.00 | A | A | A |
| 18.2 | 1.8 | -8.5 | 11.58 | A | A | A |
| 19.2 | 1.6 | -3.5 | 10.29 | A | A | A |
| 17.5 | 1 | -12.1 | 8.31 | A | A | A |
| 19.82 | 0.83 | -0.4 | 7.34 | A | A | A |
| 19.5 | 1.2 | -2.0 | 8.62 | A | A | A |
| 15.7 | 0.8 | -21.1 | 7.89 | N | N | N |
| 18.3 | 0.9 | -8.0 | 7.78 | A | A | A |
| 20.09 | 0.57 | 0.9 | 6.66 | A | A | A |
| 21.1 | 2.1 | 6.0 | 11.64 | A | A | A |
| 9.3 | 2.5 | -53.3 | 27.55 | N | N | N |
| 18.39 | 0.94 | -7.6 | 7.91 | A | A | A |
| 20.44 | 0.85 | 2.7 | 7.33 | A | A | A |
| 19.7 | 0.8 | -1.0 | 7.27 | A | A | A |
| 22.4 | 2.4 | 12.6 | 12.29 | A | A | A |
| 19.5 | 2.25 | -2.0 | 13.02 | A | A | A |
| 19.6 | 2.1 | -1.5 | 12.29 | A | A | A |
| 22.5 | 1.6 | 13.1 | 9.32 | A | A | A |

Labcode

322

323

325

326

328

329

330

331

332

338

339

341

344

346

347

348

350

351

354

355

356

357

358

363

364

368

369

371

374

376

379

381

382

384

385

386

387

388

391

392

393

395

396

398

401

402

403

404

405

407

408

410

411

412

414

417

419

420

421

423

| Reported value [Bq/kg ] 20 | Reported uncertainty (k = 1) [Bq/kg ] 1.1 | Relative bias [%] 0.5 | P-Test [%] 8.16 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 19.327 | 1.65 | -2.9 | 10.45 | A | A | A |
| 20.3 | 1.3 | 2.0 | 8.80 | A | A | A |
| 16.6 | 0.226 | -16.6 | 6.18 | A | N | W |
| 19.08 | 0.57 | -4.1 | 6.73 | A | A | A |
| 15.8 | 0.79 | -20.6 | 7.83 | N | N | N |
| 16 | 1.6 | -19.6 | 11.68 | A | A | A |
| 19.5 | 1.3 | -2.0 | 8.99 | A | A | A |
| 18.3 | 1 | -8.0 | 8.14 | A | A | A |
| 17.5 | 1 | -12.1 | 8.31 | A | A | A |
| 19.2 | 0.7 | -3.5 | 7.05 | A | A | A |
| 18.82 | 0.87 | -5.4 | 7.60 | A | A | A |
| 20.4 | 1.4 | 2.5 | 9.14 | A | A | A |
| 18 | 0.6 | -9.6 | 6.89 | A | A | A |
| 17.39 | 0.53 | -12.6 | 6.76 | A | A | A |
| 19.8 | 0.6 | -0.5 | 6.75 | A | A | A |
| 19.38 | 1.2 | -2.6 | 8.64 | A | A | A |
| 19 | 0.654 | -4.5 | 6.94 | A | A | A |
| 20.4 | 1.2 | 2.5 | 8.42 | A | A | A |
| 20.07 | 2.01 | 0.8 | 11.69 | A | A | A |
| 19.9 | 1.7 | 0.0 | 10.46 | A | A | A |
| 26 | 3 | 30.6 | 13.02 | N | A | N |
| 21.4 | 1.3 | 7.5 | 8.56 | A | A | A |
| 20.21 | 1.15 | 1.6 | 8.29 | A | A | A |
| 19.7 | 1.8 | -1.0 | 10.95 | A | A | A |
| 20.976 | 0.7375 | 5.4 | 6.98 | A | A | A |
| 19.2 | 1.1 | -3.5 | 8.32 | A | A | A |
| 18.5 | 0.79 | -7.0 | 7.39 | A | A | A |
| 21.9 | 1.7 | 10.1 | 9.83 | A | A | A |
| 21.6 | 1.1 | 8.5 | 7.89 | A | A | A |
| 17.5 | 1.3 | -12.1 | 9.57 | A | A | A |
| 20.1 | 1.1 | 1.0 | 8.14 | A | A | A |
| 20.78 | 1.8 | 4.4 | 10.55 | A | A | A |
| 17.5 | 0.8 | -12.1 | 7.57 | A | A | A |
| 17.5 | 1 | -12.1 | 8.31 | A | A | A |
| 19.7 | 0.9 | -1.0 | 7.57 | A | A | A |
| 20.1 | 0.8 | 1.0 | 7.23 | A | A | A |
| 15.531 | 2.214 | -22.0 | 15.48 | N | A | N |
| 18.05 | 0.9 | -9.3 | 7.82 | A | A | A |
| 13.39 | 0.95 | -32.7 | 9.31 | N | N | N |
| 17.75 | 0.39 | -10.8 | 6.42 | A | A | A |
| 15.97 | 1.97 | -19.8 | 13.73 | A | A | A |
| 13.7 | 1.4 | -31.2 | 11.87 | N | N | N |
| 17.47 | 1.342 | -12.2 | 9.77 | A | A | A |
| 16 | 1 | -19.6 | 8.68 | A | A | A |
| 13.311 | 0.128 | -33.1 | 6.11 | N | N | N |
| 20.47 | 0.46 | 2.9 | 6.44 | A | A | A |
| 19.73 | 0.3 | -0.8 | 6.22 | A | A | A |
| 20.5 | 1 | 3.0 | 7.76 | A | A | A |
| 20.4 | 1.4 | 2.5 | 9.14 | A | A | A |
| 18 | 0.5 | -9.6 | 6.64 | A | A | A |
| 15.8 | 0.7 | -20.6 | 7.48 | N | N | N |
| 18.9 | 1.3 | -5.0 | 9.15 | A | A | A |
| 18.6 | 1.3 | -6.5 | 9.23 | A | A | A |
| 19.3 | 1.2 | -3.0 | 8.66 | A | A | A |
| 11.98 | 0.77 | -39.8 | 8.81 | N | N | N |
| 17.175 | 0.86 | -13.7 | 7.84 | A | A | A |
| 19.81545 | 0.962996 | -0.4 | 7.74 | A | A | A |
| 12.5 | 0.7 | -37.2 | 8.23 | N | N | N |
| 17.7 | 1.8 | -11.1 | 11.82 | A | A | A |

Labcode

425

426

428

429

433

434

435

437

439

441

443

444

445

446

448

450

452

453

454

459

460

462

463

464

465

466

469

471

472

475

477

478

479

482

483

484

487

491

492

495

496

498

499

500

501

507

508

509

514

516

522

523

525

527

528

529

530

532

533

535

| Reported value [Bq/kg ] 22.2 | Reported uncertainty (k = 1) [Bq/kg ] 1.6 | Relative bias [%] 11.6 | P-Test [%] 9.40 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 21.49 | 1.14 | 8.0 | 8.03 | A | A | A |
| 19.42 | 0.86 | -2.4 | 7.48 | A | A | A |
| 20.71 | 0.62 | 4.1 | 6.73 | A | A | A |
| 19.4 | 1.1 | -2.5 | 8.28 | A | A | A |
| 21.7 | 1.5 | 9.1 | 9.17 | A | A | A |

Labcode

536

537

539

544

553

555

Co-60 determination of activity concentration in Sample 1

Table 6. Target values

| Parameter | Value |
| --- | --- |
| Activity concentration [ Bq/kg ] | 35.6 |
| Uncertainty [ Bq/kg ] | 1.5 |
| Maximum Acceptable Relative Bias [%] | 20 |

Accepted 80%	Not Accepted 3%	Warning 2%	Not Reported 15%

45

50

Figure 6. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

25

30

84

383

499

81

495

523

180

143

520

211

184

147

61

181

95

49

303

27

70

34

391

179

159

32

265

9

291

206

217

113

348

454

185

539

313

429

452

110

346

347

455

218

166

197

100

134

450

41

183

555

532

386

189

408

444

156

250

398

453

116

257

284

168

286

393

411

68

272

357

465

256

60

319

402

477

42

191

261

283

368

441

74

177

277

294

469

553

8

66

175

403

496

13

344

395

479

45

50

220

242

369

69

65

104

170

238

316

428

446

528

15

62

325

407

478

342

264

471

525

48

253

288

17

224

354

385

508

190

332

330

271

382

233

530

258

440

202

97

420

318

58

367

198

445

223

172

260

132

459

363

204

310

38

152

304

51

315

150

358

442

339

376

536

544

78

463

317

323

Laboratory code

Figure 7. Reported results and evaluation parameters for Co-60. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 7. Evaluation results for Co-60 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

Labcode

Reported value

[Bq/kg ]

Reported uncertainty

(k = 1) [Bq/kg ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

Activity concentration  [ Bq/kg ]

35

40

2	34.8	2.7	-2.2	8.83	A	A	A

8	35.68	2.13	0.2	7.31	A	A	A

9	33.7	4	-5.3	12.60	A	A	A

| Reported value [Bq/kg ] 36.5 | Reported uncertainty (k = 1) [Bq/kg ] 1.63 | Relative bias [%] 2.5 | P-Test [%] 6.14 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 34.68 | 2.85 | -2.6 | 9.24 | A | A | A |
| 35.8 | 1.4 | 0.6 | 5.75 | A | A | A |
| 36.1 | 1.7 | 1.4 | 6.32 | A | A | A |
| 36.39 | 1.04 | 2.2 | 5.09 | A | A | A |
| 36 | 1.9 | 1.1 | 6.75 | A | A | A |
| 35.5 | 0.32 | -0.3 | 4.31 | A | A | A |
| 33 | 2 | -7.3 | 7.38 | A | A | A |
| 33.4 | 1.41 | -6.2 | 5.96 | A | A | A |
| 34.9 | 1 | -2.0 | 5.10 | A | A | A |
| 33.05 | 0.97 | -7.2 | 5.13 | A | A | A |
| 35.89 | 3.089 | 0.8 | 9.58 | A | A | A |
| 35.7 | 1.12 | 0.3 | 5.25 | A | A | A |
| 33.5 | 0.33 | -5.9 | 4.33 | A | A | A |
| 33.2 | 2.2 | -6.7 | 7.85 | A | A | A |
| 34.333976 | 2.732651 | -3.6 | 9.01 | A | A | A |
| 37 | 1.1 | 3.9 | 5.16 | A | A | A |
| 38.04 | 2.93 | 6.8 | 8.78 | A | A | A |
| 35.8 | 1.59 | 0.6 | 6.12 | A | A | A |
| 34.9 | 0.444 | -2.0 | 4.40 | A | A | A |
| 35.5 | 1.6 | -0.3 | 6.17 | A | A | A |
| 35.88 | 1.2 | 0.8 | 5.38 | A | A | A |
| 36.3 | 1.8 | 2.0 | 6.51 | A | A | A |
| 33 | 4 | -7.3 | 12.83 | A | A | A |
| 35.9 | 1 | 0.8 | 5.05 | A | A | A |
| 38.27 | 1.51 | 7.5 | 5.77 | A | A | A |
| 36.1 | 1.05 | 1.4 | 5.12 | A | A | A |
| 35.72 | 1.2 | 0.3 | 5.39 | A | A | A |
| 33.72 | 0.33 | -5.3 | 4.33 | A | A | A |
| 36.9 | 3.7 | 3.6 | 10.88 | A | A | A |
| 37 | 1 | 3.9 | 5.01 | A | A | A |
| 35.4 | 2.8 | -0.6 | 8.96 | A | A | A |
| 32.54 | 1.65 | -8.6 | 6.59 | A | A | A |
| 36.1 | 2 | 1.4 | 6.96 | A | A | A |
| 36.91 | 1.87 | 3.7 | 6.59 | A | A | A |
| 36 | 2.5 | 1.1 | 8.12 | A | A | A |
| 35.7 | 1.844 | 0.3 | 6.67 | A | A | A |
| 33.8 | 4.1 | -5.1 | 12.84 | A | A | A |
| 35.3 | 0.85 | -0.8 | 4.85 | A | A | A |
| 35.922155 | 1.5 | 0.9 | 5.93 | A | A | A |
| 33.1 | 1.7 | -7.0 | 6.64 | A | A | A |
| 37.28 | 1.86 | 4.7 | 6.53 | A | A | A |
| 38.1 | 1.82 | 7.0 | 6.37 | A | A | A |
| 35.1 | 1 | -1.4 | 5.09 | A | A | A |
| 35.6 | 3.54 | 0.0 | 10.80 | A | A | A |
| 35.32 | 2.11 | -0.8 | 7.31 | A | A | A |
| 40.75 | 1.67 | 14.5 | 5.88 | A | A | A |
| 30.2 | 0.4 | -15.2 | 4.42 | A | N | W |
| 36.723 | 0.6848 | 3.1 | 4.61 | A | A | A |
| 28.967 | 0.84 | -18.6 | 5.11 | A | N | W |
| 34.99 | 0.49 | -1.7 | 4.44 | A | A | A |
| 35 | 5.1 | -1.7 | 15.17 | A | A | A |
| 14.97 | 1.07 | -58.0 | 8.30 | N | N | N |
| 41.173 | 3.071 | 15.7 | 8.57 | A | A | A |
| 36 | 1.7 | 1.1 | 6.33 | A | A | A |
| 39.17 | 1.02 | 10.0 | 4.95 | A | A | A |
| 32.9 | 2.44 | -7.6 | 8.53 | A | A | A |
| 36.9 | 1.9 | 3.6 | 6.65 | A | A | A |
| 34.73 | 1.07 | -2.4 | 5.22 | A | A | A |
| 27.9 | 0.47 | -21.6 | 4.54 | N | N | N |

Labcode

11

12

13

15

17

19

20

21

24

25

27

28

31

32

34

35

37

38

40

41

42

45

48

49

50

51

52

53

55

57

58

60

61

62

63

65

66

67

68

69

70

71

72

73

74

75

78

79

80

81

82

83

84

85

90

94

95

97

100

102

| Reported value [Bq/kg ] 36 | Reported uncertainty (k = 1) [Bq/kg ] 1 | Relative bias [%] 1.1 | P-Test [%] 5.05 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 36 | 1 | 1.1 | 5.05 | A | A | A |
| 35.3 | 2.1 | -0.8 | 7.29 | A | A | A |
| 34.4 | 0.2 | -3.4 | 4.25 | A | A | A |
| 34.455758 | 1.722635 | -3.2 | 6.54 | A | A | A |
| 35.34 | 2.85 | -0.7 | 9.10 | A | A | A |
| 34.25 | 3.43 | -3.8 | 10.86 | A | A | A |
| 35.7 | 2.2 | 0.3 | 7.47 | A | A | A |
| 35.12 | 1.77 | -1.4 | 6.57 | A | A | A |
| 36.25 | 3.15 | 1.8 | 9.66 | A | A | A |
| 34.3 | 1.9 | -3.6 | 6.96 | A | A | A |
| 37.7 | 2 | 5.9 | 6.77 | A | A | A |
| 37.399723 | 2.961246 | 5.1 | 8.97 | A | A | A |
| 32.46 | 2.12 | -8.8 | 7.77 | A | A | A |
| 36.57 | 0.89 | 2.7 | 4.87 | A | A | A |
| 37.6 | 1.8 | 5.6 | 6.38 | A | A | A |
| 36.7 | 1 | 3.1 | 5.02 | A | A | A |
| 34.8 | 0.3 | -2.2 | 4.30 | A | A | A |
| 35.6 | 1.02 | 0.0 | 5.10 | A | A | A |
| 35.5 | 2.1 | -0.3 | 7.26 | A | A | A |
| 31.472226 | 1.299474 | -11.6 | 5.90 | A | A | A |
| 35.2 | 5.62 | -1.1 | 16.51 | A | A | A |
| 32.4487 | 0.07153 | -8.8 | 4.22 | A | A | A |
| 38.6 | 2.4 | 8.4 | 7.51 | A | A | A |
| 36.39 | 3.64 | 2.2 | 10.85 | A | A | A |
| 38.1 | 2.3 | 7.0 | 7.36 | A | A | A |
| 36.3 | 2 | 2.0 | 6.94 | A | A | A |
| 33.6 | 1.6 | -5.6 | 6.36 | A | A | A |
| 35.1 | 1 | -1.4 | 5.09 | A | A | A |
| 33 | 4 | -7.3 | 12.83 | A | A | A |
| 33.4 | 1.05 | -6.2 | 5.26 | A | A | A |
| 35.9 | 1.4 | 0.8 | 5.74 | A | A | A |
| 36.85 | 1 | 3.5 | 5.01 | A | A | A |
| 34.7 | 0.99 | -2.5 | 5.09 | A | A | A |
| 35.2 | 2.8 | -1.1 | 9.00 | A | A | A |
| 36 | 1.5 | 1.1 | 5.93 | A | A | A |
| 37.4 | 1.5 | 5.1 | 5.82 | A | A | A |
| 35.7 | 1.5 | 0.3 | 5.95 | A | A | A |
| 35.6 | 0.7 | 0.0 | 4.65 | A | A | A |
| 34.9 | 2.1 | -2.0 | 7.35 | A | A | A |
| 33.36 | 0.8 | -6.3 | 4.85 | A | A | A |
| 31.14 | 0.95 | -12.5 | 5.20 | A | A | A |
| 32.8 | 2.64 | -7.9 | 9.08 | A | A | A |
| 34.9 | 1.3 | -2.0 | 5.62 | A | A | A |
| 32.33 | 3.83 | -9.2 | 12.57 | A | A | A |
| 34.37 | 0.33 | -3.5 | 4.32 | A | A | A |
| 36 | 2.5 | 1.1 | 8.12 | A | A | A |
| 35 | 7 | -1.7 | 20.44 | A | N | W |
| 36.5 | 0.9 | 2.5 | 4.88 | A | A | A |
| 35.5 | 2.9 | -0.3 | 9.19 | A | A | A |
| 32.11 | 1.52 | -9.8 | 6.34 | A | A | A |
| 34.71 | 2.78 | -2.5 | 9.05 | A | A | A |
| 37 | 3 | 3.9 | 9.14 | A | A | A |
| 35.4 | 1.1 | -0.6 | 5.24 | A | A | A |
| 36.85 | 0.28 | 3.5 | 4.28 | A | A | A |
| 34.7 | 2.6 | -2.5 | 8.60 | A | A | A |
| 38 | 1 | 6.7 | 4.97 | A | A | A |
| 37.4 | 2.01 | 5.1 | 6.83 | A | A | A |
| 34 | 3.6 | -4.5 | 11.40 | A | A | A |
| 34.2 | 2.2 | -3.9 | 7.69 | A | A | A |

Labcode

104

105

107

109

110

111

113

114

116

119

121

124

127

128

130

132

133

134

139

140

143

145

147

150

151

152

154

155

156

158

159

161

165

166

168

170

172

175

177

178

179

180

181

183

184

185

187

189

190

191

194

197

198

200

202

203

204

205

206

207

| Reported value [Bq/kg ] 36.2 | Reported uncertainty (k = 1) [Bq/kg ] 1.13 | Relative bias [%] 1.7 | P-Test [%] 5.24 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 35 | 2.5 | -1.7 | 8.29 | A | A | A |
| 32.15 | 1.02 | -9.7 | 5.27 | A | A | A |
| 34.5 | 1.66 | -3.1 | 6.40 | A | A | A |
| 38 | 0.1 | 6.7 | 4.22 | A | A | A |
| 34.1 | 4 | -4.2 | 12.46 | A | A | A |
| 34.64352 | 0.99937 | -2.7 | 5.11 | A | A | A |
| 35.9 | 1.8 | 0.8 | 6.55 | A | A | A |
| 37.3 | 2.9 | 4.8 | 8.84 | A | A | A |
| 36.4 | 3.3 | 2.2 | 10.00 | A | A | A |
| 33.16 | 0.96 | -6.8 | 5.11 | A | A | A |
| 35.1 | 1.1 | -1.4 | 5.25 | A | A | A |
| 29.6 | 0.9 | -16.9 | 5.20 | A | N | W |
| 34.083 | 1.9758 | -4.3 | 7.17 | A | A | A |
| 35.7 | 0.8 | 0.3 | 4.77 | A | A | A |
| 35.9 | 1.2 | 0.8 | 5.38 | A | A | A |
| 36.7 | 2.4 | 3.1 | 7.78 | A | A | A |
| 35.5 | 2.5 | -0.3 | 8.21 | A | A | A |
| 36 | 1.53 | 1.1 | 5.98 | A | A | A |
| 35.6 | 1.1 | 0.0 | 5.23 | A | A | A |
| 36.48 | 2.43 | 2.5 | 7.88 | A | A | A |
| 34.6 | 1.5 | -2.8 | 6.05 | A | A | A |
| 35.9 | 0.9 | 0.8 | 4.90 | A | A | A |
| 33.03 | 1.944 | -7.2 | 7.24 | A | A | A |
| 35.133278 | 2.960334 | -1.3 | 9.42 | A | A | A |
| 35.1 | 2.6 | -1.4 | 8.52 | A | A | A |
| 24.57 | 1.01 | -31.0 | 5.89 | N | N | N |
| 36.3 | 1.8 | 2.0 | 6.51 | A | A | A |
| 36.79 | 2.82 | 3.3 | 8.75 | A | A | A |
| 35.32 | 1.68 | -0.8 | 6.35 | A | A | A |
| 35.16 | 0.57 | -1.2 | 4.51 | A | A | A |
| 36.8 | 1.2 | 3.4 | 5.33 | A | A | A |
| 37.5 | 2 | 5.3 | 6.80 | A | A | A |
| 35.5 | 1 | -0.3 | 5.07 | A | A | A |
| 33.22 | 0.83 | -6.7 | 4.90 | A | A | A |
| 36.2 | 2.9 | 1.7 | 9.05 | A | A | A |
| 33.6 | 4 | -5.6 | 12.63 | A | A | A |
| 35.17 | 3.02 | -1.2 | 9.56 | A | A | A |
| 35.5 | 1.8 | -0.3 | 6.59 | A | A | A |
| 36.5 | 1.1 | 2.5 | 5.18 | A | A | A |
| 36.6 | 2.2 | 2.8 | 7.34 | A | A | A |
| 35.3 | 1.6 | -0.8 | 6.19 | A | A | A |
| 35.9 | 1.743518 | 0.8 | 6.43 | A | A | A |
| 36.1 | 1.9 | 1.4 | 6.74 | A | A | A |
| 35.6 | 1.1 | 0.0 | 5.23 | A | A | A |
| 39.8 | 4 | 11.8 | 10.90 | A | A | A |
| 36.6 | 1.2 | 2.8 | 5.34 | A | A | A |
| 43.19 | 1.37 | 21.3 | 5.27 | N | N | N |
| 34.43 | 0.56 | -3.3 | 4.52 | A | A | A |
| 35.5 | 0.8 | -0.3 | 4.78 | A | A | A |
| 35.199 | 0.4797 | -1.1 | 4.43 | A | A | A |
| 35.2 | 1.8 | -1.1 | 6.63 | A | A | A |
| 35.2 | 0.7 | -1.1 | 4.66 | A | A | A |
| 36.3 | 1 | 2.0 | 5.03 | A | A | A |
| 36.3 | 0.85 | 2.0 | 4.82 | A | A | A |
| 36 | 1.1 | 1.1 | 5.20 | A | A | A |
| 33.78 | 3.547 | -5.1 | 11.31 | A | A | A |
| 34.9 | 1 | -2.0 | 5.10 | A | A | A |
| 35.6 | 2.8 | 0.0 | 8.92 | A | A | A |
| 35.6 | 3.2 | 0.0 | 9.93 | A | A | A |

Labcode

208

209

211

213

214

217

218

220

223

224

226

227

228

229

230

232

233

235

238

239

240

241

242

247

249

250

252

253

254

256

257

258

260

261

262

264

265

267

268

270

271

272

274

276

277

278

279

280

281

283

284

285

286

287

288

290

291

292

293

294

| Reported value [Bq/kg ] 36.04 | Reported uncertainty (k = 1) [Bq/kg ] 1.2 | Relative bias [%] 1.2 | P-Test [%] 5.37 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 36.9 | 1.29 | 3.6 | 5.47 | A | A | A |
| 37.5 | 1.8 | 5.3 | 6.39 | A | A | A |
| 37 | 3 | 3.9 | 9.14 | A | A | A |
| 31.3 | 0.84 | -12.1 | 5.00 | A | A | A |
| 38.1 | 3.1 | 7.0 | 9.16 | A | A | A |
| 33 | 2 | -7.3 | 7.38 | A | A | A |
| 38.1 | 2.4 | 7.0 | 7.58 | A | A | A |
| 35.3 | 5.3 | -0.8 | 15.59 | A | A | A |
| 36.4 | 1.4 | 2.2 | 5.70 | A | A | A |
| 38 | 2.7 | 6.7 | 8.26 | A | A | A |
| 37.9 | 0.89 | 6.5 | 4.82 | A | A | A |
| 35.2 | 1.5 | -1.1 | 5.99 | A | A | A |
| 34.4 | 1.84 | -3.4 | 6.81 | A | A | A |
| 35.84 | 1.07 | 0.7 | 5.16 | A | A | A |
| 38.367 | 1.762 | 7.8 | 6.23 | A | A | A |
| 36 | 3.5 | 1.1 | 10.60 | A | A | A |
| 42.8 | 2.5 | 20.2 | 7.20 | N | N | N |
| 36.91 | 0.77 | 3.7 | 4.70 | A | A | A |
| 35.4 | 1.5 | -0.6 | 5.98 | A | A | A |
| 35.4 | 0.7 | -0.6 | 4.65 | A | A | A |
| 46.2 | 2.3 | 29.8 | 6.52 | N | N | N |
| 36.1 | 1.76 | 1.4 | 6.44 | A | A | A |
| 35.04 | 1.6 | -1.6 | 6.21 | A | A | A |
| 34.8 | 2.5 | -2.2 | 8.33 | A | A | A |
| 35.5 | 0.9 | -0.3 | 4.92 | A | A | A |
| 36.51 | 1.9 | 2.6 | 6.70 | A | A | A |
| 32.3 | 1 | -9.3 | 5.23 | A | A | A |
| 36.5 | 2.7 | 2.5 | 8.51 | A | A | A |
| 38.88 | 2.68 | 9.2 | 8.08 | A | A | A |
| 39.4 | 0.8 | 10.7 | 4.68 | A | A | A |
| 36.8 | 1 | 3.4 | 5.01 | A | A | A |
| 36.16 | 0.28 | 1.6 | 4.28 | A | A | A |
| 35.8 | 3.1 | 0.6 | 9.63 | A | A | A |
| 34.5 | 1 | -3.1 | 5.11 | A | A | A |
| 34.6 | 1.4 | -2.8 | 5.84 | A | A | A |
| 34.3 | 1.4 | -3.6 | 5.87 | A | A | A |
| 34.3 | 0.6 | -3.6 | 4.56 | A | A | A |
| 34.6 | 1.8 | -2.8 | 6.69 | A | A | A |
| 36.4 | 1.2 | 2.2 | 5.35 | A | A | A |
| 33.47 | 0.26 | -6.0 | 4.28 | A | A | A |
| 36 | 1.3 | 1.1 | 5.55 | A | A | A |
| 35.3 | 1.8 | -0.8 | 6.61 | A | A | A |
| 38.9 | 1.9 | 9.3 | 6.45 | A | A | A |
| 37.7 | 1.3 | 5.9 | 5.44 | A | A | A |
| 38.2 | 3.8 | 7.3 | 10.80 | A | A | A |
| 37 | 1.3 | 3.9 | 5.49 | A | A | A |
| 35.5 | 2.5 | -0.3 | 8.21 | A | A | A |
| 35.9 | 0.37 | 0.8 | 4.34 | A | A | A |
| 33.07 | 1.61 | -7.1 | 6.44 | A | A | A |
| 37.2 | 1.8 | 4.5 | 6.42 | A | A | A |
| 39.5 | 1.6 | 11.0 | 5.84 | A | A | A |
| 34.4 | 4.1 | -3.4 | 12.64 | A | A | A |
| 38.49 | 1.27 | 8.1 | 5.35 | A | A | A |
| 36.63 | 1.56 | 2.9 | 5.99 | A | A | A |
| 25.7 | 0.9 | -27.8 | 5.48 | N | N | N |
| 35.8 | 3.9 | 0.6 | 11.68 | A | A | A |
| 36.41 | 1.25 | 2.3 | 5.44 | A | A | A |
| 34.99 | 3.41 | -1.7 | 10.62 | A | A | A |
| 35.1 | 1 | -1.4 | 5.09 | A | A | A |

Labcode

295

296

297

299

300

302

303

304

306

307

310

311

312

313

314

315

316

317

318

319

322

323

325

326

328

329

330

331

332

338

339

341

342

344

346

347

348

350

351

354

355

356

357

358

363

364

367

368

369

371

374

376

379

381

382

383

384

385

386

387

| Reported value [Bq/kg ] 35.3 | Reported uncertainty (k = 1) [Bq/kg ] 1 | Relative bias [%] -0.8 | P-Test [%] 5.08 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 33.22 | 0.96 | -6.7 | 5.11 | A | A | A |
| 36.1 | 1.59 | 1.4 | 6.10 | A | A | A |
| 35.2 | 1.3 | -1.1 | 5.60 | A | A | A |
| 35.8 | 3.5 | 0.6 | 10.65 | A | A | A |
| 36.35 | 1.63 | 2.1 | 6.15 | A | A | A |
| 35.1 | 2.4 | -1.4 | 8.03 | A | A | A |
| 37.6 | 2.1 | 5.6 | 7.00 | A | A | A |
| 35.4 | 1.8 | -0.6 | 6.60 | A | A | A |
| 35.7 | 1.03 | 0.3 | 5.11 | A | A | A |
| 35.6 | 1.9 | 0.0 | 6.80 | A | A | A |
| 31.6 | 1.7 | -11.2 | 6.83 | A | A | A |
| 36.1 | 2 | 1.4 | 6.96 | A | A | A |
| 35.01 | 0.48 | -1.7 | 4.43 | A | A | A |
| 35.2 | 3.4 | -1.1 | 10.54 | A | A | A |
| 35.2 | 1.1 | -1.1 | 5.25 | A | A | A |
| 35.65 | 1.55 | 0.1 | 6.05 | A | A | A |
| 36.4 | 1.5 | 2.2 | 5.89 | A | A | A |
| 35.2 | 1.4 | -1.1 | 5.79 | A | A | A |
| 42.5 | 3 | 19.4 | 8.22 | A | A | A |
| 36.9 | 3.65 | 3.6 | 10.75 | A | A | A |
| 36.2 | 3.9 | 1.7 | 11.57 | A | A | A |
| 36.1 | 1.4 | 1.4 | 5.73 | A | A | A |
| 34.71 | 2 | -2.5 | 7.14 | A | A | A |
| 36 | 2.5 | 1.1 | 8.12 | A | A | A |
| 34.4 | 0.292 | -3.4 | 4.30 | A | A | A |
| 34.97 | 0.75 | -1.8 | 4.73 | A | A | A |
| 32.87 | 0.8 | -7.7 | 4.87 | A | A | A |
| 35.4 | 3 | -0.6 | 9.46 | A | A | A |
| 35.1 | 2.1 | -1.4 | 7.32 | A | A | A |
| 35.5 | 1.6 | -0.3 | 6.17 | A | A | A |
| 36.84 | 0.74 | 3.5 | 4.67 | A | A | A |
| 35.5 | 1.4 | -0.3 | 5.77 | A | A | A |
| 39.069747 | 2.121107 | 9.8 | 6.87 | A | A | A |
| 36 | 1.5 | 1.1 | 5.93 | A | A | A |
| 35.06 | 1.52 | -1.5 | 6.05 | A | A | A |
| 37.2 | 1.3 | 4.5 | 5.47 | A | A | A |
| 36 | 1.1 | 1.1 | 5.20 | A | A | A |
| 34.37 | 0.33 | -3.5 | 4.32 | A | A | A |
| 34.8 | 1 | -2.2 | 5.10 | A | A | A |
| 34.42 | 2 | -3.3 | 7.18 | A | A | A |
| 35.1 | 0.8 | -1.4 | 4.79 | A | A | A |
| 34.3 | 1.6 | -3.6 | 6.29 | A | A | A |
| 34.6 | 2.7 | -2.8 | 8.87 | A | A | A |
| 37.66 | 3.77 | 5.8 | 10.86 | A | A | A |
| 37 | 3 | 3.9 | 9.14 | A | A | A |
| 49 | 3 | 37.6 | 7.43 | N | N | N |
| 41.5 | 2.1 | 16.6 | 6.58 | A | A | A |
| 38.31 | 1.63 | 7.6 | 5.99 | A | A | A |
| 35.3 | 3.5 | -0.8 | 10.77 | A | A | A |
| 39.485 | 1.19915 | 10.9 | 5.19 | A | A | A |
| 35.6 | 1.7 | 0.0 | 6.37 | A | A | A |
| 36.2 | 1.2 | 1.7 | 5.36 | A | A | A |
| 40.5 | 2.2 | 13.8 | 6.87 | A | A | A |
| 38.9 | 1.9 | 9.3 | 6.45 | A | A | A |
| 35.4 | 1.4 | -0.6 | 5.78 | A | A | A |
| 36.1 | 1.8 | 1.4 | 6.53 | A | A | A |
| 35.81 | 2.89 | 0.6 | 9.10 | A | A | A |
| 35.1 | 1.2 | -1.4 | 5.43 | A | A | A |
| 35.5 | 1.4 | -0.3 | 5.77 | A | A | A |

Labcode

388

391

392

393

395

396

398

401

402

403

404

405

407

408

410

411

412

414

417

419

420

421

425

426

428

429

433

434

435

437

439

440

441

442

443

444

445

446

448

450

452

453

454

455

459

460

462

463

464

465

466

469

471

472

475

477

478

479

482

483

| Reported value [Bq/kg ] 36.1 | Reported uncertainty (k = 1) [Bq/kg ] 1.1 | Relative bias [%] 1.4 | P-Test [%] 5.20 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 36.2 | 1.3 | 1.7 | 5.54 | A | A | A |
| 36.5 | 1.6 | 2.5 | 6.08 | A | A | A |
| 29.76 | 0.74 | -16.4 | 4.89 | A | N | W |
| 35.71 | 0.47 | 0.3 | 4.41 | A | A | A |
| 32.44 | 2.13 | -8.9 | 7.80 | A | A | A |
| 28.2 | 2.8 | -20.8 | 10.79 | N | A | N |
| 33.31 | 1.703 | -6.4 | 6.63 | A | A | A |
| 35.6 | 1.2 | 0.0 | 5.40 | A | A | A |
| 32.6 | 1.4 | -8.4 | 6.02 | A | A | A |
| 30.733 | 0.126 | -13.7 | 4.23 | A | N | W |
| 36.49 | 0.85 | 2.5 | 4.81 | A | A | A |
| 36 | 0.6 | 1.1 | 4.53 | A | A | A |
| 38 | 1.5 | 6.7 | 5.77 | A | A | A |
| 35.8 | 2.2 | 0.6 | 7.45 | A | A | A |
| 32.05 | 0.92 | -10.0 | 5.10 | A | A | A |
| 34.4 | 1.1 | -3.4 | 5.29 | A | A | A |
| 30.3 | 1.2 | -14.9 | 5.78 | A | A | A |
| 36.2 | 2.5 | 1.7 | 8.09 | A | A | A |
| 35.9 | 2.2 | 0.8 | 7.44 | A | A | A |
| 36 | 1.3 | 1.1 | 5.55 | A | A | A |
| 28.4 | 1.6 | -20.2 | 7.04 | N | N | N |
| 36.755 | 0.91 | 3.2 | 4.89 | A | A | A |
| 34.970975 | 1.412865 | -1.8 | 5.84 | A | A | A |
| 35.7 | 1.6 | 0.3 | 6.15 | A | A | A |
| 33.6 | 1.9 | -5.6 | 7.05 | A | A | A |
| 40.1 | 1.5 | 12.6 | 5.63 | A | A | A |
| 40.11 | 2.05 | 12.7 | 6.62 | A | A | A |
| 34.39 | 1.51 | -3.4 | 6.09 | A | A | A |
| 34.64 | 3.15 | -2.7 | 10.02 | A | A | A |
| 40.19 | 1.21 | 12.9 | 5.18 | A | A | A |
| 35.6 | 1.8 | 0.0 | 6.58 | A | A | A |
| 34.9 | 2.4 | -2.0 | 8.06 | A | A | A |

Labcode

484

487

490

495

496

498

499

500

501

506

507

508

509

514

516

520

522

523

525

527

528

529

530

532

533

535

536

537

539

540

544

553

555

Sr-90 determination of activity concentration in Sample 1

Table 8. Target values

| Parameter | Value |
| --- | --- |
| Activity concentration [ Bq/kg ] | 39.5 |
| Uncertainty [ Bq/kg ] | 2.0 |
| Maximum Acceptable Relative Bias [%] | 30 |

Accepted 39%	Not Accepted 4%	Warning 2%	Not Reported 55%

60

Figure 8. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

20

30

533

127

342

384

71

232

52

383

440

313

293

419

159

555

170

49

73

12

134

145

291

204

347

443

181

393

311

41

140

401

463

152

156

133

355

525

508

283

371

32

83

364

395

192

72

277

265

312

261

414

539

287

532

445

469

65

187

205

191

231

382

238

466

31

90

369

514

527

507

166

302

91

230

523

15

432

110

297

319

437

66

189

254

264

292

58

87

276

304

33

161

475

288

429

7

109

208

303

439

60

81

175

241

250

38

178

183

317

262

63

132

516

260

151

28

374

325

25

75

379

306

386

477

45

9

459

69

42

381

168

114

213

433

21

70

255

119

417

13

220

553

6

215

177

314

55

281

392

37

122

315

53

128

100

195

19

286

285

410

368

194

370

344

223

78

197

402

529

211

346

341

278

244

85

405

491

Laboratory code

Figure 9. Reported results and evaluation parameters for Sr-90. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 9. Evaluation results for Sr-90 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

Labcode

Reported value

[Bq/kg ]

Reported uncertainty

(k = 1) [Bq/kg ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

Activity concentration  [ Bq/kg ]

40

50

6	42.5	2.11	7.6	7.09	A	A	A

7	39.21	1.96	-0.7	7.12	A	A	A

9	40.93	3.49	3.6	9.92	A	A	A

| Reported value [Bq/kg ] 33.12 | Reported uncertainty (k = 1) [Bq/kg ] 0.83 | Relative bias [%] -16.1 | P-Test [%] 5.65 | Trueness evaluation A | Precision evaluation N | Final Score W |
| --- | --- | --- | --- | --- | --- | --- |
| 42.2 | 2.9 | 6.8 | 8.54 | A | A | A |
| 38.7 | 2.7 | -2.0 | 8.62 | A | A | A |
| 44.1 | 1.1 | 11.7 | 5.64 | A | A | A |
| 41.7 | 0.3 | 5.6 | 5.11 | A | A | A |
| 40.5 | 1.7 | 2.5 | 6.58 | A | A | A |
| 40.14 | 3.09 | 1.6 | 9.21 | A | A | A |
| 38.2 | 1.53 | -3.3 | 6.46 | A | A | A |
| 36.2 | 1.9 | -8.3 | 7.29 | A | A | A |
| 39.1 | 2.1 | -1.0 | 7.38 | A | A | A |
| 43.2 | 5 | 9.4 | 12.63 | A | A | A |
| 39.46 | 3.76 | -0.1 | 10.79 | A | A | A |
| 34.5 | 1.68 | -12.7 | 7.02 | A | A | A |
| 41.2 | 1.6 | 4.3 | 6.38 | A | A | A |
| 40.92 | 0.56 | 3.6 | 5.24 | A | A | A |
| 33 | 5 | -16.5 | 15.98 | A | A | A |
| 19.2 | 1.05 | -51.4 | 7.45 | N | N | N |
| 43.25 | 9.33 | 9.5 | 22.16 | A | A | A |
| 42.72 | 1.03 | 8.2 | 5.61 | A | A | A |
| 39 | 2 | -1.3 | 7.21 | A | A | A |
| 39.4 | 3.7 | -0.2 | 10.67 | A | A | A |
| 39.8 | 1.27 | 0.8 | 5.98 | A | A | A |
| 37.9 | 2.2 | -4.0 | 7.70 | A | A | A |
| 38.89 | 2.496 | -1.5 | 8.17 | A | A | A |
| 41.1 | 2.1 | 4.0 | 7.19 | A | A | A |
| 41.7 | 2.7 | 5.6 | 8.22 | A | A | A |
| 16.5 | 0.85 | -58.2 | 7.22 | N | N | N |
| 36.9 | 1.25 | -6.6 | 6.09 | A | A | A |
| 33.1 | 2.5 | -16.2 | 9.09 | A | A | A |
| 40.5 | 4.53 | 2.5 | 12.28 | A | A | A |
| 46.5 | 2.9 | 17.7 | 8.03 | A | A | A |
| 39.4 | 2 | -0.2 | 7.17 | A | A | A |
| 36.6 | 7.15 | -7.3 | 20.18 | A | A | A |
| 74.737 | 10.957 | 89.2 | 15.51 | N | N | N |
| 38.2 | 2.3 | -3.3 | 7.87 | A | A | A |
| 38.5 | 2.3 | -2.5 | 7.83 | A | A | A |
| 43.8 | 3.5 | 10.9 | 9.46 | A | A | A |
| 39.25 | 1.13 | -0.6 | 5.82 | A | A | A |
| 38.736326 | 1.597282 | -1.9 | 6.53 | A | A | A |
| 41.5 | 2.9 | 5.1 | 8.63 | A | A | A |
| 41.9 | 5.8 | 6.1 | 14.74 | A | A | A |
| 43.2 | 6.4 | 9.4 | 15.66 | A | A | A |
| 8.7 | 0.7 | -78.0 | 9.51 | N | N | N |
| 43.4 | 3.2 | 9.9 | 8.94 | A | A | A |
| 39.8 | 2.9 | 0.8 | 8.87 | A | A | A |
| 35.3 | 0.7 | -10.6 | 5.44 | A | A | A |
| 33.5 | 1.1 | -15.2 | 6.03 | A | A | A |
| 34.7 | 3.5 | -12.2 | 11.29 | A | A | A |
| 33.6 | 6.4 | -14.9 | 19.71 | A | A | A |
| 40.11 | 2.81 | 1.5 | 8.64 | A | A | A |
| 34.8 | 2.2 | -11.9 | 8.10 | A | A | A |
| 35 | 1.27 | -11.4 | 6.23 | A | A | A |
| 31.8 | 2.06 | -19.5 | 8.22 | A | A | A |
| 39.1 | 3.7 | -1.0 | 10.73 | A | A | A |
| 38.4 | 3.4 | -2.8 | 10.20 | A | A | A |
| 41.4 | 4 | 4.8 | 10.91 | A | A | A |
| 32.6 | 1.3 | -17.5 | 6.45 | A | N | W |
| 39.4 | 1 | -0.2 | 5.66 | A | A | A |
| 42.7 | 2.82 | 8.1 | 8.32 | A | A | A |
| 39.5 | 3.2 | 0.0 | 9.55 | A | A | A |

Labcode

12

13

15

19

21

25

28

31

32

33

37

38

41

42

45

49

52

53

55

58

60

63

65

66

69

70

71

72

73

75

78

81

83

85

90

91

100

109

110

114

119

122

127

128

132

133

134

140

145

151

152

156

159

161

166

168

170

175

177

178

| Reported value [Bq/kg ] 34.01 | Reported uncertainty (k = 1) [Bq/kg ] 6.58 | Relative bias [%] -13.9 | P-Test [%] 20.00 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 39.54 | 0.73 | 0.1 | 5.39 | A | A | A |
| 37.9 | 2.2 | -4.0 | 7.70 | A | A | A |
| 38.89 | 0.63 | -1.5 | 5.32 | A | A | A |
| 38 | 3.4 | -3.8 | 10.28 | A | A | A |
| 36.7 | 2.68 | -7.1 | 8.89 | A | A | A |
| 44.44 | 2.18 | 12.5 | 7.05 | A | A | A |
| 44 | 2.2 | 11.4 | 7.12 | A | A | A |
| 46.81 | 4.67 | 18.5 | 11.19 | A | A | A |
| 34 | 2 | -13.9 | 7.76 | A | A | A |
| 37.9 | 0.677 | -4.0 | 5.37 | A | A | A |
| 39.3 | 1.55 | -0.5 | 6.42 | A | A | A |
| 48.953 | 1.438 | 23.9 | 5.85 | A | N | W |
| 41.6 | 1.26 | 5.3 | 5.90 | A | A | A |
| 42.5 | 8.9 | 7.6 | 21.54 | A | A | A |
| 42.3 | 3.1 | 7.1 | 8.91 | A | A | A |
| 45.5 | 4.2 | 15.2 | 10.53 | A | A | A |
| 38.5 | 4.1 | -2.5 | 11.79 | A | A | A |
| 38 | 5 | -3.8 | 14.10 | A | A | A |
| 17.2 | 2.6 | -56.5 | 15.94 | N | N | N |
| 38.1 | 2.82 | -3.5 | 8.97 | A | A | A |
| 39.4 | 2 | -0.2 | 7.17 | A | A | A |
| 67.4 | 10.3 | 70.6 | 16.10 | N | N | N |
| 39.4 | 1.3 | -0.2 | 6.04 | A | A | A |
| 38.9 | 0.5 | -1.5 | 5.22 | A | A | A |
| 41.72 | 2.8 | 5.6 | 8.41 | A | A | A |
| 40.1 | 3.4 | 1.5 | 9.88 | A | A | A |
| 37.1 | 3 | -6.1 | 9.54 | A | A | A |
| 39.62 | 1.04 | 0.3 | 5.70 | A | A | A |
| 38.9 | 1.95 | -1.5 | 7.12 | A | A | A |
| 37 | 3.7 | -6.3 | 11.21 | A | A | A |
| 39 | 3 | -1.3 | 9.21 | A | A | A |
| 36.9 | 3 | -6.6 | 9.58 | A | A | A |
| 60.1 | 3.03 | 52.1 | 7.15 | N | N | N |
| 43 | 0.6 | 8.9 | 5.25 | A | A | A |
| 36 | 1.34 | -8.9 | 6.28 | A | A | A |
| 44.2 | 2.6 | 11.9 | 7.76 | A | A | A |
| 44.1 | 3.4 | 11.7 | 9.22 | A | A | A |
| 37.2 | 0.5 | -5.8 | 5.24 | A | A | A |
| 39.11 | 1.95 | -1.0 | 7.11 | A | A | A |
| 33.79 | 4.73 | -14.5 | 14.89 | A | A | A |
| 38.9 | 2.7 | -1.5 | 8.59 | A | A | A |
| 29.03 | 3.15 | -26.5 | 11.97 | A | A | A |
| 38.76 | 1.48 | -1.9 | 6.34 | A | A | A |
| 38.4 | 4.2 | -2.8 | 12.05 | A | A | A |
| 39.3 | 2 | -0.5 | 7.18 | A | A | A |
| 39 | 2.3 | -1.3 | 7.77 | A | A | A |
| 40.8 | 6.4 | 3.3 | 16.48 | A | A | A |
| 34.1 | 1.52 | -13.7 | 6.75 | A | A | A |
| 37 | 4 | -6.3 | 11.94 | A | A | A |
| 27.8 | 4.901 | -29.6 | 18.34 | A | A | A |
| 42.71 | 3.35 | 8.1 | 9.34 | A | A | A |
| 43.243 | 2.668 | 9.5 | 7.98 | A | A | A |
| 39.6 | 2.9 | 0.2 | 8.90 | A | A | A |
| 38.8 | 2.54 | -1.8 | 8.28 | A | A | A |
| 40.3 | 4.24 | 2.0 | 11.68 | A | A | A |
| 59.7 | 4.8 | 51.1 | 9.50 | N | N | N |
| 13.4 | 0.27 | -66.1 | 5.45 | N | N | N |
| 45.3 | 4.1 | 14.7 | 10.37 | A | A | A |
| 49.6 | 3.1 | 25.6 | 8.04 | A | N | W |

Labcode

181

183

187

189

191

192

194

195

197

204

205

208

211

213

215

220

223

230

231

232

238

241

244

250

254

255

260

261

262

264

265

276

277

278

281

283

285

286

287

288

291

292

293

297

302

303

304

306

311

312

313

314

315

317

319

325

341

342

344

346

| Reported value [Bq/kg ] 34 | Reported uncertainty (k = 1) [Bq/kg ] 6.8 | Relative bias [%] -13.9 | P-Test [%] 20.63 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 35.66 | 1.61 | -9.7 | 6.78 | A | A | A |
| 36.6 | 2.5 | -7.3 | 8.50 | A | A | A |
| 44.3 | 5.8 | 12.2 | 14.04 | A | A | A |
| 38.3 | 0.35 | -3.0 | 5.15 | A | A | A |
| 44.5 | 4.6 | 12.7 | 11.51 | A | A | A |
| 36 | 3 | -8.9 | 9.75 | A | A | A |
| 40.2 | 3.5 | 1.8 | 10.07 | A | A | A |
| 40.5 | 4.4 | 2.5 | 11.99 | A | A | A |
| 41.32 | 1.6 | 4.6 | 6.37 | A | A | A |
| 38 | 1.9 | -3.8 | 7.12 | A | A | A |
| 24.6 | 1.7 | -37.7 | 8.57 | N | N | N |
| 15 | 3 | -62.0 | 20.63 | N | N | N |
| 40.87 | 3.39 | 3.5 | 9.72 | A | A | A |
| 43 | 8 | 8.9 | 19.28 | A | A | A |
| 34.02 | 6.58 | -13.9 | 19.99 | A | A | A |
| 36.67 | 3.66 | -7.2 | 11.19 | A | A | A |
| 34.7 | 4.4 | -12.2 | 13.65 | A | A | A |
| 46.9 | 2.9 | 18.7 | 7.99 | A | A | A |
| 75.8 | 6.1 | 91.9 | 9.51 | N | N | N |
| 44.2 | 3.8 | 11.9 | 9.98 | A | A | A |
| 37.1 | 1.5 | -6.1 | 6.48 | A | A | A |
| 42.1 | 1.8 | 6.6 | 6.63 | A | A | A |
| 30.5 | 7 | -22.8 | 23.50 | A | A | A |
| 39.2 | 1.34 | -0.8 | 6.11 | A | A | A |
| 38.7 | 2.46 | -2.0 | 8.13 | A | A | A |
| 41.6 | 1.29 | 5.3 | 5.94 | A | A | A |
| 38.8 | 3.1 | -1.8 | 9.46 | A | A | A |
| 39.3 | 1.7 | -0.5 | 6.66 | A | A | A |
| 27.073 | 0.153 | -31.5 | 5.09 | N | N | N |
| 34 | 1.6 | -13.9 | 6.91 | A | A | A |
| 37.4 | 1.6 | -5.3 | 6.63 | A | A | A |
| 41.07 | 3.46 | 4.0 | 9.83 | A | A | A |
| 34.7 | 1.8 | -12.2 | 7.25 | A | A | A |
| 38.1 | 1.52 | -3.5 | 6.45 | A | A | A |
| 37.8 | 1.4 | -4.3 | 6.27 | A | A | A |
| 39.1 | 2.9 | -1.0 | 8.98 | A | A | A |
| 40.9 | 3.4 | 3.5 | 9.73 | A | A | A |
| 2063.089 | 16.755 | 5123.0 | 5.13 | N | N | N |
| 38.367 | 2.107 | -2.9 | 7.47 | A | A | A |
| 35.8 | 1.2 | -9.4 | 6.07 | A | A | A |
| 38.3 | 2.8 | -3.0 | 8.89 | A | A | A |
| 39.8 | 2.1 | 0.8 | 7.31 | A | A | A |
| 38.6 | 1.2 | -2.3 | 5.94 | A | A | A |
| 35.7 | 1.5 | -9.6 | 6.58 | A | A | A |
| 38.3 | 2 | -3.0 | 7.27 | A | A | A |
| 47.1 | 1.4 | 19.2 | 5.87 | A | N | W |
| 37.2 | 6 | -5.8 | 16.91 | A | A | A |
| 3.71 | 0.52 | -90.6 | 14.90 | N | N | N |
| 37.1 | 2.28 | -6.1 | 7.96 | A | A | A |
| 42.3 | 2.1 | 7.1 | 7.09 | A | A | A |
| 32.4 | 5.2 | -18.0 | 16.83 | A | A | A |

Labcode

347

355

364

368

369

370

371

374

379

381

382

383

384

386

392

393

395

401

402

405

410

414

417

419

429

432

433

437

439

440

443

445

459

463

466

469

475

477

491

507

508

514

516

523

525

527

529

532

533

539

553

555

Activity concentration  [ Bq/kg ]

35

40

Ba-133 determination of activity concentration in Sample 1

Table 10. Target values

| Parameter | Value |
| --- | --- |
| Activity concentration [ Bq/kg ] | 34.3 |
| Uncertainty [ Bq/kg ] | 2.1 |
| Maximum Acceptable Relative Bias [%] | 20 |

Accepted 75%	Not Accepted 5%	Warning 1%	Not Reported 19%

45

Figure 10. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

20

25

30

455

71

529

434

74

405

364

498

499

341

533

530

179

81

205

342

207

63

277

272

477

281

249

466

203

200

184

271

329

331

381

238

395

292

522

426

180

230

429

35

8

328

391

262

168

472

300

411

113

165

304

408

190

388

49

278

398

539

194

232

306

220

323

516

119

448

444

154

224

407

527

116

261

439

19

110

100

208

368

40

134

392

13

67

487

544

73

344

465

537

159

346

469

267

187

258

419

90

152

446

50

156

414

9

265

104

241

332

525

338

133

483

202

170

214

379

72

382

183

319

421

307

25

260

325

354

198

66

70

478

213

283

41

420

97

440

111

172

442

294

501

256

145

312

479

191

51

121

454

299

247

555

48

315

464

177

274

254

311

376

462

Laboratory code

Figure 11. Reported results and evaluation parameters for Ba-133. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 11. Evaluation results for Ba-133 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

| Reported value [Bq/kg ] | Reported uncertainty (k = 1) [Bq/kg ] | Relative bias [%] | P-Test [%] | Trueness evaluation | Precision evaluation | Final Score |
| --- | --- | --- | --- | --- | --- | --- |
| 31.5 | 2.1 | -8.2 | 9.05 | A | A | A |
| 32.4 | 2.05 | -5.5 | 8.80 | A | A | A |
| 34.7 | 4.1 | 1.2 | 13.31 | A | A | A |
| 36.8 | 2.26 | 7.3 | 8.67 | A | A | A |
| 35.68 | 1.55 | 4.0 | 7.51 | A | A | A |
| 34 | 1.4 | -0.9 | 7.38 | A | A | A |
| 34.8 | 2 | 1.5 | 8.40 | A | A | A |
| 35.48 | 1.76 | 3.4 | 7.88 | A | A | A |

Labcode

2

8

9

11

12

13

15

17

| Reported value [Bq/kg ] 33.7 | Reported uncertainty (k = 1) [Bq/kg ] 1.7 | Relative bias [%] -1.8 | P-Test [%] 7.93 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 29 | 3 | -15.4 | 12.02 | A | A | A |
| 33.1 | 2.97 | -3.5 | 10.86 | A | A | A |
| 35.4 | 1.7 | 3.2 | 7.78 | A | A | A |
| 28.44 | 0.5 | -17.1 | 6.37 | A | N | W |
| 34.13 | 2.957 | -0.5 | 10.61 | A | A | A |
| 29.8 | 0.99 | -13.1 | 6.97 | A | A | A |
| 26.8 | 0.83 | -21.9 | 6.86 | N | N | N |
| 27.8 | 2 | -19.0 | 9.45 | A | A | A |
| 32.380939 | 2.981851 | -5.6 | 11.06 | A | A | A |
| 36.9 | 1 | 7.6 | 6.70 | A | A | A |
| 37.48 | 3.42 | 9.3 | 10.99 | A | A | A |
| 33.84 | 1.83 | -1.3 | 8.17 | A | A | A |
| 35.7 | 0.628 | 4.1 | 6.37 | A | A | A |
| 33.2 | 1.3 | -3.2 | 7.27 | A | A | A |
| 31.75 | 1.02 | -7.4 | 6.91 | A | A | A |
| 37.5 | 2.2 | 9.3 | 8.48 | A | A | A |
| 33 | 3 | -3.8 | 10.96 | A | A | A |
| 34.6 | 0.4 | 0.9 | 6.23 | A | A | A |
| 36.64 | 1.38 | 6.8 | 7.19 | A | A | A |
| 37.8 | 1.85 | 10.2 | 7.84 | A | A | A |
| 34 | 3 | -0.9 | 10.74 | A | A | A |
| 31.86 | 0.79 | -7.1 | 6.61 | A | A | A |
| 38.7 | 3.9 | 12.8 | 11.79 | A | A | A |
| 36 | 2 | 5.0 | 8.27 | A | A | A |
| 32.3 | 2.1 | -5.8 | 8.93 | A | A | A |
| 49.64 | 1.88 | 44.7 | 7.20 | N | N | N |
| 33 | 1.9 | -3.8 | 8.40 | A | A | A |
| 30.64 | 1.76 | -10.7 | 8.40 | A | A | A |
| 34.4 | 2.3 | 0.3 | 9.07 | A | A | A |
| 35.47 | 1.8475 | 3.4 | 8.04 | A | A | A |
| 34 | 4.1 | -0.9 | 13.52 | A | A | A |
| 32.9 | 1.1 | -4.1 | 6.98 | A | A | A |
| 33.4957 | 1.4 | -2.3 | 7.41 | A | A | A |
| 35.5 | 1.8 | 3.5 | 7.95 | A | A | A |
| 10.21 | 0.51 | -70.2 | 7.90 | N | N | N |
| 35.1 | 2.18 | 2.3 | 8.72 | A | A | A |
| 34.1 | 1.6 | -0.6 | 7.71 | A | A | A |
| 24.8 | 2.49 | -27.7 | 11.76 | N | A | N |
| 38.8 | 2.95 | 13.1 | 9.76 | A | A | A |
| 39.7 | 4.4 | 15.7 | 12.66 | A | A | A |
| 30.7 | 0.9 | -10.5 | 6.79 | A | A | A |
| 35.056 | 0.67896 | 2.2 | 6.42 | A | A | A |
| 29.933 | 1.46 | -12.7 | 7.83 | A | A | A |
| 34.06 | 0.43 | -0.7 | 6.25 | A | A | A |
| 31.6 | 4.6 | -7.9 | 15.79 | A | A | A |
| 24.78 | 2.19 | -27.8 | 10.75 | N | N | N |
| 33.725 | 2.56 | -1.7 | 9.75 | A | A | A |
| 34.5 | 1.6 | 0.6 | 7.68 | A | A | A |
| 37 | 1.27 | 7.9 | 7.02 | A | A | A |
| 35.8 | 1.9 | 4.4 | 8.10 | A | A | A |
| 33.73 | 1.09 | -1.7 | 6.92 | A | A | A |
| 21.7 | 0.75 | -36.7 | 7.03 | N | N | N |
| 34.8 | 0.4 | 1.5 | 6.23 | A | A | A |
| 34.7 | 0.4 | 1.2 | 6.23 | A | A | A |
| 34.6 | 2.1 | 0.9 | 8.62 | A | A | A |
| 33.718864 | 1.860927 | -1.7 | 8.24 | A | A | A |
| 35.91 | 2.85 | 4.7 | 10.02 | A | A | A |
| 32.74 | 3.27 | -4.6 | 11.71 | A | A | A |
| 34.5 | 2.3 | 0.6 | 9.05 | A | A | A |

Labcode

19

21

24

25

27

28

31

32

34

35

37

38

40

41

42

45

48

49

50

51

52

53

55

57

58

60

61

62

63

65

66

67

68

69

70

71

72

73

74

75

78

79

80

81

82

83

84

85

90

94

97

100

102

104

105

107

110

111

113

114

| Reported value [Bq/kg ] 33.59 | Reported uncertainty (k = 1) [Bq/kg ] 1.75 | Relative bias [%] -2.1 | P-Test [%] 8.04 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 33.22 | 3.8 | -3.1 | 12.97 | A | A | A |
| 36.8 | 2.1 | 7.3 | 8.37 | A | A | A |
| 35 | 1.6 | 2.0 | 7.64 | A | A | A |
| 37.364749 | 3.53964 | 8.9 | 11.28 | A | A | A |
| 32.53 | 1.5 | -5.2 | 7.66 | A | A | A |
| 33.4 | 1.13 | -2.6 | 7.00 | A | A | A |
| 36.6 | 2.6 | 6.7 | 9.38 | A | A | A |
| 34.9 | 1.1 | 1.8 | 6.89 | A | A | A |
| 33.9 | 0.3 | -1.2 | 6.19 | A | A | A |
| 35.7 | 3.21 | 4.1 | 10.88 | A | A | A |
| 34.8 | 2.1 | 1.5 | 8.60 | A | A | A |
| 27.196302 | 1.303109 | -20.7 | 7.77 | N | N | N |
| 36.4 | 5.82 | 6.1 | 17.12 | A | A | A |
| 30.1467 | 0.0925 | -12.1 | 6.13 | A | A | A |
| 33.6 | 1.8 | -2.0 | 8.14 | A | A | A |
| 32.41 | 2.27 | -5.5 | 9.30 | A | A | A |
| 34.5 | 2.1 | 0.6 | 8.63 | A | A | A |
| 33.4 | 2.1 | -2.6 | 8.78 | A | A | A |
| 32.4 | 2.6 | -5.5 | 10.09 | A | A | A |
| 34.6 | 1.9 | 0.9 | 8.22 | A | A | A |
| 34 | 3 | -0.9 | 10.74 | A | A | A |
| 34.3 | 0.75 | 0.0 | 6.50 | A | A | A |
| 34.3 | 1.4 | 0.0 | 7.36 | A | A | A |
| 32.8 | 1.3 | -4.4 | 7.29 | A | A | A |
| 35.87 | 1.53 | 4.6 | 7.46 | A | A | A |
| 32.6 | 2.9 | -5.0 | 10.80 | A | A | A |
| 35 | 3 | 2.0 | 10.53 | A | A | A |
| 36 | 2 | 5.0 | 8.27 | A | A | A |
| 36.5 | 1.5 | 6.4 | 7.37 | A | A | A |
| 38.2 | 1.5 | 11.4 | 7.27 | A | A | A |
| 35.6 | 2.2 | 3.8 | 8.70 | A | A | A |
| 29.43 | 0.7 | -14.2 | 6.57 | A | A | A |
| 32.22 | 1.34 | -6.1 | 7.40 | A | A | A |
| 33.84 | 3.11 | -1.3 | 11.04 | A | A | A |
| 35.11 | 1.94 | 2.4 | 8.25 | A | A | A |
| 31.51 | 3.62 | -8.1 | 13.02 | A | A | A |
| 33.26 | 0.43 | -3.0 | 6.26 | A | A | A |
| 34.4 | 2.3 | 0.3 | 9.07 | A | A | A |
| 27.62 | 5.52 | -19.5 | 20.90 | A | N | W |
| 32.9 | 1.19 | -4.1 | 7.11 | A | A | A |
| 36.6 | 3.1 | 6.7 | 10.45 | A | A | A |
| 33.09 | 1.61 | -3.5 | 7.82 | A | A | A |
| 30.29 | 2.42 | -11.7 | 10.07 | A | A | A |
| 35.4 | 2.9 | 3.2 | 10.23 | A | A | A |
| 31.5 | 1 | -8.2 | 6.90 | A | A | A |
| 34.94 | 0.26 | 1.9 | 6.17 | A | A | A |
| 31.4 | 2.6 | -8.4 | 10.30 | A | A | A |
| 35 | 1 | 2.0 | 6.76 | A | A | A |
| 30.2 | 1.66 | -11.9 | 8.23 | A | A | A |
| 32.9 | 3.8 | -4.1 | 13.07 | A | A | A |
| 30.6 | 1.9 | -10.8 | 8.72 | A | A | A |
| 33.8 | 2.17 | -1.5 | 8.87 | A | A | A |
| 35.75 | 1.15 | 4.2 | 6.92 | A | A | A |
| 35.6 | 3.37 | 3.8 | 11.27 | A | A | A |
| 35 | 0.3 | 2.0 | 6.18 | A | A | A |
| 33.2 | 2.8 | -3.2 | 10.42 | A | A | A |
| 35.2 | 3.1 | 2.6 | 10.73 | A | A | A |
| 33.5 | 3 | -2.3 | 10.85 | A | A | A |
| 30.56 | 1.36 | -10.9 | 7.57 | A | A | A |

Labcode

116

119

121

124

127

128

130

132

133

134

139

140

143

145

147

150

151

152

154

155

156

158

159

161

165

166

168

170

172

175

177

178

179

180

181

183

184

185

187

189

190

191

194

197

198

200

202

203

204

205

206

207

208

211

213

214

220

223

224

226

| Reported value [Bq/kg ] 32.8 | Reported uncertainty (k = 1) [Bq/kg ] 1 | Relative bias [%] -4.4 | P-Test [%] 6.84 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 33.9 | 1.6 | -1.2 | 7.73 | A | A | A |
| 32.387 | 1.908 | -5.6 | 8.50 | A | A | A |
| 32.3 | 0.7 | -5.8 | 6.49 | A | A | A |
| 33.1 | 2.4 | -3.5 | 9.49 | A | A | A |
| 32.2 | 2.2 | -6.1 | 9.17 | A | A | A |
| 34.4 | 3.1 | 0.3 | 10.89 | A | A | A |
| 31.9 | 2.95 | -7.0 | 11.09 | A | A | A |
| 31.7 | 1 | -7.6 | 6.89 | A | A | A |
| 35.4 | 2.72 | 3.2 | 9.82 | A | A | A |
| 34.8 | 3.2 | 1.5 | 11.05 | A | A | A |
| 33.63 | 0.84 | -2.0 | 6.61 | A | A | A |
| 37.002 | 2.7 | 7.9 | 9.53 | A | A | A |
| 31.107933 | 2.335043 | -9.3 | 9.69 | A | A | A |
| 32.3 | 2.6 | -5.8 | 10.11 | A | A | A |
| 17.52 | 0.85 | -48.9 | 7.81 | N | N | N |
| 36.6 | 2.2 | 6.7 | 8.58 | A | A | A |
| 38.79 | 1.51 | 13.1 | 7.26 | A | A | A |
| 36.32 | 1.75 | 5.9 | 7.79 | A | A | A |
| 35.23 | 0.7 | 2.7 | 6.44 | A | A | A |
| 34.4 | 1.1 | 0.3 | 6.91 | A | A | A |
| 35.4 | 2.1 | 3.2 | 8.52 | A | A | A |
| 33.6 | 1 | -2.0 | 6.81 | A | A | A |
| 32.52 | 0.79 | -5.2 | 6.59 | A | A | A |
| 35.4 | 3 | 3.2 | 10.45 | A | A | A |
| 34.7 | 4.2 | 1.2 | 13.56 | A | A | A |
| 34.35 | 3.19 | 0.1 | 11.12 | A | A | A |
| 35.6 | 1.8 | 3.8 | 7.94 | A | A | A |
| 33.5 | 3.5 | -2.3 | 12.11 | A | A | A |
| 31.6 | 1.8 | -7.9 | 8.36 | A | A | A |
| 30.8 | 1.9 | -10.2 | 8.69 | A | A | A |
| 38.5 | 2.32658 | 12.2 | 8.60 | A | A | A |
| 33.1 | 1.8 | -3.5 | 8.19 | A | A | A |
| 30.7 | 0.9 | -10.5 | 6.79 | A | A | A |
| 33 | 3.4 | -3.8 | 11.98 | A | A | A |
| 34.8 | 1.2 | 1.5 | 7.03 | A | A | A |
| 31.01 | 0.5 | -9.6 | 6.33 | A | A | A |
| 35.6 | 1.56 | 3.8 | 7.53 | A | A | A |
| 35.425 | 0.4268 | 3.3 | 6.24 | A | A | A |
| 33.2 | 1.7 | -3.2 | 7.98 | A | A | A |
| 30.9 | 0.9 | -9.9 | 6.78 | A | A | A |
| 34.1 | 1.1 | -0.6 | 6.92 | A | A | A |
| 35 | 1 | 2.0 | 6.76 | A | A | A |
| 33 | 3.5 | -3.8 | 12.25 | A | A | A |
| 32.72 | 3.436 | -4.6 | 12.16 | A | A | A |
| 32 | 1 | -6.7 | 6.87 | A | A | A |
| 32.1 | 2.3 | -6.4 | 9.42 | A | A | A |
| 36.3 | 4 | 5.8 | 12.61 | A | A | A |
| 31.99 | 1.13 | -6.7 | 7.07 | A | A | A |
| 31.9 | 1.5 | -7.0 | 7.72 | A | A | A |
| 37 | 3 | 7.9 | 10.16 | A | A | A |
| 32.7 | 0.73 | -4.7 | 6.52 | A | A | A |
| 32.5 | 2 | -5.2 | 8.68 | A | A | A |
| 32.8 | 2 | -4.4 | 8.64 | A | A | A |
| 33.1 | 5 | -3.5 | 16.30 | A | A | A |
| 35.3 | 1.4 | 2.9 | 7.29 | A | A | A |
| 33.3 | 2.4 | -2.9 | 9.46 | A | A | A |
| 39.2 | 4.07 | 14.3 | 12.05 | A | A | A |
| 36.5 | 3.8 | 6.4 | 12.08 | A | A | A |
| 33.6 | 1.82 | -2.0 | 8.17 | A | A | A |

Labcode

227

228

229

230

232

233

235

238

239

240

241

242

247

249

250

252

253

254

256

257

258

260

261

262

264

265

267

268

270

271

272

274

276

277

278

279

281

283

284

285

286

287

288

290

291

292

293

294

295

297

299

300

303

304

306

307

310

311

312

313

| Reported value [Bq/kg ] 37.533 | Reported uncertainty (k = 1) [Bq/kg ] 1.767 | Relative bias [%] 9.4 | P-Test [%] 7.72 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 36.1 | 5.3 | 5.2 | 15.91 | A | A | A |
| 30.6 | 1.7 | -10.8 | 8.27 | A | A | A |
| 34.83 | 0.67 | 1.6 | 6.42 | A | A | A |
| 35.2 | 1.26 | 2.6 | 7.09 | A | A | A |
| 35.1 | 0.8 | 2.3 | 6.53 | A | A | A |
| 33.2 | 3 | -3.2 | 10.91 | A | A | A |
| 35.4 | 1.75 | 3.2 | 7.87 | A | A | A |
| 36.32 | 1.9 | 5.9 | 8.05 | A | A | A |
| 32.4 | 1.7 | -5.5 | 8.06 | A | A | A |
| 31.7 | 1.8 | -7.6 | 8.35 | A | A | A |
| 33.07 | 1.84 | -3.6 | 8.27 | A | A | A |
| 31.8 | 1.1 | -7.3 | 7.03 | A | A | A |
| 34.8 | 2.6 | 1.5 | 9.66 | A | A | A |
| 34.85 | 2.47 | 1.6 | 9.37 | A | A | A |
| 35.8 | 1.4 | 4.4 | 7.26 | A | A | A |
| 27.7 | 0.9 | -19.2 | 6.93 | A | N | W |
| 30.55 | 0.55 | -10.9 | 6.38 | A | A | A |
| 34.1 | 2.9 | -0.6 | 10.48 | A | A | A |
| 34.3 | 1.3 | 0.0 | 7.20 | A | A | A |
| 32.38 | 1.7 | -5.6 | 8.07 | A | A | A |
| 35.4 | 1.7 | 3.2 | 7.78 | A | A | A |
| 33.8 | 1 | -1.5 | 6.80 | A | A | A |
| 31.1 | 1.7 | -9.3 | 8.21 | A | A | A |
| 35.4 | 1.5 | 3.2 | 7.45 | A | A | A |
| 34.25 | 0.34 | -0.1 | 6.20 | A | A | A |
| 34.8 | 1.8 | 1.5 | 8.01 | A | A | A |
| 32 | 1.6 | -6.7 | 7.90 | A | A | A |
| 25.7 | 1.3 | -25.1 | 7.94 | N | N | N |
| 32.6 | 0.9 | -5.0 | 6.72 | A | A | A |
| 27 | 2.7 | -21.3 | 11.73 | N | A | N |
| 27 | 0.9 | -21.3 | 6.97 | N | N | N |
| 33.8 | 3.8 | -1.5 | 12.80 | A | A | A |
| 34.4 | 0.54 | 0.3 | 6.32 | A | A | A |
| 33.21 | 1.66 | -3.2 | 7.90 | A | A | A |
| 34.5 | 1.6 | 0.6 | 7.68 | A | A | A |
| 40.8 | 1.6 | 19.0 | 7.27 | A | N | W |
| 35 | 4.2 | 2.0 | 13.47 | A | A | A |
| 31.82 | 0.96 | -7.2 | 6.83 | A | A | A |
| 35.1 | 2.22 | 2.3 | 8.80 | A | A | A |
| 32.7 | 3.3 | -4.7 | 11.80 | A | A | A |
| 32.68 | 3.3 | -4.7 | 11.81 | A | A | A |
| 32.8 | 1 | -4.4 | 6.84 | A | A | A |
| 32.9 | 1 | -4.1 | 6.84 | A | A | A |
| 32.48 | 0.69 | -5.3 | 6.48 | A | A | A |
| 33.9 | 1.51 | -1.2 | 7.57 | A | A | A |
| 33.52 | 1.75 | -2.3 | 8.05 | A | A | A |
| 31.91 | 3.15 | -7.0 | 11.62 | A | A | A |
| 31.34 | 2.56 | -8.6 | 10.21 | A | A | A |
| 33 | 2.1 | -3.8 | 8.83 | A | A | A |
| 35.2 | 2.1 | 2.6 | 8.55 | A | A | A |
| 36.5 | 3.3 | 6.4 | 10.92 | A | A | A |
| 36.16 | 2.2 | 5.4 | 8.63 | A | A | A |
| 33.8 | 1.8 | -1.5 | 8.11 | A | A | A |
| 26.3 | 1.4 | -23.3 | 8.11 | N | N | N |
| 33.5 | 1.6 | -2.3 | 7.77 | A | A | A |
| 32.8 | 0.66 | -4.4 | 6.44 | A | A | A |
| 34.6 | 3.3 | 0.9 | 11.33 | A | A | A |
| 32.7 | 1.8 | -4.7 | 8.23 | A | A | A |
| 31 | 1.33 | -9.6 | 7.48 | A | A | A |

Labcode

315

316

317

318

319

322

323

325

326

328

329

330

331

332

338

339

341

342

344

346

347

348

350

351

354

355

356

357

358

363

364

367

368

369

371

374

376

379

381

382

384

386

387

388

391

392

393

395

396

398

401

402

403

404

405

407

408

410

411

412

| Reported value [Bq/kg ] 34.6 | Reported uncertainty (k = 1) [Bq/kg ] 2.4 | Relative bias [%] 0.9 | P-Test [%] 9.25 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 34.6 | 1.4 | 0.9 | 7.34 | A | A | A |
| 34.4 | 2.2 | 0.3 | 8.85 | A | A | A |
| 35.7 | 3.15 | 4.1 | 10.74 | A | A | A |
| 35.2 | 3.9 | 2.6 | 12.66 | A | A | A |
| 38.2 | 2 | 11.4 | 8.06 | A | A | A |
| 33.2 | 1.2 | -3.2 | 7.11 | A | A | A |
| 32.13 | 2.42 | -6.3 | 9.71 | A | A | A |
| 33.9 | 2.1 | -1.2 | 8.71 | A | A | A |
| 32.3 | 0.83 | -5.8 | 6.64 | A | A | A |
| 34.87 | 0.56 | 1.7 | 6.33 | A | A | A |
| 24.54 | 1 | -28.4 | 7.35 | N | N | N |
| 37.7 | 3.5 | 9.9 | 11.12 | A | A | A |
| 34.1 | 2.2 | -0.6 | 8.89 | A | A | A |
| 33.6 | 1.7 | -2.0 | 7.94 | A | A | A |
| 35.86 | 0.692 | 4.6 | 6.42 | A | A | A |
| 34.9 | 1.3 | 1.8 | 7.17 | A | A | A |
| 36.129187 | 2.522848 | 5.3 | 9.29 | A | A | A |
| 31.5 | 1.4 | -8.2 | 7.57 | A | A | A |
| 33.34 | 1.75 | -2.8 | 8.06 | A | A | A |
| 35.4 | 1.2 | 3.2 | 7.00 | A | A | A |
| 34.5 | 1.1 | 0.6 | 6.90 | A | A | A |
| 33.26 | 0.43 | -3.0 | 6.26 | A | A | A |
| 35.1 | 1.1 | 2.3 | 6.88 | A | A | A |
| 32.76 | 2.6 | -4.5 | 10.02 | A | A | A |
| 34.3 | 0.811 | 0.0 | 6.56 | A | A | A |
| 36.9 | 1.6 | 7.6 | 7.50 | A | A | A |
| 4.7 | 1.7 | -86.3 | 36.68 | N | N | N |
| 33.01 | 2.83 | -3.8 | 10.53 | A | A | A |
| 36.3 | 3 | 5.8 | 10.29 | A | A | A |
| 56 | 3 | 63.3 | 8.14 | N | N | N |
| 35.5 | 1.8 | 3.5 | 7.95 | A | A | A |
| 37.72 | 2.16 | 10.0 | 8.38 | A | A | A |
| 34.1 | 3.1 | -0.6 | 10.96 | A | A | A |
| 31.261 | 1.85695 | -8.9 | 8.53 | A | A | A |
| 34.3 | 1.6 | 0.0 | 7.70 | A | A | A |
| 33.7 | 1 | -1.8 | 6.80 | A | A | A |
| 32.6 | 1.6 | -5.0 | 7.85 | A | A | A |
| 34.4 | 1.7 | 0.3 | 7.87 | A | A | A |
| 30.9 | 1.3 | -9.9 | 7.43 | A | A | A |
| 35.5 | 1.9 | 3.5 | 8.13 | A | A | A |
| 36.54 | 3.01 | 6.5 | 10.26 | A | A | A |
| 31.8 | 1.5 | -7.3 | 7.73 | A | A | A |
| 34.9 | 1.3 | 1.8 | 7.17 | A | A | A |
| 33.5 | 1.1 | -2.3 | 6.95 | A | A | A |
| 34 | 4 | -0.9 | 13.26 | A | A | A |
| 7.152 | 0.432 | -79.2 | 8.60 | N | N | N |
| 30.73 | 0.76 | -10.4 | 6.60 | A | A | A |
| 32.96 | 0.91 | -3.9 | 6.72 | A | A | A |
| 27.15 | 2.83 | -20.9 | 12.09 | N | A | N |
| 27.5 | 2.8 | -19.8 | 11.88 | A | A | A |
| 36.31 | 3.784 | 5.9 | 12.09 | A | A | A |
| 36.3 | 1.2 | 5.8 | 6.96 | A | A | A |
| 34 | 2 | -0.9 | 8.49 | A | A | A |
| 35.3 | 0.65 | 2.9 | 6.39 | A | A | A |
| 34.9 | 0.6 | 1.8 | 6.36 | A | A | A |
| 37.5 | 2.6 | 9.3 | 9.25 | A | A | A |
| 33.2 | 2 | -3.2 | 8.59 | A | A | A |
| 32 | 1 | -6.7 | 6.87 | A | A | A |
| 31.2 | 2 | -9.0 | 8.86 | A | A | A |

Labcode

414

417

419

420

421

423

425

426

428

429

433

434

435

437

439

440

441

442

443

444

445

446

448

450

452

453

454

455

459

460

462

463

464

465

466

469

471

472

475

477

478

479

482

483

484

487

491

495

496

498

499

500

501

506

508

509

514

516

522

523

| Reported value [Bq/kg ] 34.8 | Reported uncertainty (k = 1) [Bq/kg ] 2.2 | Relative bias [%] 1.5 | P-Test [%] 8.80 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 33.5 | 2 | -2.3 | 8.55 | A | A | A |
| 34.3 | 2.1 | 0.0 | 8.66 | A | A | A |
| 20.5 | 1.1 | -40.2 | 8.14 | N | N | N |
| 28.443 | 1.1 | -17.1 | 7.24 | A | A | A |
| 33.797968 | 1.627087 | -1.5 | 7.79 | A | A | A |
| 28.1 | 2.1 | -18.1 | 9.66 | A | A | A |
| 34.5 | 1.7 | 0.6 | 7.86 | A | A | A |
| 37 | 1.6 | 7.9 | 7.50 | A | A | A |
| 34.18 | 1.64 | -0.3 | 7.78 | A | A | A |
| 33.01 | 1.45 | -3.8 | 7.54 | A | A | A |
| 34.01 | 1.02 | -0.8 | 6.82 | A | A | A |
| 37.4 | 2.6 | 9.0 | 9.26 | A | A | A |

Labcode

525

527

528

529

530

532

533

535

536

537

539

544

555

Cs-134 determination of activity concentration in Sample 1

Table 12. Target values

| Parameter | Value |
| --- | --- |
| Activity concentration [ Bq/kg ] | 24.6 |
| Uncertainty [ Bq/kg ] | 1.0 |
| Maximum Acceptable Relative Bias [%] | 20 |

Accepted 76%	Not Accepted 4%	Warning 4%	Not Reported 16%

30

35

Figure 12. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

15

20

491

252

147

495

523

155

405

499

27

83

272

498

281

434

249

21

145

211

207

277

477

297

392

533

330

209

506

530

265

181

150

358

448

240

539

68

200

351

391

239

213

226

24

250

262

41

268

243

35

293

368

452

58

204

303

218

139

328

291

257

70

398

110

8

73

439

535

12

528

156

444

121

227

437

310

45

280

190

341

435

151

311

441

175

464

109

306

348

94

276

421

37

114

288

339

370

417

294

312

453

2

65

258

344

256

50

105

161

264

316

420

403

235

274

465

527

90

260

459

11

19

104

241

356

509

382

74

205

319

404

85

75

13

178

478

347

396

214

62

152

460

385

177

287

525

326

387

433

314

487

401

198

379

318

323

428

544

223

454

165

63

475

364

338

536

411

410

172

Laboratory code

Figure 13. Reported results and evaluation parameters for Cs-134. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 13. Evaluation results for Cs-134 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

Labcode

Reported value

[Bq/kg ]

Reported uncertainty

(k = 1) [Bq/kg ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

Activity concentration  [ Bq/kg ]

25

2	24.3	2.2	-1.2	9.92	A	A	A

8	23.25	1.39	-5.5	7.23	A	A	A

9	21.8	2.6	-11.4	12.60	A	A	A

| Reported value [Bq/kg ] 24.7 | Reported uncertainty (k = 1) [Bq/kg ] 1.1 | Relative bias [%] 0.4 | P-Test [%] 6.03 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 23.36 | 1.84 | -5.0 | 8.86 | A | A | A |
| 24.9 | 1.2 | 1.2 | 6.30 | A | A | A |
| 24.7 | 1.2 | 0.4 | 6.33 | A | A | A |
| 24.51 | 0.91 | -0.4 | 5.51 | A | A | A |
| 24.7 | 1.4 | 0.4 | 6.98 | A | A | A |
| 22.13 | 0.33 | -10.0 | 4.33 | A | A | A |
| 21 | 2 | -14.6 | 10.36 | A | A | A |
| 22.7 | 1.4 | -7.7 | 7.39 | A | A | A |
| 23.5 | 1 | -4.5 | 5.88 | A | A | A |
| 20.26 | 0.6 | -17.6 | 5.03 | A | N | W |
| 24.34 | 2.097 | -1.1 | 9.53 | A | A | A |
| 22.4 | 0.7 | -8.9 | 5.13 | A | A | A |
| 21.5 | 1.5 | -12.6 | 8.07 | A | A | A |
| 22.880698 | 1.982866 | -7.0 | 9.57 | A | A | A |
| 24.1 | 0.8 | -2.0 | 5.25 | A | A | A |
| 25.82 | 2.16 | 5.0 | 9.30 | A | A | A |
| 23.84 | 1.1 | -3.1 | 6.15 | A | A | A |
| 22.8 | 0.271 | -7.3 | 4.24 | A | A | A |
| 23.7 | 1.1 | -3.7 | 6.17 | A | A | A |
| 23.66 | 1.16 | -3.8 | 6.37 | A | A | A |
| 24.7 | 1.3 | 0.4 | 6.65 | A | A | A |
| 23 | 2 | -6.5 | 9.60 | A | A | A |
| 24.4 | 0.8 | -0.8 | 5.22 | A | A | A |
| 25.77 | 1.04 | 4.8 | 5.73 | A | A | A |
| 24.3 | 0.95 | -1.2 | 5.64 | A | A | A |
| 22.64 | 2.1 | -8.0 | 10.13 | A | A | A |
| 22.55 | 0.24 | -8.3 | 4.20 | A | A | A |
| 19.4 | 1.9 | -21.1 | 10.60 | N | A | N |
| 23 | 1 | -6.5 | 5.95 | A | A | A |
| 24.4 | 2.2 | -0.8 | 9.89 | A | A | A |
| 25.1 | 1.9 | 2.0 | 8.59 | A | A | A |
| 26.34 | 2.3 | 7.1 | 9.63 | A | A | A |
| 24.3 | 1.8 | -1.2 | 8.45 | A | A | A |
| 23.68 | 1.237 | -3.7 | 6.62 | A | A | A |
| 24.1 | 3 | -2.0 | 13.10 | A | A | A |
| 22.4 | 1 | -8.9 | 6.04 | A | A | A |
| 23.94718 | 1 | -2.6 | 5.83 | A | A | A |
| 23.2 | 1.2 | -5.7 | 6.58 | A | A | A |
| 20.9 | 1.05 | -15.0 | 6.46 | A | A | A |
| 24.9 | 4.42 | 1.2 | 18.21 | A | A | A |
| 23.3 | 0.7 | -5.3 | 5.05 | A | A | A |
| 24.8 | 2.49 | 0.8 | 10.83 | A | A | A |
| 24.88 | 2.87 | 1.1 | 12.23 | A | A | A |
| 21.83 | 0.77 | -11.3 | 5.38 | A | A | A |
| 21.2 | 0.8 | -13.8 | 5.55 | A | A | A |
| 23.885 | 0.62281 | -2.9 | 4.83 | A | A | A |
| 18.3 | 0.751 | -25.6 | 5.78 | N | N | N |
| 23.05 | 0.31 | -6.3 | 4.28 | A | A | A |
| 20.5 | 3 | -16.7 | 15.19 | A | A | A |
| 24.824 | 2.346 | 0.9 | 10.29 | A | A | A |
| 24.6 | 1.1 | 0.0 | 6.04 | A | A | A |
| 23.92 | 0.54 | -2.8 | 4.65 | A | A | A |
| 21.51 | 1.47 | -12.6 | 7.95 | A | A | A |
| 25.1 | 1.4 | 2.0 | 6.90 | A | A | A |
| 23.74 | 1.47 | -3.5 | 7.41 | A | A | A |
| 16.09 | 0.32 | -34.6 | 4.53 | N | N | N |
| 24.7 | 0.8 | 0.4 | 5.20 | A | A | A |
| 24.4 | 0.8 | -0.8 | 5.22 | A | A | A |
| 24.4 | 1.4 | -0.8 | 7.03 | A | A | A |

Labcode

11

12

13

15

17

19

20

21

24

25

27

28

31

34

35

37

38

40

41

42

45

48

49

50

51

52

53

55

57

58

60

62

63

65

66

67

68

69

70

71

72

73

74

75

78

79

80

81

82

83

85

90

94

95

97

100

102

104

105

107

| Reported value [Bq/kg ] 23.9 | Reported uncertainty (k = 1) [Bq/kg ] 0.2 | Relative bias [%] -2.9 | P-Test [%] 4.15 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 23.205703 | 1.150284 | -5.7 | 6.41 | A | A | A |
| 24.06 | 1.94 | -2.2 | 9.03 | A | A | A |
| 23.62 | 2.36 | -4.0 | 10.79 | A | A | A |
| 24.1 | 1.5 | -2.0 | 7.43 | A | A | A |
| 23.55 | 1.18 | -4.3 | 6.45 | A | A | A |
| 22.4 | 2.9 | -8.9 | 13.57 | A | A | A |
| 23.5 | 1.2 | -4.5 | 6.53 | A | A | A |
| 24.7 | 1.4 | 0.4 | 6.98 | A | A | A |
| 25.55198 | 2.211405 | 3.9 | 9.56 | A | A | A |
| 23.5 | 1.59 | -4.5 | 7.89 | A | A | A |
| 25.17 | 0.7 | 2.3 | 4.93 | A | A | A |
| 25.7 | 1.3 | 4.5 | 6.49 | A | A | A |
| 26.3 | 1.1 | 6.9 | 5.83 | A | A | A |
| 26 | 0.3 | 5.7 | 4.23 | A | A | A |
| 23.1 | 0.937 | -6.1 | 5.74 | A | A | A |
| 23.8 | 1.4 | -3.2 | 7.15 | A | A | A |
| 22.858359 | 1.124924 | -7.1 | 6.38 | A | A | A |
| 21.14 | 3.83 | -14.1 | 18.57 | A | A | A |
| 15.5265 | 0.0776 | -36.9 | 4.10 | N | N | N |
| 22.2 | 1.4 | -9.8 | 7.50 | A | A | A |
| 23.75 | 0.77 | -3.5 | 5.20 | A | A | A |
| 25.1 | 1.5 | 2.0 | 7.23 | A | A | A |
| 25 | 1.4 | 1.6 | 6.92 | A | A | A |
| 19.5 | 1.3 | -20.7 | 7.81 | N | N | N |
| 23.43 | 0.52 | -4.8 | 4.63 | A | A | A |
| 21.1 | 2.1 | -14.2 | 10.75 | A | A | A |
| 23.4 | 0.65 | -4.9 | 4.92 | A | A | A |
| 24.4 | 0.95 | -0.8 | 5.63 | A | A | A |
| 26.3 | 0.6 | 6.9 | 4.66 | A | A | A |
| 23 | 1.32 | -6.5 | 7.03 | A | A | A |
| 24.8 | 2.1 | 0.8 | 9.39 | A | A | A |
| 23.1 | 1.2 | -6.1 | 6.60 | A | A | A |
| 31 | 1.5 | 26.0 | 6.32 | N | N | N |
| 23.81 | 0.99 | -3.2 | 5.81 | A | A | A |
| 25.2 | 0.8 | 2.4 | 5.16 | A | A | A |
| 24.9 | 1.6 | 1.2 | 7.60 | A | A | A |
| 20.58 | 0.55 | -16.3 | 4.86 | A | N | W |
| 20.5 | 0.77 | -16.7 | 5.53 | A | N | W |
| 22.12 | 1.72 | -10.1 | 8.77 | A | A | A |
| 22.1 | 0.9 | -10.2 | 5.75 | A | A | A |
| 22.29 | 2.49 | -9.4 | 11.89 | A | A | A |
| 22.27 | 0.22 | -9.5 | 4.18 | A | A | A |
| 24.3 | 1.8 | -1.2 | 8.45 | A | A | A |
| 22.02 | 4.4 | -10.5 | 20.39 | A | N | W |
| 23.7 | 0.61 | -3.7 | 4.81 | A | A | A |
| 24.6 | 2 | 0.0 | 9.09 | A | A | A |
| 20.21 | 0.92 | -17.9 | 6.10 | A | N | W |
| 24.97 | 2 | 1.5 | 8.98 | A | A | A |
| 25.7 | 2.1 | 4.5 | 9.13 | A | A | A |
| 22.4 | 0.8 | -8.9 | 5.41 | A | A | A |
| 24.62 | 0.22 | 0.1 | 4.16 | A | A | A |
| 21.3 | 1.7 | -13.4 | 8.96 | A | A | A |
| 23 | 0.5 | -6.5 | 4.61 | A | A | A |
| 24.8 | 1.36 | 0.8 | 6.83 | A | A | A |
| 22.7 | 2.5 | -7.7 | 11.74 | A | A | A |
| 21.4 | 1.4 | -13.0 | 7.70 | A | A | A |
| 24.1 | 1.11 | -2.0 | 6.14 | A | A | A |
| 22 | 1.3 | -10.6 | 7.17 | A | A | A |
| 21.29 | 0.61 | -13.5 | 4.97 | A | N | W |

Labcode

109

110

111

113

114

116

119

121

124

127

128

130

132

133

134

139

140

143

145

147

150

151

152

154

155

156

158

159

161

165

166

168

170

172

175

177

178

179

180

181

183

184

185

187

189

190

191

194

197

198

200

202

203

204

205

206

207

208

209

211

| Reported value [Bq/kg ] 22.6 | Reported uncertainty (k = 1) [Bq/kg ] 1.87 | Relative bias [%] -8.1 | P-Test [%] 9.22 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 25 | 0.1 | 1.6 | 4.08 | A | A | A |
| 24.4 | 0.66 | -0.8 | 4.88 | A | A | A |
| 25.1 | 3.4 | 2.0 | 14.14 | A | A | A |
| 23.01286 | 0.71458 | -6.4 | 5.12 | A | A | A |
| 24.5 | 1.6 | -0.4 | 7.69 | A | A | A |
| 26 | 2.1 | 5.7 | 9.04 | A | A | A |
| 22.9 | 2.1 | -6.9 | 10.03 | A | A | A |
| 22.63 | 0.92 | -8.0 | 5.75 | A | A | A |
| 23.5 | 0.8 | -4.5 | 5.30 | A | A | A |
| 22.8 | 0.7 | -7.3 | 5.09 | A | A | A |
| 22.404 | 1.2855 | -8.9 | 7.03 | A | A | A |
| 23.5 | 0.4 | -4.5 | 4.41 | A | A | A |
| 22.5 | 0.9 | -8.5 | 5.70 | A | A | A |
| 25.2 | 2 | 2.4 | 8.92 | A | A | A |
| 24.5 | 2 | -0.4 | 9.12 | A | A | A |
| 24.5 | 1.38 | -0.4 | 6.95 | A | A | A |
| 22.5 | 0.8 | -8.5 | 5.40 | A | A | A |
| 22.29 | 1.52 | -9.4 | 7.94 | A | A | A |
| 24.7 | 1.6 | 0.4 | 7.65 | A | A | A |
| 25.21 | 0.54 | 2.5 | 4.59 | A | A | A |
| 22.841 | 0.409 | -7.2 | 4.44 | A | A | A |
| 24.396 | 2.047 | -0.8 | 9.32 | A | A | A |
| 20.887574 | 1.044196 | -15.1 | 6.44 | A | A | A |
| 22.7 | 1.7 | -7.7 | 8.52 | A | A | A |
| 13.92 | 0.67 | -43.4 | 6.30 | N | N | N |
| 27 | 1.6 | 9.8 | 7.19 | A | A | A |
| 21.4 | 1.3 | -13.0 | 7.31 | A | A | A |
| 24.39 | 1.18 | -0.8 | 6.32 | A | A | A |
| 23.17 | 0.39 | -5.8 | 4.40 | A | A | A |
| 24.3 | 0.7 | -1.2 | 4.98 | A | A | A |
| 24.6 | 1.3 | 0.0 | 6.67 | A | A | A |
| 26.4 | 0.7 | 7.3 | 4.85 | A | A | A |
| 22.77 | 0.56 | -7.4 | 4.75 | A | A | A |
| 24.4 | 2 | -0.8 | 9.15 | A | A | A |
| 22.1 | 2.7 | -10.2 | 12.88 | A | A | A |
| 24.26 | 2.3 | -1.4 | 10.32 | A | A | A |
| 22.8 | 1.1 | -7.3 | 6.31 | A | A | A |
| 24.6 | 2.5 | 0.0 | 10.95 | A | A | A |
| 24.7 | 1.5 | 0.4 | 7.31 | A | A | A |
| 20.5 | 1.3 | -16.7 | 7.53 | A | A | A |
| 24.5 | 2.038535 | -0.4 | 9.26 | A | A | A |
| 24 | 1.3 | -2.4 | 6.77 | A | A | A |
| 21.4 | 0.6 | -13.0 | 4.94 | A | N | W |
| 21.7 | 2.3 | -11.8 | 11.35 | A | A | A |
| 24.4 | 0.8 | -0.8 | 5.22 | A | A | A |
| 23.68 | 1.15 | -3.7 | 6.33 | A | A | A |
| 20.64 | 0.34 | -16.1 | 4.39 | A | N | W |
| 25.9 | 0.92 | 5.3 | 5.40 | A | A | A |
| 26.437 | 0.2863 | 7.5 | 4.21 | A | A | A |
| 23.9 | 1.2 | -2.9 | 6.46 | A | A | A |
| 22.4 | 0.5 | -8.9 | 4.64 | A | A | A |
| 25.2 | 0.7 | 2.4 | 4.92 | A | A | A |
| 24.1 | 0.57 | -2.0 | 4.70 | A | A | A |
| 23 | 5 | -6.5 | 22.12 | A | N | W |
| 24.5 | 2.5 | -0.4 | 10.98 | A | A | A |
| 23.12 | 2.428 | -6.0 | 11.26 | A | A | A |
| 26.3 | 0.7 | 6.9 | 4.86 | A | A | A |
| 22.9 | 1.7 | -6.9 | 8.46 | A | A | A |
| 24.2 | 2.2 | -1.6 | 9.96 | A | A | A |

Labcode

213

214

216

217

218

220

223

224

226

227

228

229

230

232

233

235

238

239

240

241

242

243

247

249

250

252

253

254

256

257

258

260

261

262

264

265

267

268

270

271

272

274

276

277

278

279

280

281

283

284

285

286

287

288

289

290

291

292

293

294

| Reported value [Bq/kg ] 23.17 | Reported uncertainty (k = 1) [Bq/kg ] 0.8 | Relative bias [%] -5.8 | P-Test [%] 5.33 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 24.2 | 0.959 | -1.6 | 5.68 | A | A | A |
| 21.6 | 1.1 | -12.2 | 6.52 | A | A | A |
| 24 | 2 | -2.4 | 9.27 | A | A | A |
| 23.3 | 0.71 | -5.3 | 5.08 | A | A | A |
| 24.1 | 1.9 | -2.0 | 8.87 | A | A | A |
| 23 | 1 | -6.5 | 5.95 | A | A | A |
| 23.2 | 1.7 | -5.7 | 8.38 | A | A | A |
| 23.9 | 3.6 | -2.9 | 15.60 | A | A | A |
| 24.9 | 0.9 | 1.2 | 5.44 | A | A | A |
| 23.6 | 1.7 | -4.1 | 8.27 | A | A | A |
| 23.8 | 0.635 | -3.2 | 4.86 | A | A | A |
| 24.2 | 2.6 | -1.6 | 11.49 | A | A | A |
| 22.3 | 1.15 | -9.3 | 6.57 | A | A | A |
| 25.58 | 0.77 | 4.0 | 5.06 | A | A | A |
| 24.8 | 1.152 | 0.8 | 6.17 | A | A | A |
| 24.4 | 3.4 | -0.8 | 14.52 | A | A | A |
| 24.2 | 1.4 | -1.6 | 7.07 | A | A | A |
| 25.86 | 0.56 | 5.1 | 4.61 | A | A | A |
| 24.8 | 0.916 | 0.8 | 5.49 | A | A | A |
| 23.7 | 0.5 | -3.7 | 4.58 | A | A | A |
| 25.9 | 2.2 | 5.3 | 9.42 | A | A | A |
| 24.3 | 1.22 | -1.2 | 6.46 | A | A | A |
| 25.21 | 1.2 | 2.5 | 6.26 | A | A | A |
| 23.1 | 1.8 | -6.1 | 8.79 | A | A | A |
| 21.89 | 1.36 | -11.0 | 7.42 | A | A | A |
| 22.2 | 0.7 | -9.8 | 5.14 | A | A | A |
| 24.8 | 1.9 | 0.8 | 8.67 | A | A | A |
| 26.72 | 1.92 | 8.6 | 8.26 | A | A | A |
| 24.1 | 0.6 | -2.0 | 4.77 | A | A | A |
| 23.7 | 0.8 | -3.7 | 5.28 | A | A | A |
| 25.14 | 0.25 | 2.2 | 4.18 | A | A | A |
| 23.9 | 1.99 | -2.9 | 9.27 | A | A | A |
| 24.3 | 1.9 | -1.2 | 8.81 | A | A | A |
| 22.81 | 0.62 | -7.3 | 4.89 | A | A | A |
| 24.92 | 1.78 | 1.3 | 8.22 | A | A | A |
| 23.9 | 1 | -2.9 | 5.83 | A | A | A |
| 19.6 | 1 | -20.3 | 6.52 | N | N | N |
| 22.4 | 1.3 | -8.9 | 7.09 | A | A | A |
| 25.2 | 0.9 | 2.4 | 5.41 | A | A | A |
| 23.22 | 0.21 | -5.6 | 4.16 | A | A | A |
| 24.7 | 1.6 | 0.4 | 7.65 | A | A | A |
| 22.9 | 1.2 | -6.9 | 6.63 | A | A | A |
| 22.2 | 1.1 | -9.8 | 6.41 | A | A | A |
| 25.9 | 0.7 | 5.3 | 4.88 | A | A | A |
| 26.5 | 2.6 | 7.7 | 10.62 | A | A | A |
| 22 | 0.8 | -10.6 | 5.45 | A | A | A |
| 22.9 | 2.2 | -6.9 | 10.43 | A | A | A |
| 24.1 | 0.21 | -2.0 | 4.16 | A | A | A |
| 24.1 | 2.5 | -2.0 | 11.14 | A | A | A |
| 23.25 | 1.15 | -5.5 | 6.40 | A | A | A |
| 24.4 | 1.3 | -0.8 | 6.70 | A | A | A |
| 27.3 | 1.1 | 11.0 | 5.72 | A | A | A |
| 25.8 | 2.9 | 4.9 | 11.95 | A | A | A |
| 23.8 | 0.79 | -3.2 | 5.25 | A | A | A |
| 24.73 | 1.44 | 0.5 | 7.10 | A | A | A |
| 26.1 | 1.1 | 6.1 | 5.86 | A | A | A |
| 22.1 | 2.4 | -10.2 | 11.60 | A | A | A |
| 25.15 | 1.19 | 2.2 | 6.24 | A | A | A |
| 22.78 | 2.31 | -7.4 | 10.92 | A | A | A |

Labcode

295

296

297

299

300

302

303

304

306

307

310

311

312

313

314

315

316

317

318

319

322

323

325

326

328

330

331

332

338

339

341

342

343

344

346

347

348

350

351

354

355

356

357

358

363

364

367

368

369

370

371

374

376

379

381

382

383

384

385

386

| Reported value [Bq/kg ] 25.5 | Reported uncertainty (k = 1) [Bq/kg ] 0.5 | Relative bias [%] 3.7 | P-Test [%] 4.51 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 25.6 | 0.5 | 4.1 | 4.51 | A | A | A |
| 22.48 | 0.6 | -8.6 | 4.86 | A | A | A |
| 21.7 | 0.97 | -11.8 | 6.04 | A | A | A |
| 25.4 | 1.65 | 3.2 | 7.66 | A | A | A |
| 23.41 | 2.3 | -4.8 | 10.63 | A | A | A |
| 24.99 | 1.2 | 1.6 | 6.29 | A | A | A |
| 23.2 | 2 | -5.7 | 9.53 | A | A | A |
| 25.7 | 1.6 | 4.5 | 7.44 | A | A | A |
| 24.1 | 1.5 | -2.0 | 7.43 | A | A | A |
| 24.48 | 1.48 | -0.5 | 7.29 | A | A | A |
| 24.8 | 1.3 | 0.8 | 6.63 | A | A | A |
| 19.9 | 1.1 | -19.1 | 6.86 | A | N | W |
| 21.9 | 1.2 | -11.0 | 6.82 | A | A | A |
| 23.45 | 0.46 | -4.7 | 4.51 | A | A | A |
| 29.9 | 2.4 | 21.5 | 9.00 | N | A | N |
| 27.5 | 1.2 | 11.8 | 5.96 | A | A | A |
| 22.73 | 0.99 | -7.6 | 5.96 | A | A | A |
| 24.8 | 1.4 | 0.8 | 6.96 | A | A | A |
| 24.1 | 1.1 | -2.0 | 6.11 | A | A | A |
| 26.5 | 1.9 | 7.7 | 8.24 | A | A | A |
| 24.4 | 2.15 | -0.8 | 9.70 | A | A | A |
| 24 | 2.6 | -2.4 | 11.57 | A | A | A |
| 23.7 | 1 | -3.7 | 5.86 | A | A | A |
| 23.32 | 1.5 | -5.2 | 7.61 | A | A | A |
| 25.9 | 1.7 | 5.3 | 7.72 | A | A | A |
| 20.8 | 0.278 | -15.4 | 4.28 | A | N | W |
| 25.52 | 0.55 | 3.7 | 4.60 | A | A | A |
| 20.76 | 0.56 | -15.6 | 4.88 | A | N | W |
| 23.7 | 2 | -3.7 | 9.37 | A | A | A |
| 23.5 | 1.4 | -4.5 | 7.21 | A | A | A |
| 23.3 | 1.1 | -5.3 | 6.23 | A | A | A |
| 20.398 | 0.534 | -17.1 | 4.84 | A | N | W |
| 23.8 | 1.2 | -3.2 | 6.48 | A | A | A |
| 25.9 | 1.3 | 5.3 | 6.46 | A | A | A |
| 23.48 | 0.99 | -4.6 | 5.86 | A | A | A |
| 25.7 | 1.4 | 4.5 | 6.80 | A | A | A |
| 23.1 | 0.8 | -6.1 | 5.34 | A | A | A |
| 22.27 | 0.22 | -9.5 | 4.18 | A | A | A |
| 24.7 | 0.7 | 0.4 | 4.96 | A | A | A |
| 22.95 | 1.46 | -6.7 | 7.55 | A | A | A |
| 24.2 | 0.479 | -1.6 | 4.52 | A | A | A |
| 26.2 | 1 | 6.5 | 5.58 | A | A | A |
| 11.7 | 1.6 | -52.4 | 14.27 | N | N | N |
| 24.61 | 1.69 | 0.0 | 7.98 | A | A | A |
| 25.1 | 2 | 2.0 | 8.95 | A | A | A |
| 32 | 1 | 30.1 | 5.13 | N | N | N |
| 28.8 | 1.4 | 17.1 | 6.34 | A | N | W |
| 23.87 | 0.88 | -3.0 | 5.49 | A | A | A |
| 24.5 | 1.7 | -0.4 | 8.04 | A | A | A |
| 24.428 | 0.822 | -0.7 | 5.28 | A | A | A |
| 22.6 | 1.2 | -8.1 | 6.69 | A | A | A |
| 22.9 | 0.8 | -6.9 | 5.36 | A | A | A |
| 23.2 | 1.6 | -5.7 | 8.01 | A | A | A |
| 26.4 | 1.3 | 7.3 | 6.39 | A | A | A |
| 21.5 | 1.3 | -12.6 | 7.29 | A | A | A |
| 24.9 | 1.3 | 1.2 | 6.62 | A | A | A |
| 25.03 | 2.02 | 1.8 | 9.04 | A | A | A |
| 24.5 | 0.8 | -0.4 | 5.21 | A | A | A |
| 23.8 | 1.2 | -3.2 | 6.48 | A | A | A |

Labcode

387

388

391

392

393

395

396

398

401

402

403

404

405

407

408

410

411

412

414

417

419

420

421

425

426

428

429

433

434

435

437

439

440

441

443

444

445

446

448

450

452

453

454

455

459

460

462

463

464

465

466

469

471

472

475

477

478

479

482

483

| Reported value [Bq/kg ] 23.9 | Reported uncertainty (k = 1) [Bq/kg ] 0.7 | Relative bias [%] -2.9 | P-Test [%] 5.01 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 25.6 | 2.8 | 4.1 | 11.67 | A | A | A |
| 23.3 | 1.1 | -5.3 | 6.23 | A | A | A |
| 6.898 | 0.455 | -72.0 | 7.75 | N | N | N |
| 23.15 | 1.3 | -5.9 | 6.93 | A | A | A |
| 17.25 | 0.62 | -29.9 | 5.43 | N | N | N |
| 20.53 | 0.27 | -16.5 | 4.27 | A | N | W |
| 20.56 | 1.73 | -16.4 | 9.34 | A | A | A |
| 20 | 2 | -18.7 | 10.79 | A | A | A |
| 24.12 | 1.39 | -2.0 | 7.05 | A | A | A |
| 20.7 | 0.8 | -15.8 | 5.61 | A | N | W |
| 22 | 1 | -10.6 | 6.10 | A | A | A |
| 15.36 | 0.074 | -37.6 | 4.09 | N | N | N |
| 24.85 | 0.74 | 1.0 | 5.04 | A | A | A |
| 24.7 | 0.4 | 0.4 | 4.38 | A | A | A |
| 25.5 | 1.3 | 3.7 | 6.52 | A | A | A |
| 24.9 | 1.6 | 1.2 | 7.60 | A | A | A |
| 23 | 0.6 | -6.5 | 4.83 | A | A | A |
| 18.8 | 0.7 | -23.6 | 5.51 | N | N | N |
| 25.2 | 1.6 | 2.4 | 7.54 | A | A | A |
| 24.5 | 1.6 | -0.4 | 7.69 | A | A | A |
| 23.4 | 1 | -4.9 | 5.90 | A | A | A |
| 19.9 | 1.1 | -19.1 | 6.86 | A | N | W |
| 22.028 | 0.54 | -10.5 | 4.75 | A | A | A |
| 24.887096 | 1.528274 | 1.2 | 7.36 | A | A | A |
| 21.8 | 1.1 | -11.4 | 6.48 | A | A | A |
| 23.3 | 1.4 | -5.3 | 7.25 | A | A | A |
| 27 | 1.1 | 9.8 | 5.76 | A | A | A |
| 24.73 | 1.16 | 0.5 | 6.21 | A | A | A |
| 22.3 | 0.98 | -9.3 | 5.99 | A | A | A |
| 30.02 | 3.59 | 22.0 | 12.63 | N | A | N |
| 25.92 | 0.78 | 5.4 | 5.06 | A | A | A |
| 24.7 | 1.3 | 0.4 | 6.65 | A | A | A |

Labcode

484

487

490

491

492

495

496

498

499

500

501

506

507

508

509

514

516

522

523

525

527

528

529

530

532

533

535

536

537

539

540

544

553

Cs-137 determination of activity concentration in Sample 1

Table 14. Target values

| Parameter | Value |
| --- | --- |
| Activity concentration [ Bq/kg ] | 17.1 |
| Uncertainty [ Bq/kg ] | 0.7 |
| Maximum Acceptable Relative Bias [%] | 20 |

Accepted 79%	Not Accepted 3%	Warning 3%	Not Reported 15%

22

24

Figure 14. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

10

12

14

491

252

254

9

520

226

83

507

499

128

79

535

351

391

147

370

27

61

452

140

384

194

355

102

363

498

371

13

250

156

35

500

41

73

168

303

425

110

111

121

213

285

381

170

241

328

100

113

42

190

238

304

386

484

116

31

189

272

312

403

249

175

496

49

154

166

208

265

290

398

528

433

8

45

15

52

203

261

286

322

344

392

465

69

404

185

382

105

150

270

307

356

421

469

318

242

509

50

114

253

283

414

478

395

495

224

471

523

530

143

80

396

82

216

239

339

450

487

119

342

202

11

72

152

315

420

63

74

271

294

319

411

506

459

132

205

313

537

165

369

127

37

57

78

139

223

410

343

401

97

429

533

314

287

434

145

446

463

311

58

289

472

544

197

85

419

323

440

358

462

Laboratory code

Figure 15. Reported results and evaluation parameters for Cs-137. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 15. Evaluation results for Cs-137 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

Labcode

Reported value

[Bq/kg ]

Reported uncertainty

(k = 1) [Bq/kg ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

Activity concentration  [ Bq/kg ]

16

18

20

2	15.9	1.2	-7.0	8.59	A	A	A

8	17.06	1.13	-0.2	7.79	A	A	A

9	15	1.8	-12.3	12.68	A	A	A

| Reported value [Bq/kg ] 17.6 | Reported uncertainty (k = 1) [Bq/kg ] 0.81 | Relative bias [%] 2.9 | P-Test [%] 6.16 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 16.96 | 1.13 | -0.8 | 7.82 | A | A | A |
| 16.4 | 0.7 | -4.1 | 5.91 | A | A | A |
| 17.1 | 0.7 | 0.0 | 5.79 | A | A | A |
| 17.6 | 1.06 | 2.9 | 7.28 | A | A | A |
| 16.7 | 1 | -2.3 | 7.25 | A | A | A |
| 16.9 | 0.25 | -1.2 | 4.35 | A | A | A |
| 17 | 2 | -0.6 | 12.46 | A | A | A |
| 17.3 | 1.08 | 1.2 | 7.47 | A | A | A |
| 16.5 | 1 | -3.5 | 7.31 | A | A | A |
| 16.13 | 0.5 | -5.7 | 5.13 | A | A | A |
| 17.57 | 1.521 | 2.8 | 9.58 | A | A | A |
| 16.9 | 0.55 | -1.2 | 5.23 | A | A | A |
| 15.68 | 0.35 | -8.3 | 4.66 | A | A | A |
| 15.6 | 1.3 | -8.8 | 9.28 | A | A | A |
| 16.450809 | 1.706696 | -3.8 | 11.15 | A | A | A |
| 18 | 0.8 | 5.3 | 6.04 | A | A | A |
| 17.92 | 1.42 | 4.8 | 8.92 | A | A | A |
| 17.1 | 0.78 | 0.0 | 6.13 | A | A | A |
| 16.5 | 0.289 | -3.5 | 4.45 | A | A | A |
| 16.8 | 0.6 | -1.8 | 5.43 | A | A | A |
| 17.07 | 0.7 | -0.2 | 5.79 | A | A | A |
| 17.2 | 0.9 | 0.6 | 6.64 | A | A | A |
| 17 | 2 | -0.6 | 12.46 | A | A | A |
| 17.3 | 0.7 | 1.2 | 5.76 | A | A | A |
| 18.8 | 0.91 | 9.9 | 6.34 | A | A | A |
| 17.1 | 1.02 | 0.0 | 7.23 | A | A | A |
| 18.25 | 1.58 | 6.7 | 9.58 | A | A | A |
| 16.16 | 0.29 | -5.5 | 4.47 | A | A | A |
| 18.1 | 1.8 | 5.8 | 10.75 | A | A | A |
| 19 | 1 | 11.1 | 6.67 | A | A | A |
| 16.4 | 1.1 | -4.1 | 7.86 | A | A | A |
| 16.17 | 0.65 | -5.4 | 5.74 | A | A | A |
| 17.5 | 1.5 | 2.3 | 9.50 | A | A | A |
| 17.64 | 0.91 | 3.2 | 6.59 | A | A | A |
| 18.1 | 1.5 | 5.8 | 9.24 | A | A | A |
| 16.85 | 0.8885 | -1.5 | 6.68 | A | A | A |
| 16.3 | 1.9 | -4.7 | 12.35 | A | A | A |
| 16.5 | 0.7 | -3.5 | 5.90 | A | A | A |
| 17.10323 | 0.7 | 0.0 | 5.79 | A | A | A |
| 16.2 | 0.8 | -5.3 | 6.41 | A | A | A |
| 16.65 | 0.83 | -2.6 | 6.45 | A | A | A |
| 17.6 | 0.88 | 2.9 | 6.46 | A | A | A |
| 16.5 | 1.1 | -3.5 | 7.82 | A | A | A |
| 17.7 | 1.78 | 3.5 | 10.86 | A | A | A |
| 16.23 | 3.2 | -5.1 | 20.14 | A | N | W |
| 18.1 | 0.64 | 5.8 | 5.41 | A | A | A |
| 15.9 | 0.3 | -7.0 | 4.51 | A | A | A |
| 17.443 | 0.57147 | 2.0 | 5.24 | A | A | A |
| 15.067 | 0.924 | -11.9 | 7.37 | A | A | A |
| 17.5 | 0.34 | 2.3 | 4.53 | A | A | A |
| 15.6 | 2.3 | -8.8 | 15.30 | A | A | A |
| 9.33 | 0.82 | -45.4 | 9.70 | N | N | N |
| 20.099 | 2.014 | 17.5 | 10.82 | A | A | A |
| 17.4 | 0.8 | 1.8 | 6.16 | A | A | A |
| 16.45 | 0.71 | -3.8 | 5.95 | A | A | A |
| 14.45 | 1.13 | -15.5 | 8.83 | A | A | A |
| 18.3 | 0.9 | 7.0 | 6.40 | A | A | A |
| 16.73 | 1.14 | -2.2 | 7.95 | A | A | A |
| 16.29 | 0.51 | -4.7 | 5.15 | A | A | A |

Labcode

11

12

13

15

17

19

20

21

24

25

27

28

31

32

34

35

37

38

40

41

42

45

48

49

50

51

52

53

55

57

58

60

61

62

63

65

66

67

68

69

70

71

72

73

74

75

78

79

80

81

82

83

84

85

90

94

95

97

100

102

| Reported value [Bq/kg ] 17.3 | Reported uncertainty (k = 1) [Bq/kg ] 0.7 | Relative bias [%] 1.2 | P-Test [%] 5.76 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 17.2 | 0.7 | 0.6 | 5.77 | A | A | A |
| 16.9 | 0.9 | -1.2 | 6.72 | A | A | A |
| 17 | 0.1 | -0.6 | 4.14 | A | A | A |
| 16.527571 | 0.849356 | -3.4 | 6.57 | A | A | A |
| 16.56 | 1 | -3.2 | 7.30 | A | A | A |
| 16.79 | 1.34 | -1.8 | 8.97 | A | A | A |
| 17.3 | 1.1 | 1.2 | 7.56 | A | A | A |
| 16.86 | 0.83 | -1.4 | 6.40 | A | A | A |
| 17.53 | 1.82 | 2.5 | 11.16 | A | A | A |
| 16.6 | 0.9 | -2.9 | 6.79 | A | A | A |
| 17.2 | 0.9 | 0.6 | 6.64 | A | A | A |
| 17.908696 | 1.920349 | 4.7 | 11.48 | A | A | A |
| 15.88 | 1.02 | -7.1 | 7.62 | A | A | A |
| 17.64 | 0.66 | 3.2 | 5.55 | A | A | A |
| 17.8 | 0.8 | 4.1 | 6.08 | A | A | A |
| 18.1 | 1 | 5.8 | 6.88 | A | A | A |
| 17.6 | 0.2 | 2.9 | 4.25 | A | A | A |
| 18.1 | 1.1 | 5.8 | 7.33 | A | A | A |
| 16.2 | 1.1 | -5.3 | 7.93 | A | A | A |
| 17.424213 | 1.323837 | 1.9 | 8.63 | A | A | A |
| 18.5 | 2.95 | 8.2 | 16.46 | A | A | A |
| 16.0824 | 0.8422 | -6.0 | 6.65 | A | A | A |
| 17.2 | 0.9 | 0.6 | 6.64 | A | A | A |
| 16.98 | 1.36 | -0.7 | 8.99 | A | A | A |
| 17.6 | 1.1 | 2.9 | 7.47 | A | A | A |
| 17 | 0.9 | -0.6 | 6.69 | A | A | A |
| 15.9 | 1.1 | -7.0 | 8.04 | A | A | A |
| 16.43 | 0.4 | -3.9 | 4.76 | A | A | A |
| 16.5 | 2.1 | -3.5 | 13.37 | A | A | A |
| 17.9 | 0.8 | 4.7 | 6.06 | A | A | A |
| 17 | 0.7 | -0.6 | 5.81 | A | A | A |
| 17.88 | 1 | 4.6 | 6.93 | A | A | A |
| 17 | 0.44 | -0.6 | 4.84 | A | A | A |
| 16.5 | 1.3 | -3.5 | 8.88 | A | A | A |
| 16.7 | 0.8 | -2.3 | 6.30 | A | A | A |
| 17.5 | 0.5 | 2.3 | 4.99 | A | A | A |
| 16.96 | 0.71 | -0.8 | 5.86 | A | A | A |
| 17.8 | 0.5 | 4.1 | 4.96 | A | A | A |
| 16.8 | 1.6 | -1.8 | 10.37 | A | A | A |
| 17.35 | 0.56 | 1.5 | 5.21 | A | A | A |
| 17.86 | 1.14 | 4.4 | 7.58 | A | A | A |
| 16.11 | 1.31 | -5.8 | 9.10 | A | A | A |
| 19.6 | 0.7 | 14.6 | 5.43 | A | N | W |
| 15.85 | 1.76 | -7.3 | 11.83 | A | A | A |
| 17.17 | 0.29 | 0.4 | 4.43 | A | A | A |
| 18.1 | 1.5 | 5.8 | 9.24 | A | A | A |
| 16.9 | 3.4 | -1.2 | 20.53 | A | N | W |
| 16.8 | 0.46 | -1.8 | 4.92 | A | A | A |
| 17.1 | 1.1 | 0.0 | 7.62 | A | A | A |
| 16.21 | 0.8 | -5.2 | 6.41 | A | A | A |
| 19.62 | 2.35 | 14.7 | 12.66 | A | A | A |
| 18.3 | 1.6 | 7.0 | 9.65 | A | A | A |
| 16.1 | 0.5 | -5.8 | 5.14 | A | A | A |
| 17.57 | 0.23 | 2.8 | 4.30 | A | A | A |
| 17.1 | 1.4 | 0.0 | 9.15 | A | A | A |
| 16.6 | 0.5 | -2.9 | 5.08 | A | A | A |
| 17.8 | 0.989 | 4.1 | 6.90 | A | A | A |
| 16.8 | 2 | -1.8 | 12.59 | A | A | A |
| 17 | 1.1 | -0.6 | 7.66 | A | A | A |

Labcode

104

105

107

109

110

111

113

114

116

119

121

124

127

128

130

132

133

134

139

140

143

145

147

150

151

152

154

155

156

158

159

161

165

166

168

170

172

175

177

178

179

180

181

183

184

185

187

189

190

191

194

197

198

200

202

203

204

205

206

207

| Reported value [Bq/kg ] 17 | Reported uncertainty (k = 1) [Bq/kg ] 1.12 | Relative bias [%] -0.6 | P-Test [%] 7.76 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 17 | 1.3 | -0.6 | 8.67 | A | A | A |
| 18.49 | 0.61 | 8.1 | 5.26 | A | A | A |
| 16.6 | 1.1 | -2.9 | 7.79 | A | A | A |
| 19 | 0.1 | 11.1 | 4.13 | A | N | W |
| 17.5 | 0.55 | 2.3 | 5.16 | A | A | A |
| 18.5 | 4 | 8.2 | 22.01 | A | N | W |
| 16.28946 | 0.99086 | -4.7 | 7.33 | A | A | A |
| 17.5 | 1.2 | 2.3 | 7.99 | A | A | A |
| 18.1 | 1.7 | 5.8 | 10.25 | A | A | A |
| 17.4 | 1.6 | 1.8 | 10.07 | A | A | A |
| 15.56 | 0.95 | -9.0 | 7.35 | A | A | A |
| 16.7 | 0.5 | -2.3 | 5.07 | A | A | A |
| 16 | 1 | -6.4 | 7.47 | A | A | A |
| 16.304 | 0.95236 | -4.7 | 7.13 | A | A | A |
| 16.5 | 0.3 | -3.5 | 4.48 | A | A | A |
| 17.3 | 0.8 | 1.2 | 6.18 | A | A | A |
| 17.8 | 1.1 | 4.1 | 7.41 | A | A | A |
| 17.1 | 1.2 | 0.0 | 8.12 | A | A | A |
| 16.8 | 1.07 | -1.8 | 7.57 | A | A | A |
| 17.5 | 0.6 | 2.3 | 5.34 | A | A | A |
| 17.57 | 1.37 | 2.8 | 8.81 | A | A | A |
| 16.7 | 1.1 | -2.3 | 7.76 | A | A | A |
| 17.26 | 0.37 | 0.9 | 4.62 | A | A | A |
| 16.79 | 0.429 | -1.8 | 4.83 | A | A | A |
| 17.133 | 1.443 | 0.2 | 9.36 | A | A | A |
| 16.946768 | 0.688071 | -0.9 | 5.77 | A | A | A |
| 16.4 | 1.2 | -4.1 | 8.38 | A | A | A |
| 10.57 | 0.45 | -38.2 | 5.91 | N | N | N |
| 17.3 | 1.1 | 1.2 | 7.56 | A | A | A |
| 14.9 | 3.07 | -12.9 | 21.01 | A | N | W |
| 16.93 | 0.82 | -1.0 | 6.34 | A | A | A |
| 18.12 | 0.55 | 6.0 | 5.10 | A | A | A |
| 17.3 | 0.8 | 1.2 | 6.18 | A | A | A |
| 17.2 | 1 | 0.6 | 7.11 | A | A | A |
| 17.1 | 0.4 | 0.0 | 4.71 | A | A | A |
| 15.8 | 0.26 | -7.6 | 4.41 | A | A | A |
| 17.7 | 1.4 | 3.5 | 8.91 | A | A | A |
| 17 | 2 | -0.6 | 12.46 | A | A | A |
| 16.76 | 1.59 | -2.0 | 10.33 | A | A | A |
| 16.9 | 0.8 | -1.2 | 6.26 | A | A | A |
| 17.2 | 1.3 | 0.6 | 8.60 | A | A | A |
| 17.7 | 0.6 | 3.5 | 5.31 | A | A | A |
| 16.9 | 1.1 | -1.2 | 7.69 | A | A | A |
| 18.4 | 0.909983 | 7.6 | 6.42 | A | A | A |
| 17 | 0.9 | -0.6 | 6.69 | A | A | A |
| 16.7 | 0.8 | -2.3 | 6.30 | A | A | A |
| 17.7 | 1.9 | 3.5 | 11.49 | A | A | A |
| 17.1 | 0.6 | 0.0 | 5.39 | A | A | A |
| 17.03 | 0.42 | -0.4 | 4.78 | A | A | A |
| 17.3 | 0.44 | 1.2 | 4.82 | A | A | A |
| 19.811 | 0.3758 | 15.8 | 4.51 | A | N | W |
| 16.62 | 0.88 | -2.8 | 6.69 | A | A | A |
| 17.1 | 0.7 | 0.0 | 5.79 | A | A | A |
| 18.4 | 0.4 | 7.6 | 4.63 | A | A | A |
| 17.6 | 0.5 | 2.9 | 4.98 | A | A | A |
| 19 | 4 | 11.1 | 21.45 | A | N | W |
| 17 | 1.3 | -0.6 | 8.67 | A | A | A |
| 17.11 | 1.797 | 0.1 | 11.27 | A | A | A |
| 16.8 | 0.4 | -1.8 | 4.74 | A | A | A |

Labcode

208

209

211

213

214

216

217

218

220

223

224

226

227

228

229

230

232

233

235

238

239

240

241

242

243

247

249

250

252

253

254

256

257

258

260

261

262

264

265

267

268

270

271

272

274

276

277

278

279

281

283

284

285

286

287

288

289

290

291

292

| Reported value [Bq/kg ] 16.4 | Reported uncertainty (k = 1) [Bq/kg ] 1.8 | Relative bias [%] -4.1 | P-Test [%] 11.71 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 17.7 | 1.6 | 3.5 | 9.92 | A | A | A |
| 17.45 | 0.63 | 2.0 | 5.46 | A | A | A |
| 18.1 | 0.712 | 5.8 | 5.68 | A | A | A |
| 17.02 | 0.94 | -0.5 | 6.87 | A | A | A |
| 17.2 | 1.7 | 0.6 | 10.70 | A | A | A |
| 16.57 | 0.39 | -3.1 | 4.72 | A | A | A |
| 17.3 | 1.4 | 1.2 | 9.07 | A | A | A |
| 16.5 | 0.8 | -3.5 | 6.35 | A | A | A |
| 16.8 | 1.4 | -1.8 | 9.28 | A | A | A |
| 17.1 | 2.6 | 0.0 | 15.75 | A | A | A |
| 17.2 | 0.7 | 0.6 | 5.77 | A | A | A |
| 16.9 | 1.2 | -1.2 | 8.20 | A | A | A |
| 18.9 | 0.87 | 10.5 | 6.16 | A | A | A |
| 16.9 | 1.8 | -1.2 | 11.41 | A | A | A |
| 17.8 | 0.984 | 4.1 | 6.88 | A | A | A |
| 18.37 | 0.55 | 7.4 | 5.07 | A | A | A |
| 17.6 | 0.86 | 2.9 | 6.37 | A | A | A |
| 17.7 | 1.9 | 3.5 | 11.49 | A | A | A |
| 20.8 | 1.7 | 21.6 | 9.14 | N | A | N |
| 17.22 | 0.25 | 0.7 | 4.34 | A | A | A |
| 17.7 | 0.771 | 3.5 | 5.98 | A | A | A |
| 17.1 | 0.6 | 0.0 | 5.39 | A | A | A |
| 20.7 | 1.3 | 21.0 | 7.50 | N | N | N |
| 17.5 | 0.96 | 2.3 | 6.84 | A | A | A |
| 17.06 | 1 | -0.2 | 7.15 | A | A | A |
| 16.7 | 1.3 | -2.3 | 8.80 | A | A | A |
| 16.2 | 0.5 | -5.3 | 5.13 | A | A | A |
| 17.1 | 1.4 | 0.0 | 9.15 | A | A | A |
| 17.78 | 1.23 | 4.0 | 8.04 | A | A | A |
| 17.5 | 0.5 | 2.3 | 4.99 | A | A | A |
| 16.9 | 0.5 | -1.2 | 5.05 | A | A | A |
| 17.55 | 0.25 | 2.6 | 4.33 | A | A | A |
| 18.2 | 1.18 | 6.4 | 7.67 | A | A | A |
| 17.1 | 1.4 | 0.0 | 9.15 | A | A | A |
| 16.6 | 1 | -2.9 | 7.28 | A | A | A |
| 16.08 | 0.98 | -6.0 | 7.34 | A | A | A |
| 17.2 | 1.2 | 0.6 | 8.09 | A | A | A |
| 15 | 0.5 | -12.3 | 5.28 | A | A | A |
| 16 | 0.8 | -6.4 | 6.46 | A | A | A |
| 17.6 | 0.7 | 2.9 | 5.71 | A | A | A |
| 16.25 | 0.23 | -5.0 | 4.33 | A | A | A |
| 17.2 | 0.8 | 0.6 | 6.20 | A | A | A |
| 16.8 | 0.9 | -1.8 | 6.74 | A | A | A |
| 22.5 | 1.1 | 31.6 | 6.38 | N | N | N |
| 16.3 | 0.7 | -4.7 | 5.93 | A | A | A |
| 18.9 | 1.9 | 10.5 | 10.85 | A | A | A |
| 17.7 | 0.7 | 3.5 | 5.69 | A | A | A |
| 18.2 | 1.6 | 6.4 | 9.70 | A | A | A |
| 17.9 | 0.24 | 4.7 | 4.31 | A | A | A |
| 16.1 | 2.2 | -5.8 | 14.26 | A | A | A |
| 16.38 | 0.82 | -4.2 | 6.47 | A | A | A |
| 17.2 | 0.87 | 0.6 | 6.51 | A | A | A |
| 20.1 | 0.8 | 17.5 | 5.71 | A | N | W |
| 17.4 | 2.1 | 1.8 | 12.74 | A | A | A |
| 16.66 | 0.6 | -2.6 | 5.45 | A | A | A |
| 17.19 | 1.03 | 0.5 | 7.26 | A | A | A |
| 20.3 | 1 | 18.7 | 6.40 | A | N | W |
| 16.2 | 1.6 | -5.3 | 10.69 | A | A | A |
| 21.2 | 0.89 | 24.0 | 5.86 | N | N | N |

Labcode

293

294

295

296

297

299

300

302

303

304

306

307

310

311

312

313

314

315

316

317

318

319

322

323

325

326

328

331

332

338

339

341

342

343

344

346

347

348

350

351

354

355

356

357

358

363

364

367

368

369

370

371

374

376

379

381

382

383

384

385

| Reported value [Bq/kg ] 16.8 | Reported uncertainty (k = 1) [Bq/kg ] 2.22 | Relative bias [%] -1.8 | P-Test [%] 13.83 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 17 | 0.4 | -0.6 | 4.72 | A | A | A |
| 17.1 | 0.4 | 0.0 | 4.71 | A | A | A |
| 16.02 | 0.69 | -6.3 | 5.94 | A | A | A |
| 17.1 | 0.78 | 0.0 | 6.13 | A | A | A |
| 16.55 | 0.92 | -3.2 | 6.90 | A | A | A |
| 17.31 | 1.7 | 1.2 | 10.64 | A | A | A |
| 17.47 | 1.12 | 2.2 | 7.61 | A | A | A |
| 17 | 1.3 | -0.6 | 8.67 | A | A | A |
| 18.2 | 1.4 | 6.4 | 8.71 | A | A | A |
| 16.2 | 0.9 | -5.3 | 6.90 | A | A | A |
| 16.92 | 1.03 | -1.0 | 7.34 | A | A | A |
| 17.11 | 0.97 | 0.1 | 6.99 | A | A | A |
| 16.5 | 0.9 | -3.5 | 6.82 | A | A | A |
| 18.3 | 0.9 | 7.0 | 6.40 | A | A | A |
| 16.37 | 0.46 | -4.3 | 4.97 | A | A | A |
| 18.1 | 2.2 | 5.8 | 12.83 | A | A | A |
| 17.7 | 0.5 | 3.5 | 4.97 | A | A | A |
| 16.17 | 0.79 | -5.4 | 6.37 | A | A | A |
| 17.3 | 1.1 | 1.2 | 7.56 | A | A | A |
| 17.1 | 0.6 | 0.0 | 5.39 | A | A | A |
| 20.2 | 2 | 18.1 | 10.71 | A | A | A |
| 17.6 | 1.45 | 2.9 | 9.20 | A | A | A |
| 17.2 | 0.7 | 0.6 | 5.77 | A | A | A |
| 18 | 0.8 | 5.3 | 6.04 | A | A | A |
| 16.5 | 0.8 | -3.5 | 6.35 | A | A | A |
| 16.46 | 1.99 | -3.7 | 12.76 | A | A | A |
| 18.3 | 0.205 | 7.0 | 4.24 | A | A | A |
| 17.02 | 0.52 | -0.5 | 5.11 | A | A | A |
| 18.42 | 0.9 | 7.7 | 6.37 | A | A | A |
| 19.3 | 1.9 | 12.9 | 10.66 | A | A | A |
| 16.8 | 1.4 | -1.8 | 9.28 | A | A | A |
| 17.6 | 0.9 | 2.9 | 6.55 | A | A | A |
| 20.879 | 0.732 | 22.1 | 5.39 | N | N | N |
| 17.5 | 1 | 2.3 | 7.03 | A | A | A |
| 23.6439 | 1.374118 | 38.3 | 7.11 | N | N | N |
| 16.7 | 0.7 | -2.3 | 5.86 | A | A | A |
| 17.24 | 0.77 | 0.8 | 6.06 | A | A | A |
| 17.9 | 0.9 | 4.7 | 6.48 | A | A | A |
| 18.7 | 0.6 | 9.4 | 5.20 | A | A | A |
| 17.17 | 0.29 | 0.4 | 4.43 | A | A | A |
| 17.5 | 0.5 | 2.3 | 4.99 | A | A | A |
| 16.18 | 0.977 | -5.4 | 7.30 | A | A | A |
| 17 | 0.656 | -0.6 | 5.63 | A | A | A |
| 17.8 | 0.8 | 4.1 | 6.08 | A | A | A |
| 14.9 | 1.1 | -12.9 | 8.44 | A | A | A |
| 17.77 | 1.38 | 3.9 | 8.78 | A | A | A |
| 17.2 | 1 | 0.6 | 7.11 | A | A | A |
| 26 | 2 | 52.0 | 8.71 | N | N | N |
| 18.8 | 1 | 9.9 | 6.71 | A | A | A |
| 17.75 | 1.59 | 3.8 | 9.85 | A | A | A |
| 17.1 | 1.5 | 0.0 | 9.68 | A | A | A |
| 19.515 | 0.70835 | 14.1 | 5.47 | A | N | W |
| 17.2 | 0.88 | 0.6 | 6.55 | A | A | A |
| 17.4 | 0.64 | 1.8 | 5.50 | A | A | A |
| 19.3 | 1.2 | 12.9 | 7.44 | A | A | A |
| 17.7 | 0.9 | 3.5 | 6.53 | A | A | A |
| 17.3 | 1.2 | 1.2 | 8.05 | A | A | A |
| 17.3 | 0.9 | 1.2 | 6.62 | A | A | A |
| 17.07 | 1.04 | -0.2 | 7.34 | A | A | A |

Labcode

386

387

388

391

392

393

395

396

398

401

402

403

404

405

407

408

410

411

412

414

417

419

420

421

423

425

426

429

433

434

435

437

439

440

441

442

443

444

445

446

448

450

452

453

454

455

459

460

462

463

464

465

466

469

471

472

475

477

478

479

| Reported value [Bq/kg ] 17.1 | Reported uncertainty (k = 1) [Bq/kg ] 0.8 | Relative bias [%] 0.0 | P-Test [%] 6.22 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 17.5 | 1 | 2.3 | 7.03 | A | A | A |
| 16.8 | 0.8 | -1.8 | 6.28 | A | A | A |
| 17.5 | 1.4 | 2.3 | 8.99 | A | A | A |
| 17.5 | 0.81 | 2.3 | 6.18 | A | A | A |
| 5.219 | 0.176 | -69.5 | 5.30 | N | N | N |
| 16 | 1 | -6.4 | 7.47 | A | A | A |
| 17.39 | 0.73 | 1.7 | 5.86 | A | A | A |
| 16.99 | 0.34 | -0.6 | 4.56 | A | A | A |
| 16.32 | 1.31 | -4.6 | 9.01 | A | A | A |
| 15.8 | 1.6 | -7.6 | 10.92 | A | A | A |
| 16.48 | 1.324 | -3.6 | 9.02 | A | A | A |
| 18.4 | 1 | 7.6 | 6.80 | A | A | A |
| 17.7 | 0.7 | 3.5 | 5.69 | A | A | A |
| 15.735 | 0.106 | -8.0 | 4.15 | A | A | A |
| 17.26 | 0.32 | 0.9 | 4.49 | A | A | A |
| 17.27 | 0.26 | 1.0 | 4.36 | A | A | A |
| 18.3 | 1 | 7.0 | 6.83 | A | A | A |
| 17.2 | 1.1 | 0.6 | 7.59 | A | A | A |
| 15.05 | 0.92 | -12.0 | 7.36 | A | A | A |
| 17.4 | 0.5 | 1.8 | 5.00 | A | A | A |
| 17.4 | 0.7 | 1.8 | 5.74 | A | A | A |
| 17.3 | 1.1 | 1.2 | 7.56 | A | A | A |
| 16.5 | 1.1 | -3.5 | 7.82 | A | A | A |
| 17 | 0.5 | -0.6 | 5.04 | A | A | A |
| 17.41 | 0.98 | 1.8 | 6.96 | A | A | A |
| 17.405 | 0.73 | 1.8 | 5.86 | A | A | A |
| 17.535575 | 0.760674 | 2.5 | 5.96 | A | A | A |
| 18.3 | 1.1 | 7.0 | 7.27 | A | A | A |
| 15.9 | 1.9 | -7.0 | 12.63 | A | A | A |
| 18.8 | 1 | 9.9 | 6.71 | A | A | A |
| 17.82 | 0.75 | 4.2 | 5.87 | A | A | A |
| 16.38 | 0.66 | -4.2 | 5.74 | A | A | A |
| 19.59 | 0.59 | 14.6 | 5.08 | A | N | W |
| 17.43 | 0.83 | 1.9 | 6.28 | A | A | A |
| 17.4 | 1.2 | 1.8 | 8.02 | A | A | A |

Labcode

482

483

484

487

490

491

492

495

496

498

499

500

501

506

507

508

509

514

516

520

522

523

525

527

528

529

530

532

533

535

536

537

539

544

553

555

Po-210 determination of activity concentration in Sample 1

Table 16. Target values

| Parameter | Value |
| --- | --- |
| Activity concentration [ Bq/kg ] | 29.9 |
| Uncertainty [ Bq/kg ] | 1.5 |
| Maximum Acceptable Relative Bias [%] | 30 |

Accepted 18%	Not Accepted 8%	Warning 2%	Not Reported 72%

40

45

Figure 16. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

15

20

25

535

71

423

217

314

528

440

156

100

170

266

303

363

58

204

38

32

72

370

414

230

128

241

250

445

466

523

161

401

63

70

65

187

134

262

264

340

166

15

412

231

312

181

315

163

232

347

393

319

90

355

439

132

443

553

151

368

66

175

537

219

283

341

530

408

238

525

177

19

52

75

25

317

178

34

197

322

191

539

291

49

516

371

527

433

255

379

192

536

152

220

432

308

529

533

31

313

85

492

207

276

441

483

405

213

455

350

403

Laboratory code

Figure 17. Reported results and evaluation parameters for Po-210. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 17. Evaluation results for Po-210 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

Labcode

Reported value

[Bq/kg ]

Reported uncertainty

(k = 1) [Bq/kg ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

Activity concentration  [ Bq/kg ]

30

35

15	26.6	1.4	-11.0	7.27	A	A	A

19	29.6	1	-1.0	6.05	A	A	A

25	30.3	2	1.3	8.29	A	A	A

| Reported value [Bq/kg ] 71.7 | Reported uncertainty (k = 1) [Bq/kg ] 4.59 | Relative bias [%] 139.8 | P-Test [%] 8.13 | Trueness evaluation N | Precision evaluation N | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 24.2 | 2.9 | -19.1 | 12.99 | A | A | A |
| 31.3 | 11.3 | 4.7 | 36.45 | A | N | W |
| 24.03 | 2.62 | -19.6 | 12.00 | A | A | A |
| 33.5 | 5.1 | 12.0 | 16.03 | A | A | A |
| 30 | 2.84 | 0.3 | 10.71 | A | A | A |
| 24 | 1 | -19.7 | 6.52 | A | N | W |
| 25.7 | 0.7 | -14.1 | 5.71 | A | A | A |
| 25.89 | 3.82 | -13.4 | 15.58 | A | A | A |
| 28.47 | 1.14 | -4.8 | 6.42 | A | A | A |
| 25.8 | 0.7 | -13.7 | 5.70 | A | A | A |
| 1.78 | 0.09 | -94.0 | 7.12 | N | N | N |
| 24.2 | 0.87 | -19.1 | 6.17 | A | N | W |
| 30.28 | 1.27 | 1.3 | 6.54 | A | A | A |
| 78.847 | 2.434 | 163.7 | 5.89 | N | N | N |
| 27.65 | 0.64 | -7.5 | 5.52 | A | A | A |
| 19 | 0.2 | -36.5 | 5.13 | N | N | N |
| 24.81 | 2.64 | -17.0 | 11.76 | A | A | A |
| 28.1 | 2.5 | -6.0 | 10.21 | A | A | A |
| 25.9 | 0.8 | -13.4 | 5.89 | A | A | A |
| 28.19 | 3.04 | -5.7 | 11.89 | A | A | A |
| 44.4 | 3.9 | 48.5 | 10.12 | N | N | N |
| 17.74 | 0.57 | -40.7 | 5.96 | N | N | N |
| 25.6 | 2.05 | -14.4 | 9.45 | A | A | A |
| 27.1 | 1.1 | -9.4 | 6.45 | A | A | A |
| 26.438392 | 3.113584 | -11.6 | 12.80 | A | A | A |
| 21.5 | 1.9 | -28.1 | 10.16 | A | N | W |
| 28.51 | 0.69 | -4.7 | 5.57 | A | A | A |
| 29.23 | 2.47 | -2.2 | 9.83 | A | A | A |
| 31 | 1.4 | 3.7 | 6.75 | A | A | A |
| 27.02 | 1.9 | -9.6 | 8.64 | A | A | A |
| 25.89 | 3.82 | -13.4 | 15.58 | A | A | A |
| 31.9 | 1.6 | 6.7 | 7.09 | A | A | A |
| 39.9 | 3.1 | 33.4 | 9.25 | N | N | N |
| 31.4 | 6.22 | 5.0 | 20.43 | A | A | A |
| 24 | 2 | -19.7 | 9.73 | A | A | A |
| 95 | 4 | 217.7 | 6.55 | N | N | N |
| 128 | 11.5 | 328.1 | 10.29 | N | N | N |
| 3.56 | 0.56 | -88.1 | 16.51 | N | N | N |
| 28.53 | 1.82 | -4.6 | 8.12 | A | A | A |
| 53.1 | 3.4 | 77.6 | 8.13 | N | N | N |
| 24.8 | 1.2 | -17.1 | 6.97 | A | A | A |
| 27 | 4 | -9.7 | 15.64 | A | A | A |
| 27.1 | 0.9 | -9.4 | 6.02 | A | A | A |
| 29.1 | 1.75 | -2.7 | 7.83 | A | A | A |
| 25.3 | 1.3 | -15.4 | 7.18 | A | A | A |
| 25.5 | 1.3 | -14.7 | 7.15 | A | A | A |
| 38 | 11 | 27.1 | 29.38 | A | A | A |
| 26.2 | 0.31 | -12.4 | 5.15 | A | A | A |
| 26.3 | 3.5 | -12.0 | 14.22 | A | A | A |
| 21.6 | 2.1 | -27.8 | 10.94 | A | A | A |
| 98 | 3 | 227.8 | 5.88 | N | N | N |
| 28.6 | 1.58 | -4.3 | 7.46 | A | A | A |
| 33.355 | 6 | 11.6 | 18.67 | A | A | A |
| 23.04 | 1.38 | -22.9 | 7.81 | A | N | W |
| 60.8 | 3.2 | 103.3 | 7.27 | N | N | N |
| 27 | 3 | -9.7 | 12.19 | A | A | A |
| 77 | 5.6 | 157.5 | 8.84 | N | N | N |
| 10.15 | 1.15 | -66.0 | 12.39 | N | N | N |
| 27.051 | 1.169 | -9.5 | 6.62 | A | A | A |

Labcode

31

32

34

38

49

52

58

63

65

66

70

71

72

75

85

90

100

128

132

134

151

152

156

161

163

166

170

175

177

178

181

187

191

192

197

204

207

213

217

219

220

230

231

232

238

241

250

255

262

264

266

276

283

291

303

308

312

313

314

315

| Reported value [Bq/kg ] 30.5 | Reported uncertainty (k = 1) [Bq/kg ] 3.3 | Relative bias [%] 2.0 | P-Test [%] 11.93 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 27.5 | 1.9 | -8.0 | 8.54 | A | A | A |
| 31.5 | 3.5 | 5.3 | 12.19 | A | A | A |
| 26.3798 | 2.5673 | -11.8 | 10.95 | A | A | A |
| 28.7 | 1.4 | -4.0 | 7.00 | A | A | A |
| 27.1 | 4.5 | -9.4 | 17.35 | A | A | A |
| 142 | 4.5 | 374.9 | 5.93 | N | N | N |
| 27.71 | 0.79 | -7.3 | 5.77 | A | A | A |
| 23.1 | 2 | -22.7 | 10.01 | A | A | A |
| 28.3 | 1.7 | -5.3 | 7.83 | A | A | A |
| 24.3 | 2.5 | -18.7 | 11.45 | A | A | A |
| 34.47 | 4.07 | 15.3 | 12.83 | A | A | A |
| 38.6 | 4.4 | 29.1 | 12.45 | A | A | A |
| 27.16 | 1.93 | -9.2 | 8.70 | A | A | A |
| 25.6 | 0.9 | -14.4 | 6.13 | A | A | A |
| 334.32 | 71.52 | 1018.1 | 21.97 | N | N | N |
| 107.6 | 4.5 | 259.9 | 6.53 | N | N | N |
| 29 | 0.5 | -3.0 | 5.30 | A | A | A |
| 26.65 | 1.66 | -10.9 | 8.00 | A | A | A |
| 24.4 | 0.8 | -18.4 | 5.99 | A | N | W |
| 2.72 | 0.22 | -90.9 | 9.52 | N | N | N |
| 56.9 | 4.23 | 90.3 | 8.97 | N | N | N |
| 36.38 | 0.25 | 21.7 | 5.06 | A | N | W |
| 27.9 | 1.8 | -6.7 | 8.17 | A | A | A |
| 14.7723 | 0.0825 | -50.6 | 5.05 | N | N | N |
| 106.8 | 5.7 | 257.2 | 7.32 | N | N | N |
| 28.1 | 2.1 | -6.0 | 9.00 | A | A | A |
| 25.5 | 2.5 | -14.7 | 11.01 | A | A | A |
| 136 | 7 | 354.9 | 7.19 | N | N | N |
| 25.5 | 1.785 | -14.7 | 8.61 | A | A | A |
| 106.8 | 5.7 | 257.2 | 7.32 | N | N | N |
| 81.7 | 3 | 173.2 | 6.22 | N | N | N |
| 34.3 | 2.3 | 14.7 | 8.37 | A | A | A |
| 25.5 | 0.7 | -14.7 | 5.72 | A | A | A |
| 29.1 | 1.7 | -2.7 | 7.70 | A | A | A |
| 35.3 | 2.4 | 18.1 | 8.45 | A | A | A |
| 12.11 | 0.64 | -59.5 | 7.29 | N | N | N |
| 61.8 | 1.9 | 106.7 | 5.88 | N | N | N |
| 28.81 | 9.046 | -3.6 | 31.80 | A | N | W |
| 64.7 | 4.6 | 116.4 | 8.70 | N | N | N |
| 0.029 | 0.003 | -99.9 | 11.50 | N | N | N |
| 40.4 | 1.4 | 35.1 | 6.10 | N | N | N |
| 28.51 | 1.37 | -4.7 | 6.95 | A | A | A |
| 32.1 | 2.71 | 7.4 | 9.82 | A | A | A |
| 28.1 | 1.8 | -6.0 | 8.14 | A | A | A |

Labcode

317

319

322

340

341

347

350

355

363

368

370

371

379

393

401

403

405

408

412

414

423

432

433

439

440

441

443

445

455

466

483

492

516

523

525

527

528

529

530

533

535

536

537

539

553

Pb-210 determination of activity concentration in Sample 1

Table 18. Target values

| Parameter | Value |
| --- | --- |
| Activity concentration [ Bq/kg ] | 29.4 |
| Uncertainty [ Bq/kg ] | 1.2 |
| Maximum Acceptable Relative Bias [%] | 30 |

Accepted 30%	Not Accepted 3%	Warning 1%	Not Reported 66%

40

45

Figure 18. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

15

20

25

499

217

401

293

287

475

90

239

202

484

370

408

363

38

70

177

283

391

454

300

128

250

72

231

355

55

326

347

404

232

393

15

207

41

70

452

110

178

205

45

25

264

290

31

256

119

421

403

204

479

487

191

276

319

238

270

443

75

247

312

356

175

465

156

20

203

151

127

159

52

235

523

528

262

509

60

258

318

525

295

272

53

501

58

214

161

166

553

242

147

414

555

537

2

371

425

433

38

317

382

227

539

156

279

411

218

155

533

412

407

48

315

299

68

498

63

466

111

460

348

152

414

73

364

170

322

536

206

197

379

57

278

419

24

286

423

249

Laboratory code

Figure 19. Reported results and evaluation parameters for Pb-210. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 19. Evaluation results for Pb-210 in Sample 1, Rep. Value and Rep. Unc. in Bq/kg

Labcode

Reported value

[Bq/kg ]

Reported uncertainty

(k = 1) [Bq/kg ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

Activity concentration  [ Bq/kg ]

30

35

2	31.5	2.9	7.1	10.07	A	A	A

15	27.6	1.9	-6.1	8.00	A	A	A

20	30.27	3.97	3.0	13.74	A	A	A

| Reported value [Bq/kg ] 45.3 | Reported uncertainty (k = 1) [Bq/kg ] 6.8 | Relative bias [%] 54.1 | P-Test [%] 15.56 | Trueness evaluation N | Precision evaluation N | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 28.5 | 4.1 | -3.1 | 14.95 | A | A | A |
| 28.6 | 1.37 | -2.7 | 6.29 | A | A | A |
| 25.67 | 1.76 | -12.7 | 7.98 | A | A | A |
| 31.82 | 5.12 | 8.2 | 16.60 | A | A | A |
| 27.7 | 2.68 | -5.8 | 10.50 | A | A | A |
| 28.47 | 2.03 | -3.2 | 8.22 | A | A | A |
| 32.8 | 2.7 | 11.6 | 9.19 | A | A | A |
| 30.5 | 3.33 | 3.7 | 11.66 | A | A | A |
| 30.82 | 3 | 4.8 | 10.56 | A | A | A |
| 27.04 | 7.76 | -8.0 | 28.99 | A | A | A |
| 40 | 4 | 36.0 | 10.80 | N | N | N |
| 31 | 2 | 5.4 | 7.63 | A | A | A |
| 30.7 | 3.8 | 4.4 | 13.03 | A | A | A |
| 33.75 | 2.01 | 14.8 | 7.22 | A | A | A |
| 33.3 | 4.4 | 13.3 | 13.83 | A | A | A |
| 27.7 | 1.5 | -5.8 | 6.78 | A | A | A |
| 25.7 | 2.2 | -12.6 | 9.48 | A | A | A |
| 26.8 | 5.2 | -8.8 | 19.83 | A | A | A |
| 35.6 | 5.3 | 21.1 | 15.44 | A | A | A |
| 29.62 | 7.44 | 0.8 | 25.45 | A | A | A |
| 24.1 | 2.5 | -18.0 | 11.15 | A | A | A |
| 27.884487 | 2.037853 | -5.2 | 8.37 | A | A | A |
| 33.81 | 3.3 | 15.0 | 10.58 | A | A | A |
| 28.7 | 3.4 | -2.4 | 12.53 | A | A | A |
| 30.361828 | 4.217186 | 3.3 | 14.48 | A | A | A |
| 26.14 | 1.99 | -11.1 | 8.64 | A | A | A |
| 31.2842 | 1.0455 | 6.4 | 5.28 | A | A | A |
| 30.32 | 3.03 | 3.1 | 10.79 | A | A | A |
| 35.1 | 5 | 19.4 | 14.82 | A | A | A |
| 32.4 | 3.9 | 10.2 | 12.71 | A | A | A |
| 32.26 | 1.59 | 9.7 | 6.40 | A | A | A |
| 30.1 | 1.1 | 2.4 | 5.48 | A | A | A |
| 30.4 | 2.6 | 3.4 | 9.48 | A | A | A |
| 31.1 | 1.5 | 5.8 | 6.32 | A | A | A |
| 31.1 | 4.9 | 5.8 | 16.28 | A | A | A |
| 36 | 6 | 22.4 | 17.16 | A | A | A |
| 29.9 | 1.8 | 1.7 | 7.27 | A | A | A |
| 25.7 | 1.9 | -12.6 | 8.44 | A | A | A |
| 27.9 | 3.4 | -5.1 | 12.85 | A | A | A |
| 29.2 | 3.6 | -0.7 | 12.99 | A | A | A |
| 38.34 | 6.52 | 30.4 | 17.49 | N | A | N |
| 24.76 | 1.27 | -15.8 | 6.56 | A | A | A |
| 30.3 | 6.6 | 3.1 | 22.16 | A | A | A |
| 29 | 1 | -1.4 | 5.34 | A | A | A |
| 28.1 | 2.08 | -4.4 | 8.45 | A | A | A |
| 37.8 | 7.6 | 28.6 | 20.52 | A | A | A |
| 27.6 | 3 | -6.1 | 11.61 | A | A | A |
| 31 | 3 | 5.4 | 10.50 | A | A | A |
| 3.56 | 0.56 | -87.9 | 16.25 | N | N | N |
| 32.36915 | 3.39379 | 10.1 | 11.25 | A | A | A |
| 32.2 | 1.6 | 9.5 | 6.43 | A | A | A |
| 27 | 5 | -8.2 | 18.96 | A | A | A |
| 27.5 | 3.5 | -6.5 | 13.37 | A | A | A |
| 30.5 | 2.7 | 3.7 | 9.75 | A | A | A |
| 29.4 | 5.68 | 0.0 | 19.75 | A | A | A |
| 24.1 | 5.2 | -18.0 | 21.96 | A | A | A |
| 31.21 | 0.99 | 6.2 | 5.17 | A | A | A |
| 29.684 | 5.012 | 1.0 | 17.37 | A | A | A |
| 60.784722 | 2.6 | 106.8 | 5.91 | N | N | N |

Labcode

24

25

31

38

38

41

45

48

52

53

55

57

58

60

63

68

70

70

72

73

75

90

110

111

119

127

128

147

151

152

155

156

156

159

161

166

170

175

177

178

191

197

202

203

204

205

206

207

214

217

218

227

231

232

235

238

239

242

247

249

| Reported value [Bq/kg ] 26.2 | Reported uncertainty (k = 1) [Bq/kg ] 2.5 | Relative bias [%] -10.9 | P-Test [%] 10.38 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 28.6 | 4.03 | -2.7 | 14.67 | A | A | A |
| 30.7 | 5.6 | 4.4 | 18.69 | A | A | A |
| 30.57 | 3.04 | 4.0 | 10.75 | A | A | A |
| 28.5 | 2.5 | -3.1 | 9.68 | A | A | A |
| 29.5 | 3 | 0.3 | 10.96 | A | A | A |
| 30.8 | 2.1 | 4.8 | 7.95 | A | A | A |
| 29.2 | 1.5 | -0.7 | 6.56 | A | A | A |
| 42.6 | 10.2 | 44.9 | 24.29 | N | A | N |
| 32.3 | 2.1 | 9.9 | 7.68 | A | A | A |
| 25.7 | 1.07 | -12.6 | 5.83 | A | A | A |
| 55.5 | 15.7 | 88.8 | 28.58 | N | N | N |
| 21.7 | 7.2 | -26.2 | 33.43 | A | N | W |
| 28.5 | 3 | -3.1 | 11.29 | A | A | A |
| 19.9 | 3.6 | -32.3 | 18.55 | N | A | N |
| 30.73 | 1.65 | 4.5 | 6.74 | A | A | A |
| 33 | 4 | 12.2 | 12.79 | A | A | A |
| 26.13 | 2.06 | -11.1 | 8.88 | A | A | A |
| 29.8 | 3 | 1.4 | 10.86 | A | A | A |
| 32.933 | 2.95 | 12.0 | 9.84 | A | A | A |
| 32 | 3.4 | 8.8 | 11.38 | A | A | A |
| 30.7 | 1.3 | 4.4 | 5.88 | A | A | A |
| 29.2 | 1.69 | -0.7 | 7.08 | A | A | A |
| 36.6 | 4.6 | 24.5 | 13.21 | A | A | A |
| 27.3 | 4.8 | -7.1 | 18.05 | A | A | A |
| 27.3 | 3.6 | -7.1 | 13.80 | A | A | A |
| 35 | 6 | 19.0 | 17.62 | A | A | A |
| 27.02 | 2.82 | -8.1 | 11.21 | A | A | A |
| 29.8 | 2.2 | 1.4 | 8.44 | A | A | A |
| 25.4 | 4.2 | -13.6 | 17.03 | A | A | A |
| 35.7 | 4 | 21.4 | 11.92 | A | A | A |
| 25.1 | 3.3 | -14.6 | 13.77 | A | A | A |
| 31.5 | 3.2 | 7.1 | 10.95 | A | A | A |
| 38.6 | 4.4 | 31.3 | 12.11 | N | N | N |
| 32 | 4.75 | 8.8 | 15.39 | A | A | A |
| 25.82 | 2.23 | -12.2 | 9.55 | A | A | A |
| 27.5 | 2.68 | -6.5 | 10.57 | A | A | A |
| 17.1 | 0.7 | -41.8 | 5.78 | N | N | N |
| 28.75 | 4.53 | -2.2 | 16.28 | A | A | A |
| 27.4 | 2.3 | -6.8 | 9.33 | A | A | A |
| 32.7 | 2.1 | 11.2 | 7.61 | A | A | A |
| 25.16 | 2.73 | -14.4 | 11.59 | A | A | A |
| 32.3 | 4.3 | 9.9 | 13.92 | A | A | A |
| 32.55 | 5.35 | 10.7 | 16.94 | A | A | A |
| 35.3 | 1.9 | 20.1 | 6.76 | A | N | W |
| 31.4 | 4.7 | 6.8 | 15.51 | A | A | A |
| 43 | 12 | 46.3 | 28.20 | N | A | N |
| 28.7 | 2.2 | -2.4 | 8.68 | A | A | A |
| 56 | 4.8 | 90.5 | 9.49 | N | N | N |
| 31.6 | 3.3 | 7.5 | 11.21 | A | A | A |
| 31.65 | 2.42 | 7.7 | 8.67 | A | A | A |
| 29.6 | 1.8 | 0.7 | 7.32 | A | A | A |
| 27.72 | 2.56 | -5.7 | 10.10 | A | A | A |
| 25.9 | 1.7 | -11.9 | 7.73 | A | A | A |
| 34.6 | 3.2 | 17.7 | 10.11 | A | A | A |
| 30 | 2.7 | 2.0 | 9.88 | A | A | A |
| 33.807 | 3.44645 | 15.0 | 10.98 | A | A | A |
| 23.4 | 1.6 | -20.4 | 7.96 | A | A | A |
| 29 | 3.62 | -1.4 | 13.13 | A | A | A |
| 24.9 | 5.8 | -15.3 | 23.65 | A | A | A |

Labcode

250

256

258

262

264

270

272

276

278

279

283

286

287

290

293

295

299

300

312

315

317

318

319

322

326

347

348

355

356

363

364

370

371

379

382

391

393

401

403

404

407

408

411

412

414

414

419

421

423

425

433

443

452

454

460

465

466

475

479

484

| Reported value [Bq/kg ] 29 | Reported uncertainty (k = 1) [Bq/kg ] 0.3 | Relative bias [%] -1.4 | P-Test [%] 4.21 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 33.3 | 4.29 | 13.3 | 13.51 | A | A | A |
| 2.6 | 0.3 | -91.2 | 12.24 | N | N | N |
| 30.9 | 3.6 | 5.1 | 12.34 | A | A | A |
| 30.6 | 0.9 | 4.1 | 5.03 | A | A | A |
| 30.5 | 2.1 | 3.7 | 8.00 | A | A | A |
| 30.7 | 2.1 | 4.4 | 7.97 | A | A | A |
| 30.5 | 1.7 | 3.7 | 6.91 | A | A | A |
| 32.5 | 2.1 | 10.5 | 7.64 | A | A | A |
| 37.5 | 3.1 | 27.6 | 9.22 | A | N | W |
| 31.41 | 2.92 | 6.8 | 10.15 | A | A | A |
| 32.24 | 1.81 | 9.7 | 6.94 | A | A | A |
| 31.1 | 2.1 | 5.8 | 7.89 | A | A | A |
| 31.4 | 2.8 | 6.8 | 9.81 | A | A | A |

Labcode

487

498

499

501

509

523

525

528

533

536

537

539

553

555

Activity concentration  [ Bq/kg ]

0.10

0.12

Ra-226 determination of activity concentration in Sample 2

Table 20. Target values

Parameter	Value

Activity concentration [ Bq/kg ]	0.0962

Uncertainty [ Bq/kg ]	0.0058

Maximum Acceptable Relative Bias [%]	30

Accepted 10%	Not Accepted 10%	Warning 2%	Not Reported 78%

0.14

Figure 20. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

0.04

0.06

0.08

434

371

217

58

204

175

295

528

156

232

401

313

87

19

516

527

255

161

253

65

187

152

231

132

533

90

66

166

38

91

230

205

507

25

213

222

256

537

443

299

319

307

410

445

514

191

32

290

529

63

119

220

215

53

122

60

109

466

31

250

440

172

303

264

291

49

433

124

492

368

379

75

344

393

128

181

347

228

276

484

274

266

491

52

94

71

113

165

Laboratory code

Figure 21. Reported results and evaluation parameters for Ra-226. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 21. Evaluation results for Ra-226 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Reported value | Reported uncertainty | Relative bias | P-Test | Trueness | Precision | Final |
| --- | --- | --- | --- | --- | --- | --- |
| [Bq/kg ] | (k = 1) [Bq/kg ] | [%] | [%] | evaluation | evaluation | Score |
| 0.075 | 0.003 | -22.0 | 7.24 | A | N | W |
| 0.093 | 0.006 | -3.3 | 8.83 | A | A | A |
| 0.125 | 0.0298 | 29.9 | 24.59 | A | A | A |

Labcode

19

25

31

| Reported value [Bq/kg ] 0.105 | Reported uncertainty (k = 1) [Bq/kg ] 0.029 | Relative bias [%] 9.2 | P-Test [%] 28.27 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 0.088 | 0.007 | -8.5 | 9.98 | A | A | A |
| 0.16 | 0.03 | 66.3 | 19.70 | N | N | N |
| 3.31 | 0.32 | 3340.8 | 11.39 | N | N | N |
| 0.111 | 0.017 | 15.4 | 16.46 | A | A | A |
| 0.045 | 0.007 | -53.2 | 16.68 | N | N | N |
| 0.117 | 0.024 | 21.6 | 21.38 | A | A | A |
| 0.106 | 0.004 | 10.2 | 7.11 | A | A | A |
| 0.084 | 0.014 | -12.7 | 17.72 | A | A | A |
| 0.0875 | 0.0066 | -9.0 | 9.66 | A | A | A |
| 3.6 | 0.18 | 3642.2 | 7.83 | N | N | N |
| 0.334 | 0.031 | 247.2 | 11.07 | N | N | N |
| 0.0744 | 0.015 | -22.7 | 21.04 | A | A | A |
| 0.0874 | 0.0049 | -9.2 | 8.23 | A | A | A |
| 0.089 | 0.008 | -7.5 | 10.82 | A | A | A |
| 3.41 | 0.54 | 3444.7 | 16.94 | N | N | N |
| 0.118 | 0.016 | 22.7 | 14.84 | A | A | A |
| 7.94 | 1.95 | 8153.6 | 25.29 | N | N | N |
| 0.109 | 0.015 | 13.3 | 15.02 | A | A | A |
| 0.112 | 0.016 | 16.4 | 15.51 | A | A | A |
| 0.18 | 0.03 | 87.1 | 17.72 | N | N | N |
| 0.52 | 0.08 | 440.5 | 16.52 | N | N | N |
| 0.087 | 0.012 | -9.6 | 15.05 | A | A | A |
| 0.0843 | 0.0073 | -12.4 | 10.55 | A | A | A |
| 0.0627 | 0.0051 | -34.8 | 10.12 | N | N | N |
| 0.0795 | 0.0054 | -17.4 | 9.08 | A | A | A |
| 37.6 | 2.4 | 38985.2 | 8.78 | N | N | N |
| 0.0878 | 0.016 | -8.7 | 19.19 | A | A | A |
| 0.14 | 0.04 | 45.5 | 29.20 | N | A | N |
| 0.058 | 0.016 | -39.7 | 28.24 | N | A | N |
| 0.52 | 0.06 | 440.5 | 13.02 | N | N | N |
| 0.084 | 0.014 | -12.7 | 17.72 | A | A | A |
| 0.101 | 0.012 | 5.0 | 13.32 | A | A | A |
| 0.05 | 0.02 | -48.0 | 40.45 | N | N | N |
| 0.0895 | 0.0087 | -7.0 | 11.44 | A | A | A |
| 0.094 | 0.0173 | -2.3 | 19.37 | A | A | A |
| 0.11 | 0.016 | 14.3 | 15.75 | A | A | A |
| 0.017205 | 0.002 | -82.1 | 13.10 | N | N | N |
| 0.109 | 0.01 | 13.3 | 10.98 | A | A | A |
| 0.095 | 0.009 | -1.2 | 11.23 | A | A | A |
| 0.56 | 0.1 | 482.1 | 18.85 | N | N | N |
| 0.089 | 0.009 | -7.5 | 11.77 | A | A | A |
| 0.085 | 0.015 | -11.6 | 18.65 | A | A | A |
| 0.066 | 0.007 | -31.4 | 12.20 | N | A | N |
| 0.128 | 0.013 | 33.1 | 11.81 | N | N | N |
| 0.08 | 0.16 | -16.8 | 200.09 | A | N | W |
| 0.077 | 0.012 | -20.0 | 16.71 | A | A | A |
| 0.096 | 0.016 | -0.2 | 17.72 | A | A | A |
| 0.141 | 0.016 | 46.6 | 12.85 | N | N | N |
| 1.93 | 0.33 | 1906.2 | 18.13 | N | N | N |
| 1.51 | 0.15 | 1469.7 | 11.62 | N | N | N |
| 0.6 | 0.03 | 523.7 | 7.83 | N | N | N |
| 0.105 | 0.042 | 9.2 | 40.45 | A | N | W |
| 0.15 | 0.03 | 55.9 | 20.89 | N | N | N |
| 0.059 | 0.009 | -38.7 | 16.40 | N | A | N |
| 0.098 | 0.005 | 1.9 | 7.90 | A | A | A |
| 0.14 | 0.04 | 45.5 | 29.20 | N | A | N |
| 0.1 | 0.01 | 4.0 | 11.68 | A | A | A |
| 0.0734 | 0.0115 | -23.7 | 16.79 | A | A | A |
| 0.0982 | 0.0125 | 2.1 | 14.08 | A | A | A |

Labcode

32

38

49

52

53

58

60

63

65

66

71

75

87

90

91

94

109

113

119

122

124

128

132

152

156

161

165

166

172

175

181

187

191

204

205

213

215

217

220

222

228

230

231

232

250

253

255

256

264

266

274

276

290

291

295

299

303

307

313

319

| Reported value [Bq/kg ] 0.35 | Reported uncertainty (k = 1) [Bq/kg ] 0.08 | Relative bias [%] 263.8 | P-Test [%] 23.64 | Trueness evaluation N | Precision evaluation N | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 0.52 | 0.12 | 440.5 | 23.85 | N | N | N |
| 0.196 | 0.014 | 103.7 | 9.35 | N | N | N |
| 0.012 | 0.003 | -87.5 | 25.72 | N | N | N |
| 0.2 | 0.026 | 107.9 | 14.33 | N | N | N |
| 0.51 | 0.08 | 430.1 | 16.81 | N | N | N |
| 0.066 | 0.059 | -31.4 | 89.60 | N | N | N |
| 0.1 | 0.08 | 4.0 | 80.23 | A | N | W |
| 0.17 | 0.058 | 76.7 | 34.65 | N | N | N |
| 0 | 0 | -100.0 |  | N |  | N |
| 0.1299 | 0.0043 | 35.0 | 6.88 | N | N | N |
| 0.097 | 0.008 | 0.8 | 10.22 | A | A | A |
| 0.1 | 0.09 | 4.0 | 90.20 | A | N | W |
| 0.119 | 0.006545 | 23.7 | 8.16 | A | N | W |
| 1.3 | 0.7 | 1251.3 | 54.18 | N | N | N |
| 2.263 | 0.377 | 2252.4 | 17.72 | N | N | N |
| 0.188 | 0.014 | 95.4 | 9.58 | N | N | N |
| 0.09 | 0.009 | -6.4 | 11.68 | A | A | A |
| 0.1 | 0.015 | 4.0 | 16.17 | A | A | A |
| 0.0766 | 0.012 | -20.4 | 16.79 | A | A | A |
| 0.0769 | 0.012 | -20.1 | 16.73 | A | A | A |
| 0.06 | 0.01 | -37.6 | 17.72 | N | A | N |
| 0.105 | 0.035 | 9.2 | 33.87 | A | N | W |
| 0.087 | 0.021 | -9.6 | 24.88 | A | A | A |
| 0.096 | 0.009 | -0.2 | 11.15 | A | A | A |

Labcode

344

347

368

371

379

393

401

410

433

434

440

443

445

466

484

491

492

507

514

516

527

528

529

533

537

  - determination of activity concentration in Sample 2

Table 22. Target values

Parameter	Value

Activity concentration [ Bq/kg ]	0.0891

Uncertainty [ Bq/kg ]	0.00446

Maximum Acceptable Relative Bias [%]	30

Accepted 21%	Not Accepted 5%	Warning 2%	Not Reported 72%

0.12

0.14

Figure 22. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

0.04

0.06

0.08

434

278

441

483

291

49

355

330

402

60

38

28

132

15

156

371

432

205

264

90

87

340

461

217

315

268

133

181

308

393

31

347

348

379

213

2

368

41

114

410

208

516

528

55

535

175

285

63

287

25

166

311

81

231

392

437

475

204

276

286

443

537

527

192

399

313

66

42

159

525

191

152

19

58

346

207

414

32

52

65

187

230

161

304

219

312

128

423

220

161

255

536

73

370

419

508

427

170

163

11

303

433

401

283

412

492

533

445

377

71

113

463

Laboratory code

Figure 23. Reported results and evaluation parameters for U-234. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 23. Evaluation results for U-234 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

Labcode

Reported value

[Bq/kg ]

Reported uncertainty

(k = 1) [Bq/kg ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

Activity concentration  [ Bq/kg ]

0.10

2	0.083	0.011	-6.8	14.17	A	A	A

11	0.124	0.0101	39.2	9.56	N	N	N

15	0.073	0.01	-18.1	14.58	A	A	A

| Reported value [Bq/kg ] 0.095 | Reported uncertainty (k = 1) [Bq/kg ] 0.005 | Relative bias [%] 6.6 | P-Test [%] 7.26 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 0.087 | 0.002 | -2.4 | 5.51 | A | A | A |
| 0.0711 | 0.0141 | -20.2 | 20.45 | A | A | A |
| 0.0802 | 0.01195 | -10.0 | 15.72 | A | A | A |
| 0.097 | 0.008 | 8.9 | 9.65 | A | A | A |
| 0.071 | 0.012 | -20.3 | 17.63 | A | A | A |
| 0.0836 | 0.00431 | -6.2 | 7.19 | A | A | A |
| 0.0915 | 0.0029 | 2.7 | 5.92 | A | A | A |
| 0.061 | 0.009 | -31.5 | 15.58 | N | A | N |
| 0.097 | 0.0048 | 8.9 | 7.04 | A | A | A |
| 0.086 | 0.015 | -3.5 | 18.15 | A | A | A |
| 0.095 | 0.003 | 6.6 | 5.92 | A | A | A |
| 0.0676 | 0.0096 | -24.1 | 15.06 | A | A | A |
| 0.0863 | 0.0036 | -3.1 | 6.52 | A | A | A |
| 0.098 | 0.012 | 10.0 | 13.23 | A | A | A |
| 0.09132 | 0.0066 | 2.5 | 8.79 | A | A | A |
| 4.47 | 0.22 | 4916.8 | 7.02 | N | N | N |
| 0.11 | 0.006 | 23.5 | 7.40 | A | N | W |
| 0.089 | 0.008 | -0.1 | 10.29 | A | A | A |
| 0.0782 | 0.0078 | -12.2 | 11.16 | A | A | A |
| 0.0779 | 0.0035 | -12.6 | 6.73 | A | A | A |
| 17 | 1.87 | 18979.7 | 12.09 | N | N | N |
| 0.084 | 0.012 | -5.7 | 15.14 | A | A | A |
| 0.102 | 0.0106 | 14.5 | 11.53 | A | A | A |
| 0.0714 | 0.0051 | -19.9 | 8.72 | A | A | A |
| 0.08 | 0.01 | -10.2 | 13.46 | A | A | A |
| 0.0944 | 0.0067 | 6.0 | 8.69 | A | A | A |
| 0.0739 | 0.0033 | -17.1 | 6.71 | A | A | A |
| 0.09178 | 0.00464 | 3.0 | 7.11 | A | A | A |
| 0.105 | 0.011 | 17.9 | 11.61 | A | A | A |
| 0.099 | 0.014 | 11.1 | 15.00 | A | A | A |
| 0.116 | 0.08 | 30.2 | 69.15 | N | N | N |
| 0.087 | 0.0096 | -2.4 | 12.12 | A | A | A |
| 0.115 | 0.007 | 29.1 | 7.88 | A | N | W |
| 0.0862 | 0.0053 | -3.2 | 7.93 | A | A | A |
| 0.08 | 0.01 | -10.2 | 13.46 | A | A | A |
| 0.098 | 0.012 | 10.0 | 13.23 | A | A | A |
| 0.0939 | 0.0086 | 5.4 | 10.44 | A | A | A |
| 0.0907 | 0.0103 | 1.8 | 12.41 | A | A | A |
| 0.09 | 0.005 | 1.0 | 7.48 | A | A | A |
| 0.0764 | 0.00644 | -14.2 | 9.80 | A | A | A |
| 0.096 | 0.006 | 7.7 | 8.01 | A | A | A |
| 0.0847 | 0.0051 | -4.9 | 7.83 | A | A | A |
| 0.0814 | 0.0099 | -8.6 | 13.15 | A | A | A |
| 0.07946 | 0.024 | -10.8 | 30.62 | A | N | W |
| 0.1 | 0.02 | 12.2 | 20.62 | A | A | A |
| 0.104 | 0.01 | 16.7 | 10.84 | A | A | A |
| 0.098 | 0.008 | 10.0 | 9.58 | A | A | A |
| 0.089 | 0.009 | -0.1 | 11.28 | A | A | A |
| 0.1054 | 0.0581 | 18.3 | 55.35 | A | N | W |
| 0.0766 | 0.00509 | -14.0 | 8.32 | A | A | A |
| 0.0798 | 0.004 | -10.4 | 7.08 | A | A | A |
| 0.09 | 0.006 | 1.0 | 8.34 | A | A | A |
| 0.014 | 0.001 | -84.3 | 8.72 | N | N | N |
| 0.168 | 0.012 | 88.6 | 8.72 | N | N | N |
| 0.0862 | 0.0052 | -3.2 | 7.84 | A | A | A |
| 0.09 | 0.007 | 1.0 | 9.25 | A | A | A |
| 0.0868 | 0.0037 | -2.6 | 6.57 | A | A | A |
| 0.0605 | 0.0089 | -32.1 | 15.54 | N | A | N |
| 0.128 | 0.013 | 43.7 | 11.32 | N | N | N |

Labcode

19

25

28

31

32

38

41

42

49

52

55

58

60

63

65

66

71

73

81

87

90

113

114

128

132

133

152

156

159

161

161

163

166

170

175

181

187

191

192

204

205

207

208

213

217

219

220

230

231

255

264

268

276

278

283

285

286

287

291

303

| Reported value [Bq/kg ] 0.099 | Reported uncertainty (k = 1) [Bq/kg ] 0.006 | Relative bias [%] 11.1 | P-Test [%] 7.86 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 0.08 | 0.01 | -10.2 | 13.46 | A | A | A |
| 0.0886 | 0.00565 | -0.6 | 8.11 | A | A | A |
| 0.1 | 0.02 | 12.2 | 20.62 | A | A | A |
| 0.09105 | 0.02759 | 2.2 | 30.71 | A | N | W |
| 0.0797 | 0.0117 | -10.6 | 15.51 | A | A | A |
| 0.064 | 0.008 | -28.2 | 13.46 | A | A | A |
| 0.0782 | 0.008 | -12.2 | 11.39 | A | A | A |
| 0.0955 | 0.0089 | 7.2 | 10.58 | A | A | A |
| 0.081 | 0.021 | -9.1 | 26.40 | A | A | A |
| 0.081 | 0.004 | -9.1 | 7.03 | A | A | A |
| 0.063 | 0.02 | -29.3 | 32.14 | A | N | W |
| 0.083 | 0.011 | -6.8 | 14.17 | A | A | A |
| 0.11 | 0.01 | 23.5 | 10.38 | A | A | A |
| 0.074 | 0.005 | -17.0 | 8.41 | A | A | A |
| 0.435 | 0.0813 | 388.2 | 19.35 | N | N | N |
| 0.081 | 0.011 | -9.1 | 14.47 | A | A | A |
| 0.089 | 0.016 | -0.1 | 18.66 | A | A | A |
| 0.08 | 0.01 | -10.2 | 13.46 | A | A | A |
| 0.091 | 0.0055 | 2.1 | 7.85 | A | A | A |
| 0.167 | 0.028 | 87.4 | 17.50 | N | N | N |
| 0.067 | 0.007 | -24.8 | 11.58 | A | A | A |
| 0.084 | 0.005 | -5.7 | 7.78 | A | A | A |
| 0.18 | 0.01 | 102.0 | 7.48 | N | N | N |
| 0.096 | 0.012 | 7.7 | 13.46 | A | A | A |
| 0.11 | 0.007 | 23.5 | 8.10 | A | N | W |
| 0.103 | 0.009 | 15.6 | 10.07 | A | A | A |
| 0.1114 | 0.0105 | 25.0 | 10.67 | A | A | A |
| 0.0763 | 0.00652 | -14.4 | 9.90 | A | A | A |
| 0.128 | 0.045 | 43.7 | 35.51 | N | N | N |
| 0 | 0 | -100.0 |  | N |  | N |
| 0.089 | 0.0052 | -0.1 | 7.69 | A | A | A |
| 0.06 | 0.007 | -32.7 | 12.70 | N | A | N |
| 0.09 | 0.006 | 1.0 | 8.34 | A | A | A |
| 0.346 | 0.062 | 288.3 | 18.61 | N | N | N |
| 0.079 | 0.0024 | -11.3 | 5.86 | A | A | A |
| 79.1 | 3.6 | 88676.7 | 6.77 | N | N | N |
| 0.089 | 0.004 | -0.1 | 6.73 | A | A | A |
| 0.06 | 0.007 | -32.7 | 12.70 | N | A | N |
| 0.208 | 0.01 | 133.4 | 6.94 | N | N | N |
| 0.11 | 0.006 | 23.5 | 7.40 | A | N | W |
| 0.0849 | 0.0032 | -4.7 | 6.27 | A | A | A |
| 0.0931 | 0.0059 | 4.5 | 8.08 | A | A | A |
| 0.0904 | 0.0035 | 1.5 | 6.33 | A | A | A |
| 0.085 | 0.009 | -4.6 | 11.71 | A | A | A |
| 0.211 | 0.072 | 136.8 | 34.49 | N | N | N |
| 0.086 | 0.005 | -3.5 | 7.67 | A | A | A |
| 0.1058 | 0.0032 | 18.7 | 5.85 | A | N | W |
| 0.09 | 0.01 | 1.0 | 12.19 | A | A | A |

Labcode

304

308

311

312

313

315

330

340

346

347

348

355

368

370

371

377

379

392

393

399

401

402

410

412

414

419

423

427

432

433

434

437

441

443

445

461

463

475

483

492

508

516

525

527

528

533

535

536

537

Mass fraction  [ ng/g ]

0.05

determination of mass fraction in Sample 2 by ICP-MS

Table 24. Target values

Parameter	Value

Mass fraction [ ng/g ]	0.0470

Uncertainty [ ng/g ]	0.0017

Maximum Acceptable Relative Bias [%]	20

Accepted 46%	Not Accepted 10%	Warning 10%	Not Reported 34%

0.06

Figure 24. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

0.03

0.04

211

71

213

131

423

70

87

133

246

325

555

123

159

114

25

37

371

15

230

363

191

91

166

516

508

150

231

232

537

262

463

527

143

405

509

52

428

491

110

402

208

38

439

69

255

42

204

306

295

19

163

313

81

21

Laboratory code

Figure 25. Reported results and evaluation parameters for U-235. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 25. Evaluation results for U-235 in Sample 2, Rep. Value and Rep. Unc. in ng/g for mass fraction

| Reported value | Reported uncertainty | Relative bias | P-Test | Trueness | Precision | Final |
| --- | --- | --- | --- | --- | --- | --- |
| [ng/g ] | (k = 1) [ng/g ] | [%] | [%] | evaluation | evaluation | Score |
| 0.045 | 0.004 | -4.3 | 5.73 | A | A | A |
| 0.056 | 0.008 | 19.1 | 8.01 | A | A | A |
| 6.38 | 0.06 | 13474.5 | 3.65 | N | N | N |

Labcode

15

19

21

Table 25. Evaluation results for U-235 in Sample 2, Rep. Value and Rep. Unc. in ng/g for mass fraction

| Reported value [ng/g ] 0.044 | Reported uncertainty (k = 1) [ng/g ] 0.001 | Relative bias [%] -6.4 | P-Test [%] 3.79 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 0.044 | 0.013 | -6.4 | 15.21 | A | A | A |
| 0.049 | 0.006 | 4.3 | 7.11 | A | A | A |
| 0.0499 | 0.0032 | 6.2 | 4.83 | A | A | A |
| 0.047 | 0.0046 | 0.0 | 6.09 | A | A | A |
| 0.0494 | 0.0045 | 5.1 | 5.82 | A | A | A |
| 0.04 | 0.002 | -14.9 | 4.40 | A | N | W |
| 0.016 | 0.0006 | -66.0 | 4.07 | N | N | N |
| 0.475 | 0.069 | 910.6 | 8.11 | N | N | N |
| 0.04 | 0.006 | -14.9 | 8.33 | A | A | A |
| 0.0456 | 0.0019 | -3.0 | 4.17 | A | A | A |
| 0.0474 | 0.0004 | 0.8 | 3.64 | A | A | A |
| 0.0436 | 0.0044 | -7.2 | 6.21 | A | A | A |
| 0.0424 | 0.002 | -9.8 | 4.32 | A | A | A |
| 0.037 | 0.0019 | -21.3 | 4.44 | N | N | N |
| 0.04 | 0.02 | -14.9 | 25.26 | A | N | W |
| 0.0468 | 0.0047 | -0.4 | 6.19 | A | A | A |
| 0.046 | 0.03 | -2.1 | 32.81 | A | N | W |
| 0.04252 | 0.00319 | -9.5 | 5.21 | A | A | A |
| 0.056 | 0.004 | 19.1 | 5.08 | A | N | W |
| 0.0456 | 0.007 | -3.0 | 8.48 | A | A | A |
| 0.0455 | 0.008 | -3.2 | 9.51 | A | A | A |
| 0.05 | 0.005 | 6.4 | 6.17 | A | A | A |
| 0.0483 | 0.00104 | 2.8 | 3.77 | A | A | A |
| 0.015 | 0.002 | -68.1 | 7.58 | N | N | N |
| 0.0308 | 0.02 | -34.5 | 32.67 | N | N | N |
| 0.045 | 0.002 | -4.3 | 4.25 | A | A | A |
| 0.046 | 0.003 | -2.1 | 4.87 | A | A | A |
| 0.046 | 0.002 | -2.1 | 4.22 | A | A | A |
| 0.04 | 0.001 | -14.9 | 3.83 | A | N | W |
| 0.0496 | 0.0084 | 5.5 | 9.21 | A | A | A |
| 0.0462 | 0.0016 | -1.7 | 4.01 | A | A | A |
| 0.053 | 0.002 | 12.8 | 4.08 | A | N | W |
| 0.0524 | 0.00036 | 11.5 | 3.63 | A | N | W |
| 0.156862 | 0.018855 | 233.8 | 7.01 | N | N | N |
| 0.04 | 0.0014 | -14.9 | 4.02 | A | N | W |
| 0.0451 | 0.0094 | -4.0 | 11.03 | A | A | A |
| 0.044 | 0.006 | -6.4 | 7.72 | A | A | A |
| 0.0475 | 0.0006 | 1.1 | 3.67 | A | A | A |
| 0.0468 | 0.0031 | -0.4 | 4.90 | A | A | A |
| 0.0372 | 0.15 | -20.9 | 201.65 | N | N | N |
| 0.047 | 0.006 | 0.0 | 7.34 | A | A | A |
| 0.0493 | 0.0036 | 4.9 | 5.14 | A | A | A |
| 0.0463 | 0.0028 | -1.5 | 4.71 | A | A | A |
| 0.047 | 0.0005 | 0.0 | 3.66 | A | A | A |
| 0.04582 | 0.00102 | -2.5 | 3.78 | A | A | A |
| 0.04685 | 0.00015 | -0.3 | 3.62 | A | A | A |
| 0.0456 | 0.0047 | -3.0 | 6.30 | A | A | A |
| 0.0463 | 0.004 | -1.5 | 5.63 | A | A | A |
| 0.046 | 0.007 | -2.1 | 8.42 | A | A | A |
| 0.04 | 0.005 | -14.9 | 7.22 | A | A | A |

Labcode

25

37

38

42

52

69

70

71

81

87

91

110

114

123

131

133

143

150

159

163

166

191

204

208

211

213

230

231

232

246

255

262

295

306

313

325

363

371

402

405

423

428

439

463

491

508

509

516

527

537

555

Activity concentration  [ Bq/kg ]

0.005

U-235 determination of activity concentration in Sample 2

Table 26. Target values

Parameter	Value

Activity concentration [ Bq/kg ]	0.00372

Uncertainty [ Bq/kg ]	0.00031

Maximum Acceptable Relative Bias [%]	30

Accepted 9%	Not Accepted 20%	Warning 4%	Not Reported 67%

0.010

0.015

Figure 26. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

−0.005

0.000

434

213

58

291

49

536

207

220

308

441

483

287

377

60

87

15

65

187

527

25

114

166

264

159

63

90

191

528

170

231

475

535

276

285

311

52

268

208

423

69

255

42

81

152

286

315

443

32

437

516

230

371

205

163

132

346

414

175

204

401

192

304

312

11

433

533

31

73

412

410

211

402

71

217

213

131

423

70

87

133

246

325

555

419

123

159

114

25

37

371

15

230

363

191

91

166

516

508

150

231

232

537

262

463

527

143

405

509

52

428

491

110

402

208

38

439

69

255

42

204

306

295

19

163

313

71

81

113

218

491

463

80

21

Laboratory code

Figure 27. Reported results and evaluation parameters for U-235. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 27. Evaluation results for U-235 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Reported value | Reported uncertainty | Relative bias | P-Test | Trueness | Precision | Final |
| --- | --- | --- | --- | --- | --- | --- |
| [Bq/kg ] | (k = 1) [Bq/kg ] | [%] | [%] | evaluation | evaluation | Score |
| 0.00718 | 0.00243 | 93.0 | 34.85 | N | N | N |
| 0.0033 | 0.0014 | -11.3 | 43.23 | A | N | W |
| 0.0035 | 0.0001 | -5.9 | 8.81 | A | A | A |

Labcode

11

15

25

| Reported value [Bq/kg ] 0.00843 | Reported uncertainty (k = 1) [Bq/kg ] 0.00313 | Relative bias [%] 126.6 | P-Test [%] 38.05 | Trueness evaluation N | Precision evaluation N | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 0.0041 | 0.0019 | 10.2 | 47.08 | A | N | W |
| 0.00399 | 0.00013 | 7.3 | 8.95 | A | A | A |
| 0.0028 | 0.0005 | -24.7 | 19.71 | A | A | A |
| 0.0038 | 0.00019 | 2.1 | 9.72 | A | A | A |
| 0.0026 | 0.0004 | -30.1 | 17.50 | N | A | N |
| 0.00311 | 0.00077 | -16.4 | 26.12 | A | A | A |
| 0.0036 | 0.0005 | -3.2 | 16.20 | A | A | A |
| 0.0033 | 0.0015 | -11.3 | 46.21 | A | N | W |
| 0.00395 | 0.0002 | 6.2 | 9.75 | A | A | A |
| 0.16 | 0.008 | 4201.1 | 9.72 | N | N | N |
| 0.0089 | 0.0011 | 139.2 | 14.91 | N | N | N |
| 3.9058 | 0.25903 | 104894.6 | 10.65 | N | N | N |
| 0.004 | 0.002 | 7.5 | 50.69 | A | N | W |
| 0.0032 | 0.0003 | -14.0 | 12.54 | A | A | A |
| 0.0036 | 0.0004 | -3.2 | 13.89 | A | A | A |
| 0.7 | 0.077 | 18717.2 | 13.80 | N | N | N |
| 0.0035 | 0.0007 | -5.9 | 21.67 | A | A | A |
| 0.00476 | 0.00075 | 28.0 | 17.82 | A | A | A |
| 0.004 | 0.0008 | 7.5 | 21.67 | A | A | A |
| 0.003576 | 0.001135 | -3.9 | 32.82 | A | N | W |
| 0.0045 | 0.0002 | 21.0 | 9.44 | A | A | A |
| 0.0035 | 0.0012 | -5.9 | 35.28 | A | N | W |
| 0.0037 | 0.0009 | -0.5 | 25.71 | A | A | A |
| 0.005 | 0.0014 | 34.4 | 29.21 | N | A | N |
| 0.0033 | 0.0015 | -11.3 | 46.21 | A | N | W |
| 0.0036 | 0.0003 | -3.2 | 11.79 | A | A | A |
| 0.00526 | 0.00192 | 41.4 | 37.44 | N | N | N |
| 0.005 | 0.001 | 34.4 | 21.67 | N | A | N |
| 0.00433 | 0.00163 | 16.4 | 38.56 | A | N | W |
| 0.003 | 0.001 | -19.4 | 34.36 | A | N | W |
| 0.00387 | 0.0007 | 4.0 | 19.92 | A | A | A |
| 0.0025 | 0.0018 | -32.8 | 72.48 | N | N | N |
| 0.0227 | 0.01 | 510.2 | 44.83 | N | N | N |
| 0.82861 | 0.16254 | 22174.5 | 21.31 | N | N | N |
| 0.003 | 0.001 | -19.4 | 34.36 | A | N | W |
| 0.0043 | 0.0012 | 15.6 | 29.12 | A | A | A |
| 0.0037 | 0.0002 | -0.5 | 9.93 | A | A | A |
| 0.00397 | 0.00067 | 6.7 | 18.82 | A | A | A |
| 0.00353 | 0.000235 | -5.1 | 10.67 | A | A | A |
| 0.0038 | 0.0005 | 2.1 | 15.57 | A | A | A |
| 0.00373 | 0.00025 | 0.3 | 10.69 | A | A | A |
| 0.00373 | 0.00076 | 0.3 | 22.01 | A | A | A |
| 0.004 | 0.001 | 7.5 | 26.35 | A | A | A |
| 0.00301 | 0.00066 | -19.1 | 23.46 | A | A | A |
| 0.0027 | 0.0005 | -27.4 | 20.31 | A | A | A |
| 0.0053 | 0.0014 | 42.5 | 27.70 | N | A | N |
| 0.003 | 0.001 | -19.4 | 34.36 | A | N | W |
| 0.00373 | 0.00113 | 0.3 | 31.42 | A | N | W |
| 0.007 | 0.002 | 88.2 | 29.76 | N | N | N |
| 0.004 | 0.0028 | 7.5 | 70.49 | A | N | W |
| 0.0048 | 0.0014 | 29.0 | 30.33 | A | N | W |
| 0.0043 | 0.0007 | 15.6 | 18.29 | A | A | A |
| 0.00306 | 0.000645 | -17.7 | 22.67 | A | A | A |
| 0.005 | 0.001 | 34.4 | 21.67 | N | A | N |
| 0.015 | 0.003 | 303.2 | 21.67 | N | N | N |
| 0.013 | 0.002 | 249.5 | 17.50 | N | N | N |
| 0.01 | 0.001 | 168.8 | 13.02 | N | N | N |
| 0.00494 | 0.00073 | 32.8 | 16.97 | N | A | N |
| 0.041 | 0.004 | 1002.1 | 12.83 | N | N | N |

Labcode

31

32

42

49

52

58

60

63

65

69

71

73

80

81

87

90

113

114

132

152

159

163

166

170

175

187

191

192

204

205

207

208

213

217

218

220

230

231

255

264

268

276

285

286

287

291

304

308

311

312

315

346

371

377

401

402

410

412

414

419

| Reported value [Bq/kg ] 0.0039 | Reported uncertainty (k = 1) [Bq/kg ] 0.0004 | Relative bias [%] 4.8 | P-Test [%] 13.22 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 0.008 | 0.01 | 115.0 | 125.28 | N | N | N |
| 0 | 0 | -100.0 |  | N |  | N |
| 0.0041 | 0.00078 | 10.2 | 20.77 | A | A | A |
| 0.003 | 0.001 | -19.4 | 34.36 | A | N | W |
| 0.004 | 0.0007 | 7.5 | 19.38 | A | A | A |
| 3.55 | 0.47 | 95330.1 | 15.64 | N | N | N |
| 0.0037 | 0.0003 | -0.5 | 11.63 | A | A | A |
| 0.003 | 0.001 | -19.4 | 34.36 | A | N | W |
| 2.652 | 0.311 | 71190.3 | 14.39 | N | N | N |
| 0.00411 | 0.00065 | 10.5 | 17.88 | A | A | A |
| 0.0033 | 0.00065 | -11.3 | 21.39 | A | A | A |
| 0.0036 | 0.0008 | -3.2 | 23.73 | A | A | A |
| 0.008 | 0.004 | 115.0 | 50.69 | N | N | N |
| 0.0037 | 0.0009 | -0.5 | 25.71 | A | A | A |
| 0.00298 | 0.00052 | -19.9 | 19.34 | A | A | A |

Labcode

423

433

434

437

441

443

463

475

483

491

516

527

528

533

535

536

U-238 determination of activity concentration in Sample 2

Table 28. Target values

Parameter	Value

Activity concentration [ Bq/kg ]	0.0818

Uncertainty [ Bq/kg ]	0.0041

Maximum Acceptable Relative Bias [%]	30

Accepted 22%	Not Accepted 21%	Warning 2%	Not Reported 55%

0.3

Figure 28. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

−0.1

0.0

434

278

213

340

21

441

483

330

132

402

31

49

291

355

161

60

379

371

87

315

133

204

219

308

537

156

348

28

2

15

516

114

192

535

287

208

25

205

264

432

347

63

191

313

268

231

38

181

276

312

393

443

527

90

52

286

437

461

475

311

159

41

166

285

410

528

525

152

175

69

81

42

128

414

55

19

65

187

304

368

32

161

58

370

412

255

207

163

423

230

419

346

66

220

536

392

303

427

399

170

401

508

217

73

11

433

147

350

283

492

533

445

377

491

71

213

71

423

131

313

70

555

87

325

133

123

307

114

159

37

256

25

246

15

363

228

371

527

230

441

483

516

38

191

463

91

537

166

150

143

232

405

52

491

231

509

110

402

508

428

304

315

58

69

439

42

128

208

204

262

306

255

295

163

19

81

113

463

Laboratory code

Figure 29. Reported results and evaluation parameters for U-238. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 29. Evaluation results for U-238 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

Labcode

Reported value

[Bq/kg ]

Reported uncertainty

(k = 1) [Bq/kg ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

Activity concentration  [ Bq/kg ]

0.1

0.2

2	0.0728	0.0095	-11.0	13.98	A	A	A

11	0.115	0.00938	40.6	9.57	N	N	N

15	0.074	0.01	-9.5	14.41	A	A	A

| Reported value [Bq/kg ] 0.088 | Reported uncertainty (k = 1) [Bq/kg ] 0.005 | Relative bias [%] 7.6 | P-Test [%] 7.58 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 0.076 | 0.002 | -7.1 | 5.66 | A | A | A |
| 0.0725 | 0.0143 | -11.4 | 20.35 | A | A | A |
| 0.0615 | 0.0097 | -24.8 | 16.55 | A | A | A |
| 0.089 | 0.008 | 8.8 | 10.29 | A | A | A |
| 0.08 | 0.0095 | -2.2 | 12.89 | A | A | A |
| 0.0818 | 0.00313 | 0.0 | 6.31 | A | A | A |
| 0.0858 | 0.0035 | 4.9 | 6.46 | A | A | A |
| 0.0625 | 0.0094 | -23.6 | 15.85 | A | A | A |
| 0.081 | 0.004 | -1.0 | 7.04 | A | A | A |
| 0.0871 | 0.015 | 6.5 | 17.94 | A | A | A |
| 0.09 | 0.003 | 10.0 | 6.02 | A | A | A |
| 0.0655 | 0.0093 | -19.9 | 15.06 | A | A | A |
| 0.0785 | 0.0033 | -4.0 | 6.54 | A | A | A |
| 0.088 | 0.011 | 7.6 | 13.47 | A | A | A |
| 0.0944 | 0.00667 | 15.4 | 8.66 | A | A | A |
| 0.085 | 0.0039 | 3.9 | 6.80 | A | A | A |
| 4.47 | 0.22 | 5364.6 | 7.02 | N | N | N |
| 0.112 | 0.007 | 36.9 | 8.01 | N | N | N |
| 0.085 | 0.007 | 3.9 | 9.64 | A | A | A |
| 0.0686 | 0.0069 | -16.1 | 11.24 | A | A | A |
| 0.0808 | 0.0049 | -1.2 | 7.87 | A | A | A |
| 16 | 1.76 | 19459.9 | 12.09 | N | N | N |
| 0.075 | 0.011 | -8.3 | 15.50 | A | A | A |
| 0.086 | 0.0089 | 5.1 | 11.50 | A | A | A |
| 0.061 | 0.0053 | -25.4 | 10.03 | A | A | A |
| 0.07 | 0.01 | -14.4 | 15.14 | A | A | A |
| 0.1329 | 0.0165 | 62.5 | 13.39 | N | N | N |
| 0.0834 | 0.006 | 2.0 | 8.77 | A | A | A |
| 0.0716 | 0.0031 | -12.5 | 6.62 | A | A | A |
| 0.08132 | 0.004265 | -0.6 | 7.25 | A | A | A |
| 0.089 | 0.0095 | 8.8 | 11.79 | A | A | A |
| 0.064 | 0.011 | -21.8 | 17.90 | A | A | A |
| 0.092 | 0.03 | 12.5 | 32.99 | A | N | W |
| 0.082 | 0.009 | 0.2 | 12.07 | A | A | A |
| 0.103 | 0.006 | 25.9 | 7.68 | A | N | W |
| 0.0849 | 0.0052 | 3.8 | 7.91 | A | A | A |
| 0.08 | 0.01 | -2.2 | 13.47 | A | A | A |
| 0.088 | 0.011 | 7.6 | 13.47 | A | A | A |
| 0.0789 | 0.0054 | -3.5 | 8.48 | A | A | A |
| 0.075 | 0.00893 | -8.3 | 12.92 | A | A | A |
| 0.07 | 0.005 | -14.4 | 8.73 | A | A | A |
| 0.0761 | 0.00643 | -7.0 | 9.82 | A | A | A |
| 0.091 | 0.006 | 11.2 | 8.28 | A | A | A |
| 0.0756 | 0.0046 | -7.6 | 7.88 | A | A | A |
| 0.0444 | 0.007 | -45.7 | 16.54 | N | N | N |
| 0.11068 | 0.023 | 35.3 | 21.38 | N | A | N |
| 0.07 | 0.02 | -14.4 | 29.01 | A | A | A |
| 0.096 | 0.009 | 17.4 | 10.63 | A | A | A |
| 0.093 | 0.008 | 13.7 | 9.96 | A | A | A |
| 0.0799 | 0.0016 | -2.3 | 5.40 | A | A | A |
| 0.0901 | 0.0069 | 10.2 | 9.15 | A | A | A |
| 0.0766 | 0.00509 | -6.4 | 8.32 | A | A | A |
| 0.0798 | 0.004 | -2.4 | 7.09 | A | A | A |
| 0.08 | 0.005 | -2.2 | 8.01 | A | A | A |
| 0.011 | 0.001 | -86.6 | 10.38 | N | N | N |
| 0.161 | 0.012 | 96.8 | 8.98 | N | N | N |
| 0.083 | 0.0051 | 1.5 | 7.93 | A | A | A |
| 0.081 | 0.007 | -1.0 | 9.99 | A | A | A |
| 0.0755 | 0.0034 | -7.7 | 6.74 | A | A | A |

Labcode

19

25

28

31

32

38

41

42

49

52

55

58

60

63

65

66

69

71

73

81

87

90

113

114

128

132

133

147

152

156

159

161

161

163

166

170

175

181

187

191

192

204

205

207

208

213

217

219

220

230

231

255

264

268

276

278

283

285

286

287

| Reported value [Bq/kg ] 0.0626 | Reported uncertainty (k = 1) [Bq/kg ] 0.0094 | Relative bias [%] -23.5 | P-Test [%] 15.83 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 0.098 | 0.01 | 19.8 | 11.37 | A | A | A |
| 0.088 | 0.006 | 7.6 | 8.46 | A | A | A |
| 0.07 | 0.01 | -14.4 | 15.14 | A | A | A |
| 0.0812 | 0.0055 | -0.7 | 8.43 | A | A | A |
| 0.08 | 0.02 | -2.2 | 25.50 | A | A | A |
| 0.079 | 0.02378 | -3.4 | 30.52 | A | N | W |
| 0.0696 | 0.0109 | -14.9 | 16.44 | A | A | A |
| 0.06 | 0.007 | -26.6 | 12.70 | A | A | A |
| 0.0456 | 0.0058 | -44.2 | 13.67 | N | N | N |
| 0.0941 | 0.0088 | 15.0 | 10.61 | A | A | A |
| 0.078 | 0.019 | -4.7 | 24.87 | A | A | A |
| 0.072 | 0.004 | -12.0 | 7.48 | A | A | A |
| 0.15 | 0.02 | 83.4 | 14.24 | N | N | N |
| 0.063 | 0.02 | -23.0 | 32.14 | A | N | W |
| 0.088 | 0.011 | 7.6 | 13.47 | A | A | A |
| 0.09 | 0.01 | 10.0 | 12.19 | A | A | A |
| 0.068 | 0.005 | -16.9 | 8.90 | A | A | A |
| 0.435 | 0.0813 | 431.8 | 19.35 | N | N | N |
| 0.067 | 0.009 | -18.1 | 14.34 | A | A | A |
| 0.097 | 0.016 | 18.6 | 17.24 | A | A | A |
| 0.08 | 0.01 | -2.2 | 13.47 | A | A | A |
| 0.1 | 0.006 | 22.2 | 7.82 | A | N | W |
| 0.105 | 0.022 | 28.4 | 21.54 | A | A | A |
| 0.061 | 0.006 | -25.4 | 11.04 | A | A | A |
| 0.083 | 0.005 | 1.5 | 7.84 | A | A | A |
| 0.09 | 0.01 | 10.0 | 12.19 | A | A | A |
| 0.087 | 0.011 | 6.4 | 13.60 | A | A | A |
| 0.094 | 0.006 | 14.9 | 8.12 | A | A | A |
| 0.092 | 0.009 | 12.5 | 10.99 | A | A | A |
| 0.0992 | 0.0009 | 21.3 | 5.09 | A | N | W |
| 0.0778 | 0.00775 | -4.9 | 11.15 | A | A | A |
| 0.116 | 0.04 | 41.8 | 34.85 | N | N | N |
| 0 | 0 | -100.0 |  | N |  | N |
| 0.081 | 0.0048 | -1.0 | 7.76 | A | A | A |
| 0.049 | 0.006 | -40.1 | 13.23 | N | N | N |
| 0.08 | 0.005 | -2.2 | 8.01 | A | A | A |
| 0.299 | 0.054 | 265.5 | 18.74 | N | N | N |
| 0.081 | 0.0026 | -1.0 | 5.95 | A | A | A |
| 75.6 | 3.6 | 92320.5 | 6.91 | N | N | N |
| 0.081 | 0.004 | -1.0 | 7.04 | A | A | A |
| 0.049 | 0.006 | -40.1 | 13.23 | N | N | N |
| 2.129 | 0.567 | 2502.7 | 27.10 | N | N | N |
| 0.18 | 0.01 | 120.0 | 7.48 | N | N | N |
| 0.105 | 0.005 | 28.4 | 6.91 | A | N | W |
| 0.0748 | 0.003 | -8.6 | 6.42 | A | A | A |
| 0.0831 | 0.0056 | 1.6 | 8.40 | A | A | A |
| 0.0802 | 0.0032 | -2.0 | 6.41 | A | A | A |
| 0.083 | 0.009 | 1.5 | 11.95 | A | A | A |
| 0.185 | 0.042 | 126.2 | 23.25 | N | N | N |
| 0.075 | 0.006 | -8.3 | 9.44 | A | A | A |
| 0.0966 | 0.003 | 18.1 | 5.90 | A | N | W |
| 0.07 | 0.01 | -14.4 | 15.14 | A | A | A |

Labcode

291

303

304

308

311

312

313

315

330

340

346

347

348

350

355

368

370

371

377

379

392

393

399

401

402

410

412

414

419

423

427

432

433

434

437

441

443

445

461

463

475

483

491

492

508

516

525

527

528

533

535

536

537

U-238 determination of mass fraction in Sample 2 by ICP-MS

Table 30. Target values

| Parameter | Value |
| --- | --- |
| Mass fraction [ ng/g ] | 6.56 |
| Uncertainty [ ng/g ] | 0.21 |
| Maximum Acceptable Relative Bias [%] | 20 |

Accepted 55%	Not Accepted 6%	Warning 13%	Not Reported 26%

8

9

Figure 30. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

4

5

6

21

71

213

423

131

313

70

555

87

325

133

123

307

114

159

37

256

25

246

15

363

228

371

527

230

441

483

516

38

191

463

91

537

166

150

143

232

405

52

491

231

509

110

402

508

428

304

315

58

69

439

42

128

208

204

262

306

255

295

163

19

81

Laboratory code

Figure 31. Reported results and evaluation parameters for U-238. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 31. Evaluation results for U-238 in Sample 2, Rep. Value and Rep. Unc. in ng/g for mass fraction

Labcode

Reported value

[ng/g ]

Reported uncertainty

(k = 1) [ng/g ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

Mass fraction  [ ng/g ]

7

15	6.18	0.48	-5.8	5.03	A	A	A

19	7.61	0.047	16.0	3.22	A	N	W

21	0.04662	0.0008	-99.3	3.31	N	N	N

| Reported value [ng/g ] 6.12 | Reported uncertainty (k = 1) [ng/g ] 0.11 | Relative bias [%] -6.7 | P-Test [%] 3.32 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 6.08 | 1.8 | -7.3 | 15.14 | A | A | A |
| 6.33 | 0.673 | -3.5 | 6.21 | A | A | A |
| 6.9 | 0.558 | 5.2 | 5.16 | A | A | A |
| 6.49 | 0.64 | -1.1 | 5.88 | A | A | A |
| 6.8 | 0.6 | 3.7 | 5.45 | A | A | A |
| 6.83 | 0.62 | 4.1 | 5.55 | A | A | A |
| 5.45 | 0.23 | -16.9 | 3.83 | A | N | W |
| 2.238 | 0.068 | -65.9 | 3.54 | N | N | N |
| 9.6 | 1.3 | 46.3 | 7.49 | N | N | N |
| 5.52 | 0.828 | -15.8 | 8.15 | A | A | A |
| 6.3604 | 0.1526 | -3.0 | 3.42 | A | A | A |
| 6.57 | 0.1 | 0.1 | 3.29 | A | A | A |
| 6.01 | 0.36 | -8.4 | 4.38 | A | A | A |
| 5.86 | 0.29 | -10.7 | 4.05 | A | N | W |
| 6.95 | 0.71 | 6.0 | 6.03 | A | A | A |
| 5.26 | 0.203 | -19.8 | 3.74 | A | N | W |
| 5.66 | 0.12 | -13.7 | 3.37 | A | N | W |
| 6.444 | 0.644 | -1.8 | 5.93 | A | A | A |
| 6.4 | 0.3 | -2.4 | 3.97 | A | A | A |
| 6.0548 | 0.4541 | -7.7 | 4.93 | A | A | A |
| 7.36 | 0.47 | 12.2 | 4.52 | A | N | W |
| 6.38 | 0.95 | -2.7 | 8.10 | A | A | A |
| 6.336 | 0.862 | -3.4 | 7.52 | A | A | A |
| 7.1 | 0.7 | 8.2 | 5.88 | A | A | A |
| 7.0687 | 0.1925 | 7.8 | 3.48 | A | A | A |
| 3.57 | 0.563 | -45.6 | 8.51 | N | N | N |
| 6.2 | 2.2 | -5.5 | 18.03 | A | A | A |
| 6.23 | 0.21 | -5.0 | 3.62 | A | A | A |
| 6.5 | 0.4 | -0.9 | 4.44 | A | A | A |
| 6.45 | 0.26 | -1.7 | 3.78 | A | A | A |
| 6.12 | 0.17 | -6.7 | 3.49 | A | A | A |
| 7.28 | 0.39 | 11.0 | 4.17 | A | N | W |
| 6.1 | 1.8 | -7.0 | 15.10 | A | A | A |
| 7.14 | 0.2332 | 8.8 | 3.59 | A | A | A |
| 7.358 | 0.25 | 12.2 | 3.62 | A | N | W |
| 6.7 | 0.5 | 2.1 | 4.92 | A | A | A |
| 7.21 | 0.86 | 9.9 | 6.77 | A | A | A |
| 5.9 | 0.6 | -10.1 | 6.01 | A | A | A |
| 5.294478 | 0.636396 | -19.3 | 6.81 | A | N | W |
| 6.754 | 1.351 | 3.0 | 10.50 | A | A | A |
| 5.56 | 0.19 | -15.2 | 3.63 | A | N | W |
| 6.18 | 0.52 | -5.8 | 5.29 | A | A | A |
| 6.22 | 0.44 | -5.2 | 4.77 | A | A | A |
| 6.632 | 0.082 | 1.1 | 3.26 | A | A | A |
| 6.46 | 0.43 | -1.5 | 4.62 | A | A | A |
| 4.806 | 0.1 | -26.7 | 3.37 | N | N | N |
| 6.69 | 0.46 | 2.0 | 4.70 | A | A | A |
| 6.88 | 0.49 | 4.9 | 4.79 | A | A | A |
| 6.265 | 1.002 | -4.5 | 8.61 | A | A | A |
| 6.35 | 0.37 | -3.2 | 4.33 | A | A | A |
| 6.265 | 1.002 | -4.5 | 8.61 | A | A | A |
| 6.491 | 0.0664 | -1.0 | 3.24 | A | A | A |
| 6.64 | 0.54 | 1.2 | 5.18 | A | A | A |
| 6.538 | 0.021 | -0.3 | 3.21 | A | A | A |
| 6.29 | 0.6 | -4.1 | 5.74 | A | A | A |
| 6.22 | 0.5 | -5.2 | 5.14 | A | A | A |
| 6.374 | 1.021 | -2.8 | 8.63 | A | A | A |
| 5.46 | 0.5 | -16.8 | 5.59 | A | N | W |

Table 31. Evaluation results for U-238 in Sample 2, Rep. Value and Rep. Unc. in ng/g for mass fraction

Labcode

25

37

38

42

52

58

69

70

71

81

87

91

110

114

123

128

131

133

143

150

159

163

166

191

204

208

213

228

230

231

232

246

255

256

262

295

304

306

307

313

315

325

363

371

402

405

423

428

439

441

463

483

491

508

509

516

527

537

555

Total U determination of mass fraction in Sample 2 by ICP-MS

Table 32. Target values

| Parameter | Value |
| --- | --- |
| Mass fraction [ ng/g ] | 6.60 |
| Uncertainty [ ng/g ] | 0.21 |
| Maximum Acceptable Relative Bias [%] | 20 |

Accepted 63%	Not Accepted 10%	Warning 16%	Not Reported 11%

8

9

Figure 32. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

4

5

6

71

213

330

423

131

470

70

555

87

180

313

133

299

235

123

114

238

466

159

25

246

171

363

15

527

371

230

239

516

191

38

463

150

91

537

166

53

232

143

21

231

405

491

52

37

509

110

40

402

508

242

428

454

528

58

145

69

42

128

208

262

204

306

255

295

81

163

170

19

288

83

12

492

Laboratory code

Figure 33. Reported results and evaluation parameters for Total U. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 33. Evaluation Results for Total U in Sample 2, Rep. Value and Rep. Unc. in ng/g for mass fraction

Labcode

Reported value

[ng/g ]

Reported uncertainty

(k = 1) [ng/g ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

Mass fraction  [ ng/g ]

7

12	13.56	0.14	105.4	3.22	N	N	N

15	6.23	0.5	-5.6	5.12	A	A	A

19	7.67	0.064	16.2	3.21	A	N	W

| Reported value [ng/g ] 6.5 | Reported uncertainty (k = 1) [ng/g ] 0.3 | Relative bias [%] -1.5 | P-Test [%] 3.93 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 6.16 | 0.11 | -6.7 | 3.30 | A | A | A |
| 6.56 | 1.4 | -0.6 | 11.14 | A | A | A |
| 6.38 | 0.68 | -3.3 | 6.21 | A | A | A |
| 6.63 | 0.86 | 0.4 | 7.22 | A | A | A |
| 6.95 | 0.558 | 5.3 | 5.12 | A | A | A |
| 6.54 | 0.94 | -0.9 | 7.86 | A | A | A |
| 6.47 | 1.29 | -2.0 | 10.46 | A | A | A |
| 6.8 | 0.6 | 3.0 | 5.44 | A | A | A |
| 6.88 | 0.92 | 4.2 | 7.40 | A | A | A |
| 5.49 | 0.23 | -16.8 | 3.81 | A | N | W |
| 2.254 | 0.0068 | -65.8 | 3.19 | N | N | N |
| 7.415 | 1.31 | 12.3 | 9.39 | A | A | A |
| 11.9 | 2.3 | 80.3 | 10.17 | N | N | N |
| 5.56 | 0.834 | -15.8 | 8.15 | A | A | A |
| 6.406 | 0.1537 | -2.9 | 3.40 | A | A | A |
| 6.62 | 0.1 | 0.3 | 3.27 | A | A | A |
| 6.05 | 0.36 | -8.3 | 4.36 | A | A | A |
| 5.9027 | 0.29 | -10.6 | 4.02 | A | N | W |
| 6.95 | 0.71 | 5.3 | 6.02 | A | A | A |
| 5.3 | 0.2 | -19.7 | 3.70 | A | N | W |
| 5.71 | 0.12 | -13.5 | 3.35 | A | N | W |
| 6.491 | 0.649 | -1.6 | 5.93 | A | A | A |
| 6.81 | 1.7 | 3.2 | 12.88 | A | A | A |
| 6.4 | 0.3 | -3.0 | 3.95 | A | A | A |
| 6.0984 | 0.4574 | -7.6 | 4.92 | A | A | A |
| 7.42 | 0.47 | 12.4 | 4.49 | A | N | W |
| 6.43 | 0.96 | -2.6 | 8.11 | A | A | A |
| 7.6 | 0.4 | 15.2 | 4.13 | A | N | W |
| 6.17 | 0.11 | -6.5 | 3.30 | A | A | A |
| 5.6 | 0.112 | -15.2 | 3.34 | A | N | W |
| 6.35 | 0.582 | -3.8 | 5.58 | A | A | A |
| 7.2 | 0.7 | 9.1 | 5.81 | A | A | A |
| 7.1174 | 0.1932 | 7.8 | 3.46 | A | A | A |
| 3.6 | 0.585 | -45.5 | 8.73 | N | N | N |
| 6.27 | 0.21 | -5.0 | 3.60 | A | A | A |
| 6.5 | 0.3 | -1.5 | 3.93 | A | A | A |
| 6.49 | 0.27 | -1.7 | 3.80 | A | A | A |
| 5.805 | 0.639 | -12.1 | 6.36 | A | A | A |
| 6.09 | 0.858 | -7.7 | 7.73 | A | A | A |
| 6.3 | 0.2 | -4.6 | 3.56 | A | A | A |
| 6.691 | 0.052 | 1.4 | 3.21 | A | A | A |
| 6.16 | 0.17 | -6.7 | 3.47 | A | A | A |
| 7.33 | 0.39 | 11.1 | 4.15 | A | N | W |
| 7.19 | 0.172 | 8.9 | 3.40 | A | N | W |
| 9.2 | 1.84 | 39.4 | 10.49 | N | N | N |
| 7.411 | 0.251 | 12.3 | 3.60 | A | N | W |
| 5.8 | 1.2 | -12.1 | 10.82 | A | A | A |
| 7.26 | 0.94 | 10.0 | 7.21 | A | A | A |
| 5.690275 | 0.000684 | -13.8 | 3.18 | A | N | W |
| 4.57 | 0.61 | -30.8 | 7.39 | N | N | N |
| 6.22 | 0.47 | -5.8 | 4.94 | A | A | A |
| 6.26 | 0.4 | -5.2 | 4.51 | A | A | A |
| 6.68 | 0.083 | 1.2 | 3.24 | A | A | A |
| 6.51 | 0.44 | -1.4 | 4.64 | A | A | A |
| 4.842 | 0.1 | -26.6 | 3.35 | N | N | N |
| 6.74 | 0.46 | 2.1 | 4.67 | A | A | A |
| 6.78 | 0.92 | 2.7 | 7.49 | A | A | A |
| 6.396 | 0.37 | -3.1 | 4.30 | A | A | A |
| 6.09 | 0.307545 | -7.7 | 4.06 | A | A | A |

Labcode

21

25

37

38

40

42

52

53

58

69

70

71

81

83

87

91

110

114

123

128

131

133

143

145

150

159

163

166

170

171

180

191

204

208

213

230

231

232

235

238

239

242

246

255

262

288

295

299

306

313

330

363

371

402

405

423

428

454

463

466

| Reported value [ng/g ] 5.3 | Reported uncertainty (k = 1) [ng/g ] 1 | Relative bias [%] -19.7 | P-Test [%] 9.96 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 6.538 | 0.067 | -0.9 | 3.22 | A | A | A |
| 14.55 | 0.7 | 120.4 | 3.99 | N | N | N |
| 6.68 | 0.54 | 1.2 | 5.14 | A | A | A |
| 6.586 | 0.021 | -0.2 | 3.19 | A | A | A |
| 6.31 | 0.6 | -4.4 | 5.72 | A | A | A |
| 6.24 | 0.5 | -5.4 | 5.12 | A | A | A |
| 6.78 | 1.36 | 2.7 | 10.52 | A | A | A |
| 6.42 | 1.023 | -2.7 | 8.58 | A | A | A |
| 5.5 | 0.5 | -16.7 | 5.55 | A | N | W |

Labcode

470

491

492

508

509

516

527

528

537

555

K-40 determination of activity concentration in Sample 3

Table 34. Target values

| Parameter | Value |
| --- | --- |
| Activity concentration [ Bq/kg ] | 485 |
| Uncertainty [ Bq/kg ] | 11 |
| Maximum Acceptable Relative Bias [%] | 20 |

Accepted 65%	Not Accepted 9%	Warning 6%	Not Reported 20%

600

Figure 34. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

300

400

37

94

303

163

499

38

535

206

346

66

211

84

350

395

72

134

230

185

293

386

34

218

294

204

453

238

151

121

50

189

200

452

555

40

408

514

178

487

159

475

192

450

116

215

13

219

306

175

272

242

122

194

435

465

52

325

393

232

119

152

220

165

166

347

53

270

102

479

492

319

506

67

549

105

17

257

107

398

478

111

2

307

265

420

544

384

516

197

311

60

8

170

354

443

344

68

127

65

202

329

402

24

484

441

356

387

124

407

260

198

328

527

364

213

158

191

292

463

286

258

256

226

227

21

302

278

71

532

376

297

480

339

318

32

279

371

412

31

460

429

20

214

367

501

273

377

133

533

262

85

326

462

520

455

495

323

491

333

Laboratory code

Figure 35. Reported results and evaluation parameters for K-40. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 35. Evaluation results for K-40 in Sample 3, Rep. Value and Rep. Unc. in Bq/kg

Labcode

Reported value

[Bq/kg ]

Reported uncertainty

(k = 1) [Bq/kg ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

Activity concentration  [ Bq/kg ]

500

2	492	56	1.4	11.61	A	A	A

8	499.64	30.79	3.0	6.57	A	A	A

9	464.3	55.7	-4.3	12.21	A	A	A

| Reported value [Bq/kg ] 502 | Reported uncertainty (k = 1) [Bq/kg ] 34.6 | Relative bias [%] 3.5 | P-Test [%] 7.26 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 409.12 | 44.16 | -15.7 | 11.03 | A | A | A |
| 470.1 | 19.1 | -3.1 | 4.65 | A | A | A |
| 479 | 30 | -1.2 | 6.66 | A | A | A |
| 489.59 | 22.01 | 0.9 | 5.04 | A | A | A |
| 529 | 42 | 9.1 | 8.26 | A | A | A |
| 562.5 | 10.7 | 16.0 | 2.96 | A | N | W |
| 527 | 12 | 8.7 | 3.21 | A | N | W |
| 507 | 21.7 | 4.5 | 4.84 | A | A | A |
| 469 | 19 | -3.3 | 4.64 | A | A | A |
| 516.3 | 45.35 | 6.4 | 9.07 | A | A | A |
| 548 | 19.9 | 13.0 | 4.28 | A | N | W |
| 537.4 | 21.1 | 10.8 | 4.53 | A | A | A |
| 447 | 27 | -7.8 | 6.45 | A | A | A |
| 490 | 39 | 1.0 | 8.28 | A | A | A |
| 2.61 | 0.31 | -99.5 | 12.09 | N | N | N |
| 409.32 | 37.75 | -15.6 | 9.50 | A | A | A |
| 462 | 31.84 | -4.7 | 7.26 | A | A | A |
| 447.05 | 21.1 | -7.8 | 5.24 | A | A | A |
| 492 | 30 | 1.4 | 6.51 | A | A | A |
| 475.9 | 22.2 | -1.9 | 5.19 | A | A | A |
| 490.5 | 18.5 | 1.1 | 4.40 | A | A | A |
| 500 | 60 | 3.1 | 12.21 | A | A | A |
| 458 | 15.2 | -5.6 | 4.02 | A | A | A |
| 477 | 20.1 | -1.6 | 4.79 | A | A | A |
| 482 | 30 | -0.6 | 6.62 | A | A | A |
| 527 | 26 | 8.7 | 5.43 | A | A | A |
| 507.35 | 21.83 | 4.6 | 4.86 | A | A | A |
| 522.2 | 52.2 | 7.7 | 10.25 | A | A | A |
| 521 | 26 | 7.4 | 5.48 | A | A | A |
| 498 | 58 | 2.7 | 11.87 | A | A | A |
| 430 | 35 | -11.3 | 8.45 | A | A | A |
| 527.81 | 12.45 | 8.8 | 3.27 | A | N | W |
| 485.8 | 43.6 | 0.2 | 9.26 | A | A | A |
| 504.4 | 25 | 4.0 | 5.45 | A | A | A |
| 424.7 | 23.045 | -12.4 | 5.88 | A | A | A |
| 486.8 | 60.5 | 0.4 | 12.63 | A | A | A |
| 502 | 20.1 | 3.5 | 4.60 | A | A | A |
| 565 | 26 | 16.5 | 5.13 | A | N | W |
| 480 | 16 | -1.0 | 4.03 | A | A | A |
| 528.18 | 26.41 | 8.9 | 5.49 | A | A | A |
| 441 | 22.2 | -9.1 | 5.52 | A | A | A |
| 578.8 | 4.85 | 19.3 | 2.42 | A | N | W |
| 473 | 11 | -2.5 | 3.25 | A | A | A |
| 421.75 | 47.25 | -13.0 | 11.43 | A | A | A |
| 764.11 | 16.48 | 57.5 | 3.13 | N | N | N |
| 497.2 | 72 | 2.5 | 14.66 | A | A | A |
| 430 | 33 | -11.3 | 8.00 | A | A | A |
| 626.094 | 40.482 | 29.1 | 6.85 | N | N | N |
| 518 | 26 | 6.8 | 5.51 | A | A | A |
| 272.5 | 19.8 | -43.8 | 7.61 | N | N | N |
| 467 | 51 | -3.7 | 11.15 | A | A | A |
| 475 | 24 | -2.1 | 5.54 | A | A | A |
| 482.2 | 43 | -0.6 | 9.20 | A | A | A |
| 470.7 | 17 | -3.0 | 4.26 | A | A | A |
| 488.8 | 16.1 | 0.8 | 4.00 | A | A | A |
| 490 | 28 | 1.0 | 6.15 | A | A | A |
| 515 | 4 | 6.2 | 2.40 | A | N | W |
| 486 | 29 | 0.2 | 6.38 | A | A | A |
| 491 | 30 | 1.2 | 6.52 | A | A | A |

Labcode

11

12

13

15

17

19

20

21

24

25

28

31

32

34

35

37

38

40

41

42

45

48

49

50

52

53

54

55

57

58

60

62

63

64

65

66

67

68

69

70

71

72

78

79

81

82

83

84

85

90

94

97

100

102

104

105

107

109

110

111

| Reported value [Bq/kg ] 573.4 | Reported uncertainty (k = 1) [Bq/kg ] 86 | Relative bias [%] 18.2 | P-Test [%] 15.17 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 455 | 31 | -6.2 | 7.18 | A | A | A |
| 469.62 | 23.49 | -3.2 | 5.49 | A | A | A |
| 480 | 32 | -1.0 | 7.04 | A | A | A |
| 457 | 35 | -5.8 | 7.99 | A | A | A |
| 475 | 22 | -2.1 | 5.16 | A | A | A |
| 509 | 27 | 5.0 | 5.77 | A | A | A |
| 480 | 31.52 | -1.0 | 6.95 | A | A | A |
| 504 | 15 | 3.9 | 3.74 | A | A | A |
| 509.59 | 19.85 | 5.1 | 4.51 | A | A | A |
| 457.2827 | 13.67027 | -5.7 | 3.75 | A | A | A |
| 471 | 50 | -2.9 | 10.86 | A | A | A |
| 604 | 57 | 24.5 | 9.71 | N | A | N |
| 441.3 | 3.7 | -9.0 | 2.42 | A | N | W |
| 501 | 21.4 | 3.3 | 4.84 | A | A | A |
| 463 | 23 | -4.5 | 5.46 | A | A | A |
| 426.8 | 68.29 | -12.0 | 16.16 | A | A | A |
| 611.755 | 24.943 | 26.1 | 4.67 | N | N | N |
| 497 | 25 | 2.5 | 5.52 | A | A | A |
| 455.19 | 31.85 | -6.2 | 7.36 | A | A | A |
| 480 | 29 | -1.0 | 6.45 | A | A | A |
| 470 | 27 | -3.1 | 6.18 | A | A | A |
| 451 | 27 | -7.0 | 6.40 | A | A | A |
| 475 | 6.3 | -2.1 | 2.63 | A | A | A |
| 519 | 73 | 7.0 | 14.25 | A | A | A |
| 467 | 21.5 | -3.7 | 5.13 | A | A | A |
| 519 | 33 | 7.0 | 6.75 | A | A | A |
| 387 | 12 | -20.2 | 3.84 | N | N | N |
| 480.56 | 20.26 | -0.9 | 4.79 | A | A | A |
| 481 | 24 | -0.8 | 5.48 | A | A | A |
| 482 | 43 | -0.6 | 9.20 | A | A | A |
| 500 | 40 | 3.1 | 8.32 | A | A | A |
| 540 | 28 | 11.3 | 5.66 | A | A | A |
| 472 | 21 | -2.7 | 4.99 | A | A | A |
| 465 | 34.1 | -4.1 | 7.68 | A | A | A |
| 505.03 | 0.59 | 4.1 | 2.27 | A | A | A |
| 458 | 19.9 | -5.6 | 4.90 | A | A | A |
| 569.77 | 40.27 | 17.5 | 7.42 | A | A | A |
| 498 | 30 | 2.7 | 6.44 | A | A | A |
| 460 | 6.5 | -5.2 | 2.67 | A | A | A |
| 443.27 | 8.49 | -8.6 | 2.97 | A | N | W |
| 504.4 | 25 | 4.0 | 5.45 | A | A | A |
| 459 | 23 | -5.4 | 5.50 | A | A | A |
| 519 | 32 | 7.0 | 6.57 | A | A | A |
| 468 | 45 | -3.5 | 9.88 | A | A | A |
| 824.28 | 25.2 | 70.0 | 3.81 | N | N | N |
| 475 | 10.2 | -2.1 | 3.12 | A | A | A |
| 496.09 | 69.45 | 2.3 | 14.18 | A | A | A |
| 513 | 42 | 5.8 | 8.50 | A | A | A |
| 460 | 15 | -5.2 | 3.97 | A | A | A |
| 504.86 | 9.51 | 4.1 | 2.95 | A | A | A |
| 603.4 | 55.6 | 24.4 | 9.49 | N | A | N |
| 450 | 20 | -7.2 | 4.99 | A | A | A |
| 548 | 30.3 | 13.0 | 5.98 | A | A | A |
| 421 | 47 | -13.2 | 11.39 | A | A | A |
| 526 | 34 | 8.4 | 6.85 | A | A | A |
| 486.2 | 22.3 | 0.2 | 5.12 | A | A | A |
| 428.27 | 13.53 | -11.7 | 3.89 | A | N | W |
| 518 | 25.2 | 6.8 | 5.37 | A | A | A |
| 565 | 10 | 16.5 | 2.88 | A | N | W |

Labcode

113

114

116

119

121

122

124

125

127

128

131

132

133

134

139

140

145

147

150

151

152

154

155

156

158

159

161

163

165

166

168

170

172

175

178

179

180

181

183

184

185

187

189

191

192

193

194

197

198

200

202

203

204

205

206

207

208

211

213

214

| Reported value [Bq/kg ] 470 | Reported uncertainty (k = 1) [Bq/kg ] 25 | Relative bias [%] -3.1 | P-Test [%] 5.78 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 447.13067 | 18.96132 | -7.8 | 4.81 | A | A | A |
| 470.9 | 18.6 | -2.9 | 4.55 | A | A | A |
| 480.3 | 24.8 | -1.0 | 5.64 | A | A | A |
| 605 | 60 | 24.7 | 10.17 | N | A | N |
| 445.2 | 40.1 | -8.2 | 9.29 | A | A | A |
| 523 | 21.75 | 7.8 | 4.74 | A | A | A |
| 525 | 20 | 8.2 | 4.43 | A | A | A |
| 442 | 21 | -8.9 | 5.26 | A | A | A |
| 479 | 24 | -1.2 | 5.50 | A | A | A |
| 465 | 32 | -4.1 | 7.25 | A | A | A |
| 481 | 43 | -0.8 | 9.22 | A | A | A |
| 455 | 20.2 | -6.2 | 4.99 | A | A | A |
| 473.1 | 11.82 | -2.5 | 3.37 | A | A | A |
| 437.3 | 38.684 | -9.8 | 9.13 | A | A | A |
| 441.67 | 25.531328 | -8.9 | 6.21 | A | A | A |
| 449.48 | 16.71 | -7.3 | 4.35 | A | A | A |
| 478 | 29 | -1.4 | 6.48 | A | A | A |
| 691.74 | 51 | 42.6 | 7.71 | N | N | N |
| 521.99 | 26.62 | 7.6 | 5.58 | A | A | A |
| 489.91 | 21.52 | 1.0 | 4.94 | A | A | A |
| 521 | 24 | 7.4 | 5.13 | A | A | A |
| 511 | 27 | 5.4 | 5.75 | A | A | A |
| 552 | 10 | 13.8 | 2.90 | A | N | W |
| 613.4 | 18 | 26.5 | 3.71 | N | N | N |
| 493 | 59 | 1.6 | 12.18 | A | A | A |
| 490 | 44 | 1.0 | 9.26 | A | A | A |
| 461 | 23.1 | -5.0 | 5.50 | A | A | A |
| 482 | 20 | -0.6 | 4.73 | A | A | A |
| 398.4 | 12.9 | -17.9 | 3.95 | A | N | W |
| 473 | 18 | -2.5 | 4.43 | A | A | A |
| 583 | 26 | 20.2 | 5.00 | N | N | N |
| 524 | 26.118031 | 8.0 | 5.48 | A | A | A |
| 481.43 | 33.6 | -0.7 | 7.34 | A | A | A |
| 504 | 15 | 3.9 | 3.74 | A | A | A |
| 527.5 | 18.2 | 8.8 | 4.13 | A | A | A |
| 540 | 24 | 11.3 | 4.99 | A | A | A |
| 1071.6 | 105.24 | 121.0 | 10.08 | N | N | N |
| 518 | 11 | 6.8 | 3.11 | A | A | A |
| 369 | 9.2 | -23.9 | 3.37 | N | N | N |
| 535 | 18 | 10.3 | 4.06 | A | A | A |
| 520.7 | 40.7 | 7.4 | 8.14 | A | A | A |
| 476.5 | 12 | -1.8 | 3.39 | A | A | A |
| 473 | 48 | -2.5 | 10.40 | A | A | A |
| 475 | 20 | -2.1 | 4.78 | A | A | A |
| 502 | 65.26 | 3.5 | 13.20 | A | A | A |
| 520 | 11 | 7.2 | 3.10 | A | A | A |
| 444.7 | 43.5 | -8.3 | 10.04 | A | A | A |
| 449 | 42 | -7.4 | 9.63 | A | A | A |
| 506.72 | 24.58 | 4.5 | 5.35 | A | A | A |
| 533 | 35 | 9.9 | 6.95 | A | A | A |
| 530 | 40 | 9.3 | 7.88 | A | A | A |
| 444.7 | 5.33 | -8.3 | 2.57 | A | N | W |
| 527 | 42 | 8.7 | 8.29 | A | A | A |
| 350 | 60 | -27.8 | 17.29 | N | A | N |
| 507 | 18 | 4.5 | 4.21 | A | A | A |
| 471 | 47 | -2.9 | 10.23 | A | A | A |
| 492 | 27 | 1.4 | 5.94 | A | A | A |
| 527 | 30 | 8.7 | 6.13 | A | A | A |
| 497.1 | 14.3 | 2.5 | 3.66 | A | A | A |

Labcode

215

218

219

220

223

224

226

227

230

232

233

235

238

242

247

249

252

253

254

256

257

258

260

261

262

265

267

268

270

271

272

273

274

276

277

278

279

280

281

283

284

286

287

288

290

291

292

293

294

295

297

299

300

302

303

304

306

307

310

311

| Reported value [Bq/kg ] 480.8 | Reported uncertainty (k = 1) [Bq/kg ] 20.5 | Relative bias [%] -0.9 | P-Test [%] 4.83 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 460 | 30.1 | -5.2 | 6.93 | A | A | A |
| 493 | 8 | 1.6 | 2.79 | A | A | A |
| 537 | 26 | 10.7 | 5.35 | A | A | A |
| 440.6 | 38.2 | -9.2 | 8.96 | A | A | A |
| 625 | 35 | 28.9 | 6.04 | N | N | N |
| 537 | 12.06 | 10.7 | 3.19 | A | N | W |
| 486 | 20 | 0.2 | 4.70 | A | A | A |
| 455 | 27 | -6.2 | 6.35 | A | A | A |
| 897.4 | 40.3 | 85.0 | 5.03 | N | N | N |
| 478 | 34.3 | -1.4 | 7.53 | A | A | A |
| 691.42 | 40.8 | 42.6 | 6.32 | N | N | N |
| 515 | 41 | 6.2 | 8.28 | A | A | A |
| 505.5 | 45 | 4.2 | 9.19 | A | A | A |
| 500 | 41 | 3.1 | 8.51 | A | A | A |
| 1758 | 37.24 | 262.5 | 3.10 | N | N | N |
| 519 | 31.5 | 7.0 | 6.48 | A | A | A |
| 536.2 | 18.6 | 10.6 | 4.14 | A | A | A |
| 488 | 14 | 0.6 | 3.66 | A | A | A |
| 501.3 | 45.3 | 3.4 | 9.32 | A | A | A |
| 424 | 20 | -12.6 | 5.23 | A | A | A |
| 481.2 | 24.6 | -0.8 | 5.59 | A | A | A |
| 430 | 30 | -11.3 | 7.34 | A | A | A |
| 431 | 16 | -11.1 | 4.35 | A | A | A |
| 500 | 20 | 3.1 | 4.60 | A | A | A |
| 419.28 | 4.73 | -13.6 | 2.53 | A | N | W |
| 508 | 20 | 4.7 | 4.54 | A | A | A |
| 464 | 23 | -4.3 | 5.45 | A | A | A |
| 537 | 53 | 10.7 | 10.13 | A | A | A |
| 467 | 21 | -3.7 | 5.04 | A | A | A |
| 517 | 28 | 6.6 | 5.87 | A | A | A |
| 570 | 70 | 17.5 | 12.49 | A | A | A |
| 468 | 17.4 | -3.5 | 4.36 | A | A | A |
| 508 | 22 | 4.7 | 4.89 | A | A | A |
| 541 | 54 | 11.6 | 10.24 | A | A | A |
| 532 | 15 | 9.7 | 3.62 | A | N | W |
| 530.8 | 21.2 | 9.4 | 4.59 | A | A | A |
| 601.1694 | 26.60554 | 24.0 | 4.97 | N | N | N |
| 480 | 50 | -1.0 | 10.66 | A | A | A |
| 491.95 | 32.08 | 1.4 | 6.90 | A | A | A |
| 516 | 23 | 6.4 | 5.00 | A | A | A |
| 496 | 50 | 2.3 | 10.33 | A | A | A |
| 444.99 | 54.9 | -8.2 | 12.54 | A | A | A |
| 508 | 20 | 4.7 | 4.54 | A | A | A |
| 520 | 20 | 7.2 | 4.47 | A | A | A |
| 448.34 | 14.34 | -7.6 | 3.92 | A | A | A |
| 478.6 | 25.3 | -1.3 | 5.75 | A | A | A |
| 440 | 44 | -9.3 | 10.25 | A | A | A |
| 298.87 | 12.09 | -38.4 | 4.64 | N | N | N |
| 490 | 45 | 1.0 | 9.46 | A | A | A |
| 259 | 2 | -46.6 | 2.40 | N | N | N |
| 841 | 43 | 73.4 | 5.59 | N | N | N |
| 506 | 30 | 4.3 | 6.35 | A | A | A |
| 478.5 | 19.55 | -1.3 | 4.67 | A | A | A |
| 487 | 21 | 0.4 | 4.87 | A | A | A |
| 510 | 25 | 5.2 | 5.40 | A | A | A |
| 463 | 15 | -4.5 | 3.95 | A | A | A |
| 521.3 | 35.5 | 7.5 | 7.18 | A | A | A |
| 489.6 | 16.6 | 0.9 | 4.08 | A | A | A |
| 541.8 | 50.37 | 11.7 | 9.57 | A | A | A |

Labcode

312

313

314

315

316

317

318

319

322

323

325

326

328

329

332

333

338

339

341

344

346

347

348

350

354

355

356

357

358

363

364

367

369

370

371

373

376

377

379

381

382

384

386

387

388

391

393

395

397

398

400

401

402

403

405

407

408

410

411

412

| Reported value [Bq/kg ] 533 | Reported uncertainty (k = 1) [Bq/kg ] 22 | Relative bias [%] 9.9 | P-Test [%] 4.71 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 486 | 30 | 0.2 | 6.58 | A | A | A |
| 493 | 33 | 1.6 | 7.07 | A | A | A |
| 483 | 35 | -0.4 | 7.59 | A | A | A |
| 500 | 34 | 3.1 | 7.17 | A | A | A |
| 492.3 | 41.39 | 1.5 | 8.71 | A | A | A |
| 494.9 | 35 | 2.0 | 7.43 | A | A | A |
| 551 | 6.3 | 13.6 | 2.54 | A | N | W |
| 540 | 31 | 11.3 | 6.17 | A | A | A |
| 583.59 | 17.22 | 20.3 | 3.72 | N | N | N |
| 410.8 | 26.6 | -15.3 | 6.86 | A | A | A |
| 475 | 20 | -2.1 | 4.78 | A | A | A |
| 442 | 29 | -8.9 | 6.94 | A | A | A |
| 1258.3 | 16.2 | 159.4 | 2.61 | N | N | N |
| 507.6 | 27.1 | 4.7 | 5.80 | A | A | A |
| 505.76495 | 34.604526 | 4.3 | 7.21 | A | A | A |
| 500 | 40 | 3.1 | 8.32 | A | A | A |
| 482 | 30.7 | -0.6 | 6.76 | A | A | A |
| 496 | 26 | 2.3 | 5.71 | A | A | A |
| 493.7 | 28 | 1.8 | 6.11 | A | A | A |
| 443.27 | 8.49 | -8.6 | 2.97 | A | N | W |
| 468 | 14 | -3.5 | 3.75 | A | A | A |
| 460 | 29.4 | -5.2 | 6.78 | A | A | A |
| 451 | 16.8 | -7.0 | 4.36 | A | A | A |
| 820 | 41 | 69.1 | 5.49 | N | N | N |
| 480.3 | 37.6 | -1.0 | 8.15 | A | A | A |
| 549 | 34 | 13.2 | 6.60 | A | A | A |
| 702 | 35 | 44.7 | 5.48 | N | N | N |
| 520 | 26 | 7.2 | 5.49 | A | A | A |
| 511.93 | 48.16 | 5.6 | 9.68 | A | A | A |
| 476 | 45 | -1.9 | 9.72 | A | A | A |
| 545.49 | 109.3 | 12.5 | 20.16 | A | N | W |
| 467 | 23 | -3.7 | 5.42 | A | A | A |
| 490 | 10 | 1.0 | 3.05 | A | A | A |
| 490 | 26 | 1.0 | 5.77 | A | A | A |
| 483.6 | 29.9 | -0.3 | 6.59 | A | A | A |
| 533 | 35 | 9.9 | 6.95 | A | A | A |
| 510 | 25 | 5.2 | 5.40 | A | A | A |
| 507.6 | 27.1 | 4.7 | 5.80 | A | A | A |
| 507 | 24 | 4.5 | 5.25 | A | A | A |
| 465 | 20 | -4.1 | 4.86 | A | A | A |
| 1088.221 | 26.785 | 124.4 | 3.35 | N | N | N |
| 485.9 | 20 | 0.2 | 4.70 | A | A | A |
| 831 | 20 | 71.3 | 3.31 | N | N | N |
| 541.7 | 15.2 | 11.7 | 3.61 | A | N | W |
| 471.52 | 31.37 | -2.8 | 7.03 | A | A | A |
| 399 | 20 | -17.7 | 5.50 | A | N | W |
| 424.5 | 32.61 | -12.5 | 8.01 | A | A | A |
| 575 | 26 | 18.6 | 5.06 | A | N | W |
| 486 | 41 | 0.2 | 8.74 | A | A | A |
| 642.33 | 6.741 | 32.4 | 2.50 | N | N | N |
| 489 | 12 | 0.8 | 3.34 | A | A | A |
| 4723 | 8 | 873.8 | 2.27 | N | N | N |
| 464 | 40 | -4.3 | 8.91 | A | A | A |
| 496 | 29 | 2.3 | 6.27 | A | A | A |
| 753.83 | 31.77 | 55.4 | 4.79 | N | N | N |
| 470 | 16 | -3.1 | 4.09 | A | A | A |
| 467 | 18 | -3.7 | 4.47 | A | A | A |
| 496 | 26 | 2.3 | 5.71 | A | A | A |
| 516 | 30 | 6.4 | 6.24 | A | A | A |

Labcode

414

417

420

421

425

426

428

429

430

433

434

435

437

440

441

442

443

444

445

446

448

450

452

453

455

459

460

462

463

464

465

467

475

477

478

479

480

482

483

484

487

491

492

495

496

498

499

500

501

506

507

508

509

514

516

520

522

523

525

527

| Reported value [Bq/kg ] 441 | Reported uncertainty (k = 1) [Bq/kg ] 22 | Relative bias [%] -9.1 | P-Test [%] 5.48 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 722 | 10 | 48.9 | 2.66 | N | N | N |
| 550.4 | 29.25 | 13.5 | 5.78 | A | A | A |
| 529.3 | 34.31 | 9.1 | 6.87 | A | A | A |
| 610 | 32 | 25.8 | 5.72 | N | N | N |
| 418 | 26 | -13.8 | 6.62 | A | A | A |
| 520 | 39 | 7.2 | 7.84 | A | A | A |
| 508 | 24 | 4.7 | 5.24 | A | A | A |
| 455.51 | 24.09 | -6.1 | 5.75 | A | A | A |
| 494.083 | 40.34 | 1.9 | 8.47 | A | A | A |
| 487 | 15 | 0.4 | 3.83 | A | A | A |
| 460 | 22 | -5.2 | 5.29 | A | A | A |
| 460.3 | 41.4 | -5.1 | 9.28 | A | A | A |

Labcode

528

529

530

532

533

535

536

537

539

544

549

553

555

Cs-137 determination of activity concentration in Sample 3

Table 36. Target values

| Parameter | Value |
| --- | --- |
| Activity concentration [ Bq/kg ] | 10.3 |
| Uncertainty [ Bq/kg ] | 0.3 |
| Maximum Acceptable Relative Bias [%] | 20 |

Accepted 66%	Not Accepted 8%	Warning 5%	Not Reported 21%

14

Figure 36. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

6

8

37

300

400

83

268

145

185

491

34

283

81

95

66

294

410

369

401

535

528

395

455

38

437

35

213

79

487

391

252

211

297

453

238

24

67

348

496

155

21

168

184

290

514

40

314

119

235

270

277

319

445

452

166

363

100

479

2

102

204

274

347

357

465

342

8

249

70

152

313

420

403

110

111

208

398

523

97

495

53

32

42

57

302

408

484

509

256

216

233

328

430

478

71

295

242

150

306

356

379

555

267

498

107

139

370

539

48

124

303

425

522

11

260

198

15

315

376

428

475

13

279

28

397

284

114

506

226

338

109

482

20

532

292

411

133

354

178

172

214

273

462

501

158

288

388

387

281

463

311

113

203

440

165

323

333

181

326

442

Laboratory code

Figure 37. Reported results and evaluation parameters for Cs-137. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 37. Evaluation results for Cs-137 in Sample 3, Rep. Value and Rep. Unc. in Bq/kg

Labcode

Reported value

[Bq/kg ]

Reported uncertainty

(k = 1) [Bq/kg ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

Activity concentration  [ Bq/kg ]

10

12

2	10.3	1	0.0	10.14	A	A	A

8	10.36	0.72	0.6	7.54	A	A	A

9	8.7	1	-15.5	11.86	A	A	A

| Reported value [Bq/kg ] 11.1 | Reported uncertainty (k = 1) [Bq/kg ] 0.773 | Relative bias [%] 7.8 | P-Test [%] 7.55 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 15 | 2.015 | 45.6 | 13.75 | N | N | N |
| 11.3 | 0.5 | 9.7 | 5.30 | A | A | A |
| 11.2 | 0.6 | 8.7 | 6.10 | A | A | A |
| 10.67 | 0.85 | 3.6 | 8.48 | A | A | A |
| 11.54 | 0.45 | 12.0 | 4.87 | A | A | A |
| 10 | 0.9 | -2.9 | 9.46 | A | A | A |
| 9.88 | 0.65 | -4.1 | 7.19 | A | A | A |
| 10 | 0.6 | -2.9 | 6.67 | A | A | A |
| 11.31 | 1.002 | 9.8 | 9.33 | A | A | A |
| 10.8 | 0.48 | 4.8 | 5.31 | A | A | A |
| 10.59 | 1.24 | 2.8 | 12.07 | A | A | A |
| 9 | 0.7 | -12.6 | 8.31 | A | A | A |
| 9.64 | 1.4 | -6.4 | 14.81 | A | A | A |
| 0.065 | 0.013 | -99.4 | 20.21 | N | N | N |
| 9.58 | 1.91 | -7.0 | 20.15 | A | N | W |
| 10.06 | 0.7 | -2.3 | 7.54 | A | A | A |
| 10.06 | 0.778 | -2.3 | 8.26 | A | A | A |
| 10.6 | 0.4 | 2.9 | 4.77 | A | A | A |
| 9.62 | 0.49 | -6.6 | 5.87 | A | A | A |
| 11 | 0.5 | 6.8 | 5.40 | A | A | A |
| 11.1 | 1.3 | 7.8 | 12.07 | A | A | A |
| 10.4 | 0.7 | 1.0 | 7.33 | A | A | A |
| 10.6 | 0.66 | 2.9 | 6.87 | A | A | A |
| 10.58 | 0.51 | 2.7 | 5.63 | A | A | A |
| 11.3 | 0.6 | 9.7 | 6.06 | A | A | A |
| 11.63 | 0.74 | 12.9 | 7.00 | A | A | A |
| 10.6 | 1 | 2.9 | 9.87 | A | A | A |
| 9.6 | 0.4 | -6.8 | 5.08 | A | A | A |
| 10.9 | 1.6 | 5.8 | 14.97 | A | A | A |
| 9.98 | 1.2 | -3.1 | 12.37 | A | A | A |
| 10.55 | 0.27 | 2.4 | 3.88 | A | A | A |
| 9.8 | 0.7 | -4.8 | 7.71 | A | A | A |
| 11.4 | 0.9 | 10.7 | 8.41 | A | A | A |
| 9.257 | 0.5085 | -10.1 | 6.22 | A | A | A |
| 9.9 | 1.2 | -3.9 | 12.47 | A | A | A |
| 10.3 | 0.54 | 0.0 | 6.00 | A | A | A |
| 10.1 | 0.6 | -1.9 | 6.62 | A | A | A |
| 10.4 | 0.3 | 1.0 | 4.10 | A | A | A |
| 10.71 | 0.54 | 4.0 | 5.82 | A | A | A |
| 10.2 | 0.43 | -1.0 | 5.12 | A | A | A |
| 10.5 | 0.4 | 1.9 | 4.80 | A | A | A |
| 9.7 | 0.3 | -5.8 | 4.25 | A | A | A |
| 9.2 | 1.223 | -10.7 | 13.61 | A | A | A |
| 19.73 | 0.59 | 91.6 | 4.17 | N | N | N |
| 8.3 | 1.2 | -19.4 | 14.75 | A | A | A |
| 9.41 | 0.69 | -8.6 | 7.89 | A | A | A |
| 13.522 | 1.593 | 31.3 | 12.14 | N | A | N |
| 11 | 0.5 | 6.8 | 5.40 | A | A | A |
| 12.13 | 1.12 | 17.8 | 9.68 | A | A | A |
| 9.23 | 0.32 | -10.4 | 4.53 | A | A | A |
| 10.52 | 1.2 | 2.1 | 11.77 | A | A | A |
| 10.21 | 0.75 | -0.9 | 7.90 | A | A | A |
| 10.3 | 0.9 | 0.0 | 9.21 | A | A | A |
| 10.4 | 0.6 | 1.0 | 6.46 | A | A | A |
| 10.6 | 0.7 | 2.9 | 7.22 | A | A | A |
| 10.9 | 0.9 | 5.8 | 8.76 | A | A | A |
| 11.5 | 0.2 | 11.7 | 3.39 | A | N | W |
| 10.46 | 0.62 | 1.6 | 6.60 | A | A | A |
| 10.5 | 0.7 | 1.9 | 7.28 | A | A | A |

Labcode

11

12

13

15

19

20

21

24

25

28

31

32

34

35

37

38

40

41

42

45

48

49

50

52

53

54

55

57

58

60

62

63

64

65

66

67

68

69

70

71

72

78

79

81

82

83

84

85

90

94

95

97

100

102

104

105

107

109

110

111

| Reported value [Bq/kg ] 12.85 | Reported uncertainty (k = 1) [Bq/kg ] 1.29 | Relative bias [%] 24.8 | P-Test [%] 10.45 | Trueness evaluation N | Precision evaluation A | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 11.4 | 0.9 | 10.7 | 8.41 | A | A | A |
| 10.27 | 0.72 | -0.3 | 7.59 | A | A | A |
| 10.1 | 0.8 | -1.9 | 8.44 | A | A | A |
| 10.35 | 0.65 | 0.5 | 6.92 | A | A | A |
| 11 | 1 | 6.8 | 9.55 | A | A | A |
| 10.9 | 0.4 | 5.8 | 4.69 | A | A | A |
| 10.65 | 0.34 | 3.4 | 4.32 | A | A | A |
| 10.77 | 0.67 | 4.6 | 6.87 | A | A | A |
| 9.255779 | 1.086739 | -10.1 | 12.10 | A | A | A |
| 10.3 | 0.5 | 0.0 | 5.66 | A | A | A |
| 11.7 | 1 | 13.6 | 9.03 | A | A | A |
| 10.75 | 0.2 | 4.4 | 3.46 | A | A | A |
| 10.9 | 0.69 | 5.8 | 6.97 | A | A | A |
| 9.75 | 0.54 | -5.3 | 6.26 | A | A | A |
| 8.53 | 1.36 | -17.2 | 16.21 | A | A | A |
| 13.364 | 0.8289 | 29.8 | 6.85 | N | N | N |
| 10.8 | 0.8 | 4.8 | 7.96 | A | A | A |
| 9.31 | 0.93 | -9.6 | 10.41 | A | A | A |
| 10.4 | 0.7 | 1.0 | 7.33 | A | A | A |
| 10.1 | 0.5 | -1.9 | 5.74 | A | A | A |
| 9.98 | 0.72 | -3.1 | 7.78 | A | A | A |
| 10.38 | 0.28 | 0.8 | 3.97 | A | A | A |
| 12.1 | 1.6 | 17.5 | 13.54 | A | A | A |
| 10.5 | 0.549 | 1.9 | 5.99 | A | A | A |
| 11.5 | 0.7 | 11.7 | 6.75 | A | A | A |
| 13.72 | 0.8 | 33.2 | 6.52 | N | N | N |
| 10.2 | 1.94 | -1.0 | 19.24 | A | A | A |
| 10 | 0.8 | -2.9 | 8.51 | A | A | A |
| 10.7 | 0.9 | 3.9 | 8.90 | A | A | A |
| 12 | 1 | 16.5 | 8.83 | A | A | A |
| 11.07 | 0.48 | 7.5 | 5.22 | A | A | A |
| 11.9 | 1.2 | 15.5 | 10.50 | A | A | A |
| 10 | 0.65 | -2.9 | 7.12 | A | A | A |
| 15.11 | 1.22 | 46.7 | 8.58 | N | N | N |
| 11 | 0.5 | 6.8 | 5.40 | A | A | A |
| 10 | 1.5 | -2.9 | 15.28 | A | A | A |
| 8.7 | 0.24 | -15.5 | 4.01 | A | N | W |
| 11.4 | 0.9 | 10.7 | 8.41 | A | A | A |
| 9.22 | 0.89 | -10.5 | 10.08 | A | A | A |
| 11.41 | 0.72 | 10.8 | 6.95 | A | A | A |
| 10.4 | 1.1 | 1.0 | 10.97 | A | A | A |
| 9.89 | 0.54 | -4.0 | 6.19 | A | A | A |
| 11.52 | 1.61 | 11.8 | 14.28 | A | A | A |
| 11.1 | 1 | 7.8 | 9.47 | A | A | A |
| 10 | 0.4 | -2.9 | 4.95 | A | A | A |
| 10.83 | 0.2 | 5.2 | 3.45 | A | A | A |
| 13.1 | 1.2 | 27.2 | 9.61 | N | N | N |
| 10.3 | 0.5 | 0.0 | 5.66 | A | A | A |
| 12 | 0.711 | 16.5 | 6.60 | A | A | A |
| 9 | 1.2 | -12.6 | 13.65 | A | A | A |
| 11.31 | 0.76 | 9.8 | 7.32 | A | A | A |
| 10.5 | 0.68 | 1.9 | 7.10 | A | A | A |
| 9.77 | 0.31 | -5.2 | 4.31 | A | A | A |
| 9.68 | 0.86 | -6.0 | 9.35 | A | A | A |
| 12 | 1 | 16.5 | 8.83 | A | A | A |
| 10.3 | 0.5 | 0.0 | 5.66 | A | A | A |
| 10.69 | 0.4 | 3.8 | 4.74 | A | A | A |
| 9.3 | 2 | -9.7 | 21.70 | A | N | W |
| 9.90008 | 0.6212 | -3.9 | 6.92 | A | A | A |

Labcode

113

114

116

119

122

124

127

128

130

131

132

133

134

139

140

145

147

150

151

152

154

155

156

158

159

161

165

166

168

170

172

175

178

180

181

183

184

185

187

190

191

192

194

197

198

200

202

203

204

205

206

207

208

211

213

214

215

216

217

218

| Reported value [Bq/kg ] 10.9 | Reported uncertainty (k = 1) [Bq/kg ] 0.8 | Relative bias [%] 5.8 | P-Test [%] 7.90 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 12.5 | 1.2 | 21.4 | 10.03 | N | A | N |
| 11.45 | 0.69 | 11.2 | 6.69 | A | A | A |
| 11.6 | 0.4 | 12.6 | 4.51 | A | N | W |
| 9.4584 | 0.62501 | -8.2 | 7.22 | A | A | A |
| 9.35 | 0.38 | -9.2 | 5.00 | A | A | A |
| 9.78 | 1.01 | -5.0 | 10.73 | A | A | A |
| 10.7 | 0.8 | 3.9 | 8.02 | A | A | A |
| 10.1 | 1 | -1.9 | 10.32 | A | A | A |
| 9.82 | 0.674 | -4.7 | 7.46 | A | A | A |
| 10.79 | 0.76 | 4.8 | 7.62 | A | A | A |
| 10.26 | 0.695 | -0.4 | 7.37 | A | A | A |
| 11.132 | 0.95 | 8.1 | 9.02 | A | A | A |
| 10.38 | 0.176536 | 0.8 | 3.37 | A | A | A |
| 9.76 | 0.25 | -5.2 | 3.88 | A | A | A |
| 12 | 0.9 | 16.5 | 8.05 | A | A | A |
| 11.71 | 3.14 | 13.7 | 26.97 | A | N | W |
| 10.65 | 0.55 | 3.4 | 5.93 | A | A | A |
| 9.76 | 0.86 | -5.2 | 9.28 | A | A | A |
| 10.2 | 0.5 | -1.0 | 5.70 | A | A | A |
| 11.1 | 0.6 | 7.8 | 6.14 | A | A | A |
| 12.1 | 0.3 | 17.5 | 3.82 | A | N | W |
| 12.7 | 0.4 | 23.3 | 4.29 | N | N | N |
| 10.1 | 1.2 | -1.9 | 12.23 | A | A | A |
| 10.85 | 1.06 | 5.3 | 10.19 | A | A | A |
| 8.38 | 0.42 | -18.6 | 5.80 | A | N | W |
| 10.1 | 0.5 | -1.9 | 5.74 | A | A | A |
| 9.1 | 0.4 | -11.7 | 5.27 | A | A | A |
| 10.1 | 0.6 | -1.9 | 6.62 | A | A | A |
| 12 | 1 | 16.5 | 8.83 | A | A | A |
| 10.3 | 0.58035 | 0.0 | 6.34 | A | A | A |
| 10.3 | 0.68 | 0.0 | 7.22 | A | A | A |
| 10.1 | 0.3 | -1.9 | 4.16 | A | A | A |
| 9.77 | 0.6 | -5.2 | 6.80 | A | A | A |
| 11.3 | 0.6 | 9.7 | 6.06 | A | A | A |
| 15.2 | 0.73 | 47.6 | 5.62 | N | N | N |
| 12.37 | 0.72 | 20.1 | 6.51 | N | N | N |
| 9.05 | 0.27 | -12.1 | 4.17 | A | N | W |
| 11.38 | 0.67 | 10.5 | 6.57 | A | A | A |
| 10.97 | 0.77 | 6.5 | 7.60 | A | A | A |
| 9.7 | 1.5 | -5.8 | 15.74 | A | A | A |
| 12 | 0.5 | 16.5 | 5.08 | A | N | W |
| 12.1 | 0.61 | 17.5 | 5.82 | A | N | W |
| 10 | 0.5 | -2.9 | 5.79 | A | A | A |
| 10.15 | 1.218 | -1.5 | 12.35 | A | A | A |
| 11.6 | 0.4 | 12.6 | 4.51 | A | N | W |
| 10.8 | 1.2 | 4.8 | 11.49 | A | A | A |
| 9.3 | 0.8 | -9.7 | 9.08 | A | A | A |
| 10.75 | 0.65 | 4.4 | 6.71 | A | A | A |
| 9.77 | 0.61 | -5.2 | 6.89 | A | A | A |
| 10.7 | 0.7 | 3.9 | 7.16 | A | A | A |
| 0.56 | 0.25 | -94.6 | 44.74 | N | N | N |
| 10.6 | 0.8 | 2.9 | 8.09 | A | A | A |
| 11 | 1.5 | 6.8 | 13.94 | A | A | A |
| 10.5 | 0.4 | 1.9 | 4.80 | A | A | A |
| 10.8 | 1.6 | 4.8 | 15.10 | A | A | A |
| 11.3 | 0.8 | 9.7 | 7.66 | A | A | A |
| 11.2 | 0.8 | 8.7 | 7.71 | A | A | A |
| 12.8 | 0.436 | 24.3 | 4.48 | N | N | N |
| 10.1 | 0.5 | -1.9 | 5.74 | A | A | A |

Labcode

220

223

226

227

229

230

232

233

235

238

242

243

247

249

252

253

254

256

257

258

260

261

262

265

267

268

270

271

272

273

274

276

277

278

279

280

281

283

284

285

286

287

288

290

291

292

293

294

295

297

299

300

302

303

304

306

307

310

311

312

| Reported value [Bq/kg ] 10.4 | Reported uncertainty (k = 1) [Bq/kg ] 0.827 | Relative bias [%] 1.0 | P-Test [%] 8.47 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 10.07 | 0.35 | -2.2 | 4.53 | A | A | A |
| 11.2 | 0.57 | 8.7 | 5.86 | A | A | A |
| 9.9 | 1.3 | -3.9 | 13.45 | A | A | A |
| 13.8 | 1.1 | 34.0 | 8.49 | N | N | N |
| 11.1 | 0.25 | 7.8 | 3.68 | A | A | A |
| 10.1 | 0.4 | -1.9 | 4.92 | A | A | A |
| 11.6 | 0.6 | 12.6 | 5.94 | A | A | A |
| 13.9 | 0.6 | 35.0 | 5.21 | N | N | N |
| 10.6 | 0.78 | 2.9 | 7.91 | A | A | A |
| 16.2 | 1.1 | 57.3 | 7.39 | N | N | N |
| 10.7 | 0.8 | 3.9 | 8.02 | A | A | A |
| 9.5 | 0.5 | -7.8 | 6.02 | A | A | A |
| 11 | 1 | 6.8 | 9.55 | A | A | A |
| 14.37 | 1.22 | 39.5 | 8.98 | N | N | N |
| 11.48 | 0.7 | 11.5 | 6.76 | A | A | A |
| 11.2 | 0.8 | 8.7 | 7.71 | A | A | A |
| 10.57 | 0.31 | 2.6 | 4.13 | A | A | A |
| 10.32 | 0.48 | 0.2 | 5.49 | A | A | A |
| 10.8 | 0.8 | 4.8 | 7.96 | A | A | A |
| 10.3 | 0.51 | 0.0 | 5.74 | A | A | A |
| 9.9 | 0.8 | -3.9 | 8.59 | A | A | A |
| 10.3 | 1 | 0.0 | 10.14 | A | A | A |
| 11.8 | 0.6 | 14.6 | 5.86 | A | A | A |
| 8.35 | 0.17 | -18.9 | 3.55 | A | N | W |
| 10.8 | 0.4 | 4.8 | 4.71 | A | A | A |
| 10.3 | 0.6 | 0.0 | 6.51 | A | A | A |
| 12.1 | 1.2 | 17.5 | 10.34 | A | A | A |
| 10.2 | 0.4 | -1.0 | 4.88 | A | A | A |
| 10 | 0.6 | -2.9 | 6.67 | A | A | A |
| 10.8 | 1.9 | 4.8 | 17.83 | A | A | A |
| 9.31 | 0.43 | -9.6 | 5.46 | A | A | A |
| 10.9 | 0.6 | 5.8 | 6.23 | A | A | A |
| 11.21 | 0.67 | 8.8 | 6.65 | A | A | A |
| 12.8 | 0.5 | 24.3 | 4.87 | N | N | N |
| 11.2 | 0.45 | 8.7 | 4.96 | A | A | A |
| 12.30795 | 0.801424 | 19.5 | 7.13 | A | N | W |
| 10.8 | 1.1 | 4.8 | 10.59 | A | A | A |
| 10.59 | 0.67 | 2.8 | 6.96 | A | A | A |
| 10.8 | 0.7 | 4.8 | 7.11 | A | A | A |
| 10.7 | 1.3 | 3.9 | 12.49 | A | A | A |
| 4.3 | 0.46 | -58.2 | 11.09 | N | N | N |
| 9.3 | 1.21 | -9.7 | 13.33 | A | A | A |
| 12.2 | 0.3 | 18.4 | 3.81 | A | N | W |
| 12.1 | 0.3 | 17.5 | 3.82 | A | N | W |
| 9.72 | 0.31 | -5.6 | 4.32 | A | A | A |
| 10.51 | 0.55 | 2.0 | 5.99 | A | A | A |
| 9.47 | 0.95 | -8.1 | 10.45 | A | A | A |
| 11.36 | 0.54 | 10.3 | 5.57 | A | A | A |
| 10.5 | 0.9 | 1.9 | 9.05 | A | A | A |
| 6 | 2 | -41.8 | 33.46 | N | N | N |
| 9.36 | 1.61 | -9.1 | 17.45 | A | A | A |
| 11.2 | 0.7 | 8.7 | 6.90 | A | A | A |
| 10.42 | 0.65 | 1.2 | 6.88 | A | A | A |
| 10.1 | 0.5 | -1.9 | 5.74 | A | A | A |
| 10.4 | 0.5 | 1.0 | 5.62 | A | A | A |
| 10.6 | 0.3 | 2.9 | 4.06 | A | A | A |
| 9.3 | 0.7 | -9.7 | 8.07 | A | A | A |
| 11.6 | 0.3 | 12.6 | 3.90 | A | N | W |
| 8.84 | 1.46 | -14.2 | 16.77 | A | A | A |

Labcode

313

314

315

316

317

318

319

322

323

325

326

328

329

332

333

338

339

341

342

344

347

348

350

354

355

356

357

358

363

364

367

369

370

371

373

376

377

379

381

382

384

385

386

387

388

391

393

395

397

398

400

401

402

403

405

407

408

410

411

412

| Reported value [Bq/kg ] 11.49 | Reported uncertainty (k = 1) [Bq/kg ] 0.7 | Relative bias [%] 11.6 | P-Test [%] 6.75 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 9.71 | 0.39 | -5.7 | 4.96 | A | A | A |
| 0.16 | 0.02 | -98.4 | 12.83 | N | N | N |
| 10.4 | 0.75 | 1.0 | 7.78 | A | A | A |
| 10.6 | 0.8 | 2.9 | 8.09 | A | A | A |
| 11 | 0.8 | 6.8 | 7.83 | A | A | A |
| 11.45 | 1.45 | 11.2 | 12.99 | A | A | A |
| 11.2 | 1.1 | 8.7 | 10.24 | A | A | A |
| 9.99 | 0.381 | -3.0 | 4.80 | A | A | A |
| 10.7 | 0.88 | 3.9 | 8.72 | A | A | A |
| 12.09 | 0.72 | 17.4 | 6.63 | A | N | W |
| 8.48 | 0.9 | -17.7 | 11.01 | A | A | A |
| 10.5 | 0.8 | 1.9 | 8.16 | A | A | A |
| 9.6 | 0.8 | -6.8 | 8.83 | A | A | A |
| 13.405 | 0.699 | 30.1 | 5.97 | N | N | N |
| 10.7 | 0.6 | 3.9 | 6.32 | A | A | A |
| 20.0412 | 1.374832 | 94.6 | 7.45 | N | N | N |
| 10.2 | 0.7 | -1.0 | 7.46 | A | A | A |
| 10.61 | 0.86 | 3.0 | 8.61 | A | A | A |
| 10.1 | 0.9 | -1.9 | 9.37 | A | A | A |
| 11.2 | 0.8 | 8.7 | 7.71 | A | A | A |
| 8.7 | 0.24 | -15.5 | 4.01 | A | N | W |
| 10.3 | 0.3 | 0.0 | 4.12 | A | A | A |
| 10.19 | 0.64197 | -1.1 | 6.94 | A | A | A |
| 9.78 | 0.435 | -5.0 | 5.32 | A | A | A |
| 9.5 | 1.1 | -7.8 | 11.94 | A | A | A |
| 9.56 | 0.96 | -7.2 | 10.46 | A | A | A |
| 12 | 1 | 16.5 | 8.83 | A | A | A |
| 12.6 | 0.8 | 22.3 | 6.99 | N | N | N |
| 11.95 | 0.98 | 16.0 | 8.70 | A | A | A |
| 10.3 | 1 | 0.0 | 10.14 | A | A | A |
| 11.2 | 0.6 | 8.7 | 6.10 | A | A | A |
| 9.4 | 0.5 | -8.7 | 6.06 | A | A | A |
| 10.7 | 0.6 | 3.9 | 6.32 | A | A | A |
| 10.26 | 0.66 | -0.4 | 7.06 | A | A | A |
| 12 | 0.6 | 16.5 | 5.79 | A | N | W |
| 11.5 | 0.5 | 11.7 | 5.23 | A | A | A |
| 10.7 | 0.6 | 3.9 | 6.32 | A | A | A |
| 10.6 | 0.6 | 2.9 | 6.37 | A | A | A |
| 9.7 | 0.5 | -5.8 | 5.92 | A | A | A |
| 8.834 | 0.4389 | -14.2 | 5.76 | A | A | A |
| 11.55 | 0.7 | 12.1 | 6.72 | A | A | A |
| 10.55 | 0.58 | 2.4 | 6.22 | A | A | A |
| 9.91 | 0.49 | -3.8 | 5.74 | A | A | A |
| 10.88 | 0.89 | 5.6 | 8.68 | A | A | A |
| 8 | 0.6 | -22.3 | 8.05 | N | N | N |
| 10.95 | 1.095 | 6.3 | 10.42 | A | A | A |
| 12 | 0.7 | 16.5 | 6.52 | A | A | A |
| 11.4 | 0.7 | 10.7 | 6.80 | A | A | A |
| 11.873 | 0.127 | 15.3 | 3.10 | A | N | W |
| 10.45 | 0.21 | 1.5 | 3.54 | A | A | A |
| 10.62 | 0.16 | 3.1 | 3.28 | A | A | A |
| 10 | 0.6 | -2.9 | 6.67 | A | A | A |
| 11 | 0.74 | 6.8 | 7.33 | A | A | A |
| 9.67 | 0.61 | -6.1 | 6.95 | A | A | A |
| 11 | 0.4 | 6.8 | 4.66 | A | A | A |
| 10.5 | 0.5 | 1.9 | 5.58 | A | A | A |
| 10.4 | 0.8 | 1.0 | 8.23 | A | A | A |
| 10.3 | 0.71 | 0.0 | 7.48 | A | A | A |
| 9.45 | 0.33 | -8.2 | 4.55 | A | A | A |

Labcode

414

417

419

420

421

425

426

428

429

430

433

434

435

437

440

441

442

443

444

445

446

448

450

452

453

455

459

462

463

464

465

475

477

478

479

480

482

483

484

487

491

492

495

496

498

499

500

501

506

507

508

509

514

516

520

522

523

525

527

528

| Reported value [Bq/kg ] 14.29 | Reported uncertainty (k = 1) [Bq/kg ] 0.86 | Relative bias [%] 38.7 | P-Test [%] 6.69 | Trueness evaluation N | Precision evaluation N | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 10.85 | 0.61 | 5.3 | 6.33 | A | A | A |
| 11.58 | 0.84 | 12.4 | 7.82 | A | A | A |
| 13 | 0.9 | 26.2 | 7.51 | N | N | N |
| 9.4 | 0.6 | -8.7 | 7.02 | A | A | A |
| 11.37 | 0.64 | 10.4 | 6.34 | A | A | A |
| 10 | 0.6 | -2.9 | 6.67 | A | A | A |
| 10.95 | 0.56 | 6.3 | 5.89 | A | A | A |
| 40.15 | 1.22 | 289.8 | 4.21 | N | N | N |
| 9.69 | 0.39 | -5.9 | 4.97 | A | A | A |
| 10.58 | 0.32 | 2.7 | 4.20 | A | A | A |
| 9.86 | 0.44 | -4.3 | 5.33 | A | A | A |
| 10.8 | 0.9 | 4.8 | 8.83 | A | A | A |

Labcode

529

530

532

533

535

536

537

539

540

544

549

553

555

Tl-208 determination of activity concentration in Sample 3

Table 38. Target values

| Parameter | Value |
| --- | --- |
| Activity concentration [ Bq/kg ] | 12.9 |
| Uncertainty [ Bq/kg ] | 0.4 |
| Maximum Acceptable Relative Bias [%] | 20 |

Accepted 40%	Not Accepted 8%	Warning 2%	Not Reported 50%

16

18

Figure 38. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

8

10

12

204

314

355

185

448

57

224

330

102

283

54

58

81

300

441

483

303

271

318

24

166

249

412

232

219

319

369

417

429

437

159

310

555

128

306

79

254

278

371

83

384

211

452

430

114

443

41

304

370

408

256

288

358

276

53

68

72

100

168

445

479

45

215

233

270

312

391

487

537

426

496

25

205

293

281

459

2

21

90

122

134

290

478

202

395

110

235

272

402

465

508

111

151

62

287

316

357

8

291

49

119

124

150

260

299

450

484

242

78

495

403

109

154

344

428

247

218

13

407

341

356

421

64

184

315

328

332

339

393

463

516

382

107

208

347

425

140

536

60

63

178

258

286

348

65

187

398

197

28

15

460

226

325

462

527

11

227

284

191

48

411

55

532

85

161

261

203

482

267

155

499

32

113

94

501

313

133

262

253

66

156

97

220

311

71

333

326

34

179

131

410

376

170

533

467

414

446

Laboratory code

Figure 39. Reported results and evaluation parameters for Tl-208. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 39. Evaluation results for Tl-208 in Sample 3, Rep. Value and Rep. Unc. in Bq/kg

Labcode

Reported value

[Bq/kg ]

Reported uncertainty

(k = 1) [Bq/kg ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

Activity concentration  [ Bq/kg ]

14

2	12.7	1.4	-1.6	11.45	A	A	A

8	12.91	0.91	0.1	7.70	A	A	A

11	14.1	1.03	9.3	7.94	A	A	A

| Reported value [Bq/kg ] 13.3 | Reported uncertainty (k = 1) [Bq/kg ] 0.6 | Relative bias [%] 3.1 | P-Test [%] 5.47 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 13.9 | 0.8 | 7.8 | 6.54 | A | A | A |
| 12.7 | 0.6 | -1.6 | 5.65 | A | A | A |
| 11.3 | 0.73 | -12.4 | 7.17 | A | A | A |
| 12.6 | 0.5 | -2.3 | 5.04 | A | A | A |
| 13.85 | 1.235 | 7.4 | 9.44 | A | A | A |
| 15.09 | 0.91 | 17.0 | 6.78 | A | A | A |
| 19.4 | 1.4 | 50.4 | 7.85 | N | N | N |
| 12.12 | 0.75 | -6.0 | 6.92 | A | A | A |
| 12.47 | 0.55 | -3.3 | 5.39 | A | A | A |
| 14.3 | 0.7 | 10.8 | 5.79 | A | A | A |
| 13 | 2 | 0.8 | 15.69 | A | A | A |
| 12.39 | 0.99 | -4.0 | 8.57 | A | A | A |
| 10.4 | 0.6 | -19.4 | 6.55 | A | N | W |
| 14.31 | 0.68 | 10.9 | 5.67 | A | A | A |
| 9.8 | 1 | -24.0 | 10.66 | N | A | N |
| 10.6 | 0.5 | -17.8 | 5.64 | A | N | W |
| 13.7 | 1.5 | 6.2 | 11.38 | A | A | A |
| 12.9 | 1.1 | 0.0 | 9.07 | A | A | A |
| 13.7 | 0.35 | 6.2 | 4.02 | A | A | A |
| 13.5 | 0.7 | 4.7 | 6.04 | A | A | A |
| 13.8 | 1 | 7.0 | 7.88 | A | A | A |
| 16.78 | 0.992 | 30.1 | 6.68 | N | N | N |
| 12.4 | 0.59 | -3.9 | 5.68 | A | A | A |
| 18.33 | 0.92 | 42.1 | 5.90 | N | N | N |
| 12.4 | 0.45 | -3.9 | 4.77 | A | A | A |
| 13.03 | 0.51 | 1.0 | 4.99 | A | A | A |
| 11.8 | 1.3 | -8.5 | 11.44 | A | A | A |
| 10.61 | 1.645 | -17.8 | 15.81 | A | A | A |
| 11.9 | 1.7 | -7.8 | 14.62 | A | A | A |
| 14.386 | 1.475 | 11.5 | 10.71 | A | A | A |
| 12.7 | 1.2 | -1.6 | 9.94 | A | A | A |
| 15.4 | 1.35 | 19.4 | 9.30 | A | A | A |
| 17.4 | 1.9 | 34.9 | 11.35 | N | N | N |
| 12.4 | 1.1 | -3.9 | 9.40 | A | A | A |
| 9.99 | 0.9 | -22.6 | 9.53 | N | A | N |
| 13.6 | 1.3 | 5.4 | 10.05 | A | A | A |
| 13.2 | 0.2 | 2.3 | 3.45 | A | A | A |
| 12.79 | 0.76 | -0.8 | 6.70 | A | A | A |
| 12.86 | 1.15 | -0.3 | 9.46 | A | A | A |
| 15.32 | 3.83 | 18.8 | 25.19 | A | N | W |
| 12.1 | 1.2 | -6.2 | 10.39 | A | A | A |
| 13 | 1.4 | 0.8 | 11.21 | A | A | A |
| 12.7 | 0.7 | -1.6 | 6.32 | A | A | A |
| 13 | 1 | 0.8 | 8.29 | A | A | A |
| 11.63 | 0.58 | -9.8 | 5.87 | A | A | A |
| 28.0718 | 1.321835 | 117.6 | 5.64 | N | N | N |
| 16 | 2 | 24.0 | 12.88 | N | A | N |
| 12.7 | 0.3 | -1.6 | 3.90 | A | A | A |
| 13.62 | 0.78 | 5.6 | 6.51 | A | A | A |
| 13 | 0.6 | 0.8 | 5.56 | A | A | A |
| 12.88 | 0.9 | -0.2 | 7.64 | A | A | A |
| 13.2 | 0.6 | 2.3 | 5.50 | A | A | A |
| 14.96 | 0.91 | 16.0 | 6.83 | A | A | A |
| 16.83 | 0.42 | 30.5 | 3.98 | N | N | N |
| 11.6 | 0.463 | -10.1 | 5.05 | A | A | A |
| 14.4 | 0.9 | 11.6 | 6.98 | A | A | A |
| 11.3 | 1.72 | -12.4 | 15.53 | A | A | A |
| 12.4 | 1.2 | -3.9 | 10.16 | A | A | A |
| 39 | 3 | 202.3 | 8.29 | N | N | N |

Labcode

13

15

21

24

25

28

32

34

41

45

48

49

53

54

55

57

58

60

62

63

64

65

66

68

71

72

78

79

81

83

85

90

94

97

100

102

107

109

110

111

113

114

119

122

124

128

131

133

134

140

150

151

154

155

156

159

161

166

168

170

| Reported value [Bq/kg ] 13.7 | Reported uncertainty (k = 1) [Bq/kg ] 1.3 | Relative bias [%] 6.2 | P-Test [%] 9.98 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 25.71 | 1.01 | 99.3 | 5.00 | N | N | N |
| 13.5 | 2 | 4.7 | 15.14 | A | A | A |
| 9.32 | 0.5 | -27.8 | 6.20 | N | N | N |
| 13.8 | 1 | 7.0 | 7.88 | A | A | A |
| 14.2 | 1.3 | 10.1 | 9.67 | A | A | A |
| 13.82 | 1.94 | 7.1 | 14.38 | A | A | A |
| 12.72 | 0.21 | -1.4 | 3.51 | A | A | A |
| 14.5 | 1.5 | 12.4 | 10.80 | A | A | A |
| 7 | 0.4 | -45.7 | 6.50 | N | N | N |
| 12.6 | 0.724 | -2.3 | 6.53 | A | A | A |
| 13.6 | 0.97 | 5.4 | 7.78 | A | A | A |
| 11.98 | 0.37 | -7.1 | 4.38 | A | A | A |
| 12.5 | 0.9 | -3.1 | 7.84 | A | A | A |
| 13.27344 | 0.64368 | 2.9 | 5.76 | A | A | A |
| 11.5 | 0.5 | -10.8 | 5.34 | A | A | A |
| 17.4 | 1.2 | 34.9 | 7.56 | N | N | N |
| 9.8 | 0.9 | -24.0 | 9.69 | N | A | N |
| 13.97 | 0.88 | 8.3 | 7.02 | A | A | A |
| 14.1 | 0.5 | 9.3 | 4.71 | A | A | A |
| 11.4 | 0.4 | -11.6 | 4.68 | A | A | A |
| 12.5 | 1 | -3.1 | 8.58 | A | A | A |
| 12.8 | 1.2 | -0.8 | 9.87 | A | A | A |
| 13.02 | 0.35 | 0.9 | 4.10 | A | A | A |
| 13.245 | 1.132 | 2.7 | 9.09 | A | A | A |
| 11.36 | 0.181728 | -11.9 | 3.49 | A | N | W |
| 16.5 | 1.7 | 27.9 | 10.76 | N | N | N |
| 11.8 | 0.9 | -8.5 | 8.23 | A | A | A |
| 12.3 | 0.66 | -4.7 | 6.20 | A | A | A |
| 13.7 | 0.4 | 6.2 | 4.26 | A | A | A |
| 13 | 0.82 | 0.8 | 7.03 | A | A | A |
| 14.4 | 0.5 | 11.6 | 4.66 | A | A | A |
| 16.3 | 0.4 | 26.4 | 3.95 | N | N | N |
| 14.74 | 1.54 | 14.3 | 10.90 | A | A | A |
| 12.5 | 0.5 | -3.1 | 5.06 | A | A | A |
| 11.1 | 1 | -13.9 | 9.53 | A | A | A |
| 12.8 | 0.8 | -0.8 | 6.98 | A | A | A |
| 12.37 | 0.84 | -4.1 | 7.47 | A | A | A |
| 11.8 | 0.7 | -8.5 | 6.69 | A | A | A |
| 12.65 | 0.66 | -1.9 | 6.07 | A | A | A |
| 10.1 | 0.56 | -21.7 | 6.35 | N | N | N |
| 14.125 | 0.96 | 9.5 | 7.47 | A | A | A |
| 13.7 | 1.9 | 6.2 | 14.21 | A | A | A |
| 12.9 | 0.3 | 0.0 | 3.88 | A | A | A |
| 12.3 | 0.7 | -4.7 | 6.48 | A | A | A |
| 12.7 | 0.5 | -1.6 | 5.01 | A | A | A |
| 12.995 | 1.624 | 0.7 | 12.88 | A | A | A |
| 12.6 | 1.4 | -2.3 | 11.54 | A | A | A |
| 13 | 1 | 0.8 | 8.29 | A | A | A |
| 10.8 | 0.25 | -16.3 | 3.87 | A | N | W |
| 11 | 2 | -14.7 | 18.44 | A | A | A |
| 12.2 | 0.5 | -5.4 | 5.14 | A | A | A |
| 11.7 | 1.8 | -9.3 | 15.69 | A | A | A |
| 11.6 | 0.9 | -10.1 | 8.36 | A | A | A |
| 17.4 | 0.765 | 34.9 | 5.38 | N | N | N |
| 12.5 | 0.5 | -3.1 | 5.06 | A | A | A |
| 15.8 | 1.13 | 22.5 | 7.80 | N | N | N |
| 7.6 | 0.21 | -41.1 | 4.15 | N | N | N |
| 13.5 | 0.71 | 4.7 | 6.11 | A | A | A |
| 12.9 | 2 | 0.0 | 15.81 | A | A | A |

Labcode

178

179

184

185

187

191

197

202

203

204

205

208

211

215

218

219

220

224

226

227

232

233

235

242

247

249

253

254

256

258

260

261

262

267

270

271

272

276

278

281

283

284

286

287

288

290

291

293

299

300

303

304

306

310

311

312

313

314

315

316

| Reported value [Bq/kg ] 11.22 | Reported uncertainty (k = 1) [Bq/kg ] 0.25 | Relative bias [%] -13.0 | P-Test [%] 3.82 | Trueness evaluation A | Precision evaluation N | Final Score W |
| --- | --- | --- | --- | --- | --- | --- |
| 11.5 | 0.5 | -10.8 | 5.34 | A | A | A |
| 14 | 1.19 | 8.5 | 9.05 | A | A | A |
| 18.83 | 1 | 46.0 | 6.15 | N | N | N |
| 13.5 | 2 | 4.7 | 15.14 | A | A | A |
| 9.96 | 0.58 | -22.8 | 6.60 | N | N | N |
| 13.5 | 1.1 | 4.7 | 8.72 | A | A | A |
| 18.57 | 1.56 | 44.0 | 8.95 | N | N | N |
| 13.5 | 0.8 | 4.7 | 6.69 | A | A | A |
| 13.4 | 0.5 | 3.9 | 4.85 | A | A | A |
| 13.2 | 1.1 | 2.3 | 8.89 | A | A | A |
| 13.6 | 0.75 | 5.4 | 6.33 | A | A | A |
| 13.7 | 0.7 | 6.2 | 5.98 | A | A | A |
| 9.03 | 0.2 | -30.0 | 3.81 | N | N | N |
| 13.4 | 0.8 | 3.9 | 6.73 | A | A | A |
| 12.9 | 1 | 0.0 | 8.35 | A | A | A |
| 12.3 | 1.2 | -4.7 | 10.24 | A | A | A |
| 11.5 | 0.43 | -10.8 | 4.86 | A | A | A |
| 12.2 | 0.7 | -5.4 | 6.52 | A | A | A |
| 11.89 | 1.19 | -7.8 | 10.48 | A | A | A |
| 37.9 | 1.5 | 193.8 | 5.03 | N | N | N |
| 13.59 | 0.87 | 5.3 | 7.11 | A | A | A |
| 11.9 | 1.3 | -7.8 | 11.36 | A | A | A |
| 12.5 | 0.3 | -3.1 | 3.92 | A | A | A |
| 13.5 | 0.8 | 4.7 | 6.69 | A | A | A |
| 12.76 | 1.47 | -1.1 | 11.93 | A | A | A |
| 13.8 | 2.1 | 7.0 | 15.53 | A | A | A |
| 12.8 | 0.8 | -0.8 | 6.98 | A | A | A |
| 13.19 | 0.84 | 2.2 | 7.08 | A | A | A |
| 13.3 | 0.8 | 3.1 | 6.77 | A | A | A |
| 12.2 | 0.4 | -5.4 | 4.51 | A | A | A |
| 34.3 | 0.8 | 165.9 | 3.88 | N | N | N |
| 14.3 | 0.7 | 10.8 | 5.79 | A | A | A |
| 11.36 | 1.56 | -11.9 | 14.08 | A | A | A |
| 52.2 | 3.2 | 304.6 | 6.87 | N | N | N |
| 11.5 | 0.6 | -10.8 | 6.07 | A | A | A |
| 13.4 | 1 | 3.9 | 8.08 | A | A | A |
| 13.6 | 0.9 | 5.4 | 7.31 | A | A | A |
| 12.52 | 1.39 | -3.0 | 11.53 | A | A | A |
| 13.2 | 1.5 | 2.3 | 11.78 | A | A | A |
| 11.5 | 0.343 | -10.8 | 4.30 | A | A | A |
| 12.03 | 0.72 | -6.7 | 6.74 | A | A | A |
| 11.5 | 1.3 | -10.8 | 11.72 | A | A | A |
| 10.9 | 0.7 | -15.5 | 7.13 | A | A | A |
| 12.1 | 0.7 | -6.2 | 6.56 | A | A | A |
| 12.4 | 1.1 | -3.9 | 9.40 | A | A | A |
| 69.5 | 4.7 | 438.8 | 7.44 | N | N | N |
| 9.32 | 0.5 | -27.8 | 6.20 | N | N | N |
| 13 | 0.4 | 0.8 | 4.37 | A | A | A |
| 11.98 | 0.70682 | -7.1 | 6.67 | A | A | A |
| 12.68 | 1.02 | -1.7 | 8.62 | A | A | A |
| 13.9 | 1 | 7.8 | 7.83 | A | A | A |
| 14 | 1 | 8.5 | 7.79 | A | A | A |
| 13.5 | 0.8 | 4.7 | 6.69 | A | A | A |
| 12.8 | 1.3 | -0.8 | 10.62 | A | A | A |
| 46.78 | 11.6 | 262.6 | 24.99 | N | N | N |
| 12.7 | 0.6 | -1.6 | 5.65 | A | A | A |
| 12.43 | 1.16 | -3.6 | 9.83 | A | A | A |
| 14.5 | 1.1 | 12.4 | 8.20 | A | A | A |
| 10.9 | 0.7 | -15.5 | 7.13 | A | A | A |

Labcode

318

319

325

326

328

330

332

333

339

341

344

347

348

355

356

357

358

369

370

371

376

382

384

391

393

395

398

402

403

407

408

410

411

412

414

417

421

425

426

428

429

430

437

441

443

445

446

448

450

452

459

460

462

463

465

467

478

479

482

483

| Reported value [Bq/kg ] 13 | Reported uncertainty (k = 1) [Bq/kg ] 0.7 | Relative bias [%] 0.8 | P-Test [%] 6.21 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 12.5 | 0.5 | -3.1 | 5.06 | A | A | A |
| 13.06 | 0.5 | 1.2 | 4.93 | A | A | A |
| 12.58 | 0.51 | -2.5 | 5.10 | A | A | A |
| 15 | 1.5 | 16.3 | 10.47 | A | A | A |
| 15.6 | 0.8 | 20.9 | 5.99 | N | N | N |
| 12.81 | 0.26 | -0.7 | 3.71 | A | A | A |
| 13.5 | 0.88 | 4.7 | 7.22 | A | A | A |
| 14 | 0.9 | 8.5 | 7.14 | A | A | A |
| 14.36 | 1.02 | 11.3 | 7.75 | A | A | A |
| 44.7 | 2.4 | 246.5 | 6.20 | N | N | N |
| 13.63 | 0.72 | 5.7 | 6.13 | A | A | A |
| 12.5 | 0.5 | -3.1 | 5.06 | A | A | A |
| 11.6 | 1.2 | -10.1 | 10.80 | A | A | A |

Labcode

484

487

495

496

499

501

508

516

527

532

533

536

537

555

Pb-210 determination of activity concentration in Sample 3

Table 40. Target values

| Parameter | Value |
| --- | --- |
| Activity concentration [ Bq/kg ] | 43.4 |
| Uncertainty [ Bq/kg ] | 1.7 |
| Maximum Acceptable Relative Bias [%] | 20 |

Accepted 26%	Not Accepted 11%	Warning 4%	Not Reported 59%

55

60

Figure 40. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

25

30

35

40

397

64

131

20

555

163

350

300

377

412

34

165

437

530

446

525

168

475

165

315

230

379

204

348

207

330

102

294

232

45

536

443

322

272

291

319

527

55

151

49

79

318

304

441

483

258

421

537

498

175

192

370

32

155

2

60

459

124

356

242

119

128

295

214

293

217

363

218

215

509

528

110

452

276

202

433

382

122

152

408

159

25

227

53

72

90

407

48

97

314

21

166

235

391

66

425

496

341

11

247

299

465

127

288

313

121

63

260

393

161

487

347

312

401

553

41

15

156

191

270

290

58

19

411

428

326

68

516

170

371

203

529

357

262

479

78

532

220

231

460

373

403

226

307

317

405

462

24

111

278

283

38

328

414

238

398

501

52

484

533

249

286

Laboratory code

Figure 41. Reported results and evaluation parameters for Pb-210. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 41. Evaluation results for Pb-210 in Sample 3, Rep. Value and Rep. Unc. in Bq/kg

Labcode

Reported value

[Bq/kg ]

Reported uncertainty

(k = 1) [Bq/kg ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

Activity concentration  [ Bq/kg ]

45

50

2	40.8	4	-6.0	10.56	A	A	A

11	45	4.88	3.7	11.53	A	A	A

15	46.6	2.9	7.4	7.35	A	A	A

| Reported value [Bq/kg ] 47.5 | Reported uncertainty (k = 1) [Bq/kg ] 3.8 | Relative bias [%] 9.4 | P-Test [%] 8.91 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 25.7 | 4.99 | -40.8 | 19.81 | N | A | N |
| 44 | 16 | 1.4 | 36.57 | A | N | W |
| 54.2 | 8.2 | 24.9 | 15.63 | N | A | N |
| 43.2 | 6.3 | -0.5 | 15.10 | A | A | A |
| 40.56 | 6.94 | -6.5 | 17.55 | A | A | A |
| 31.1 | 9 | -28.3 | 29.20 | N | N | N |
| 56.85 | 17.83 | 31.0 | 31.61 | N | N | N |
| 46.54 | 7.41 | 7.2 | 16.40 | A | A | A |
| 36.88 | 4.25 | -15.0 | 12.17 | A | A | A |
| 43.9 | 5.9 | 1.1 | 14.00 | A | A | A |
| 39 | 7 | -10.1 | 18.37 | A | A | A |
| 58.8 | 8.22 | 35.5 | 14.52 | N | A | N |
| 43.48 | 4.35 | 0.2 | 10.74 | A | A | A |
| 38.46 | 10.94 | -11.4 | 28.71 | A | N | W |
| 47.3 | 6 | 9.0 | 13.28 | A | A | A |
| 40.8 | 5.7 | -6.0 | 14.51 | A | A | A |
| 45.41 | 1.97 | 4.6 | 5.84 | A | A | A |
| 24.1 | 2.1 | -44.5 | 9.55 | N | N | N |
| 44.11 | 3.357 | 1.6 | 8.56 | A | A | A |
| 48.9 | 10.4 | 12.7 | 21.63 | A | N | W |
| 43.7 | 2.21 | 0.7 | 6.40 | A | A | A |
| 51.51 | 3.1 | 18.7 | 7.18 | A | N | W |
| 39 | 3.4 | -10.1 | 9.56 | A | A | A |
| 43.7 | 4.6 | 0.7 | 11.23 | A | A | A |
| 43.9 | 5.1 | 1.1 | 12.26 | A | A | A |
| 36.6 | 5.1 | -15.7 | 14.47 | A | A | A |
| 42.7 | 2.7 | -1.6 | 7.44 | A | A | A |
| 55.31 | 7.31 | 27.4 | 13.78 | N | A | N |
| 41.31 | 5.3 | -4.8 | 13.41 | A | A | A |
| 45.2 | 4.6 | 4.2 | 10.90 | A | A | A |
| 43 | 4 | -0.9 | 10.09 | A | A | A |
| 41 | 15 | -5.5 | 36.79 | A | N | W |
| 45.1 | 3.7 | 3.9 | 9.09 | A | A | A |
| 41.51 | 2.36 | -4.3 | 6.90 | A | A | A |
| 24.16666 | 2.690446 | -44.3 | 11.80 | N | N | N |
| 38.8 | 3.53 | -10.6 | 9.91 | A | A | A |
| 43 | 6.4 | -0.9 | 15.39 | A | A | A |
| 40.7 | 5 | -6.2 | 12.89 | A | A | A |
| 46.6 | 1.7 | 7.4 | 5.35 | A | A | A |
| 43.1 | 3.63 | -0.7 | 9.29 | A | A | A |
| 46 | 5.5 | 6.0 | 12.58 | A | A | A |
| 28 | 4 | -35.5 | 14.81 | N | A | N |
| 31.57 | 8.5 | -27.3 | 27.21 | N | N | N |
| 34 | 2 | -21.7 | 7.07 | N | N | N |
| 44 | 19 | 1.4 | 43.36 | A | N | W |
| 33 | 4.9 | -24.0 | 15.36 | N | A | N |
| 50 | 4 | 15.2 | 8.91 | A | A | A |
| 40 | 2.4 | -7.8 | 7.17 | A | A | A |
| 46.8 | 4.9 | 7.8 | 11.18 | A | A | A |
| 40 | 8 | -7.8 | 20.38 | A | N | W |
| 42.91 | 1.9 | -1.1 | 5.91 | A | A | A |
| 50.4 | 13.1 | 16.1 | 26.29 | A | N | W |
| 35 | 4 | -19.4 | 12.08 | A | A | A |
| 35.6 | 3.2 | -18.0 | 9.81 | A | A | A |
| 42 | 3 | -3.2 | 8.15 | A | A | A |
| 42.5 | 5.2 | -2.1 | 12.85 | A | A | A |
| 42.3 | 9 | -2.5 | 21.63 | A | N | W |
| 42.3025 | 4.83934 | -2.5 | 12.09 | A | A | A |
| 52 | 13 | 19.8 | 25.31 | A | N | W |

Labcode

19

20

21

24

25

32

34

38

41

45

48

49

52

53

55

58

60

63

64

66

68

72

78

79

90

97

102

110

111

119

121

122

124

127

128

131

151

152

155

156

159

161

163

165

165

166

168

170

175

191

192

202

203

204

207

214

215

217

218

220

| Reported value [Bq/kg ] 53.34 | Reported uncertainty (k = 1) [Bq/kg ] 8.32 | Relative bias [%] 22.9 | P-Test [%] 16.08 | Trueness evaluation N | Precision evaluation A | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 43.4 | 2.4 | 0.0 | 6.78 | A | A | A |
| 34.3 | 4.6 | -21.0 | 13.97 | N | A | N |
| 52 | 10 | 19.8 | 19.63 | A | A | A |
| 36.7 | 5.1 | -15.4 | 14.44 | A | A | A |
| 44 | 4.4 | 1.4 | 10.74 | A | A | A |
| 57.4 | 7.6 | 32.3 | 13.81 | N | A | N |
| 41.24 | 2.02 | -5.0 | 6.27 | A | A | A |
| 45 | 7.514 | 3.7 | 17.15 | A | A | A |
| 78.79 | 1.763597 | 81.5 | 4.51 | N | N | N |
| 39.6 | 10 | -8.8 | 25.55 | A | N | W |
| 45.6 | 3.5 | 5.1 | 8.62 | A | A | A |
| 51 | 7 | 17.5 | 14.27 | A | A | A |
| 47 | 4 | 8.3 | 9.37 | A | A | A |
| 38.2 | 2.4 | -12.0 | 7.40 | A | A | A |
| 42.82 | 3.35 | -1.3 | 8.75 | A | A | A |
| 55.9 | 12 | 28.8 | 21.82 | N | N | N |
| 56.1 | 3.21 | 29.3 | 6.93 | N | N | N |
| 122.8 | 58.9 | 182.9 | 48.12 | N | N | N |
| 45.1 | 6.2 | 3.9 | 14.29 | A | A | A |
| 47 | 4 | 8.3 | 9.37 | A | A | A |
| 38.219 | 7.45 | -11.9 | 19.88 | A | A | A |
| 42.1 | 6.9 | -3.0 | 16.85 | A | A | A |
| 36.6 | 4 | -15.7 | 11.61 | A | A | A |
| 41.56 | 3.07 | -4.2 | 8.36 | A | A | A |
| 45 | 7 | 3.7 | 16.04 | A | A | A |
| 28.06 | 1.92 | -35.4 | 7.88 | N | N | N |
| 39.1 | 4.8 | -9.9 | 12.89 | A | A | A |
| 53.4 | 4.5 | 23.0 | 9.29 | N | A | N |
| 46.3 | 5.2 | 6.7 | 11.89 | A | A | A |
| 45.1 | 5.36 | 3.9 | 12.51 | A | A | A |
| 43.92 | 4.62 | 1.2 | 11.22 | A | A | A |
| 34.15 | 3.63 | -21.3 | 11.33 | N | A | N |
| 53.4 | 2 | 23.0 | 5.42 | N | N | N |
| 39.02 | 1.51 | -10.1 | 5.51 | A | A | A |
| 38.4 | 2.7 | -11.5 | 8.05 | A | A | A |
| 37.6 | 8.2 | -13.4 | 22.16 | A | N | W |
| 48.51 | 4.6 | 11.8 | 10.26 | A | A | A |
| 57 | 10 | 31.3 | 17.98 | N | A | N |
| 36.1 | 1.63 | -16.8 | 5.98 | A | N | W |
| 44.8 | 6.2 | 3.2 | 14.38 | A | A | A |
| 46.1 | 2.5 | 6.2 | 6.69 | A | A | A |
| 35 | 6 | -19.4 | 17.58 | A | A | A |
| 28 | 2 | -35.5 | 8.15 | N | N | N |
| 41.2 | 4.2 | -5.1 | 10.92 | A | A | A |
| 50.8 | 3 | 17.0 | 7.09 | A | A | A |
| 42.3 | 4.3 | -2.5 | 10.89 | A | A | A |
| 40.5 | 3.1 | -6.7 | 8.60 | A | A | A |
| 50.15 | 5.43 | 15.6 | 11.51 | A | A | A |
| 52.3 | 4 | 20.5 | 8.59 | N | A | N |
| 29.20244 | 2.255174 | -32.7 | 8.66 | N | N | N |
| 34.8 | 4.1 | -19.8 | 12.42 | A | A | A |
| 42.93 | 4.72 | -1.1 | 11.67 | A | A | A |
| 44.06 | 2.17 | 1.5 | 6.29 | A | A | A |
| 45.8 | 2.4 | 5.5 | 6.54 | A | A | A |
| 22.28 | 2.78 | -48.7 | 13.08 | N | N | N |
| 58 | 10 | 33.6 | 17.68 | N | A | N |
| 46.3 | 7.4 | 6.7 | 16.46 | A | A | A |
| 53.08 | 8.71 | 22.3 | 16.87 | N | A | N |
| 53.9 | 3.5 | 24.2 | 7.58 | N | N | N |

Labcode

226

227

230

231

232

235

238

242

247

249

258

260

262

270

272

276

278

283

286

288

290

291

293

294

295

299

300

304

307

312

313

314

315

317

318

319

322

326

328

330

341

347

348

350

356

357

363

370

371

373

377

379

382

391

393

397

398

401

403

405

| Reported value [Bq/kg ] 43.8 | Reported uncertainty (k = 1) [Bq/kg ] 2.8 | Relative bias [%] 0.9 | P-Test [%] 7.50 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 43 | 3 | -0.9 | 8.00 | A | A | A |
| 47.8 | 3.3 | 10.1 | 7.94 | A | A | A |
| 30.66 | 9.94 | -29.4 | 32.66 | N | N | N |
| 57 | 8.6 | 31.3 | 15.59 | N | A | N |
| 39.6 | 4.1 | -8.8 | 11.07 | A | A | A |
| 44.3 | 4.1 | 2.1 | 10.05 | A | A | A |
| 48.4 | 5 | 11.5 | 11.05 | A | A | A |
| 42.91 | 3.04 | -1.1 | 8.10 | A | A | A |
| 31.6 | 7.3 | -27.2 | 23.43 | N | N | N |
| 39.4 | 3.5 | -9.2 | 9.71 | A | A | A |
| 37 | 5 | -14.8 | 14.07 | A | A | A |
| 31.8 | 14.1 | -26.7 | 44.51 | N | N | N |
| 42.74 | 3.03454 | -1.5 | 8.11 | A | A | A |
| 40.89 | 3.83 | -5.8 | 10.15 | A | A | A |
| 52.2 | 6.3 | 20.3 | 12.69 | N | A | N |
| 54 | 6 | 24.4 | 11.78 | N | A | N |
| 45 | 4.4 | 3.7 | 10.53 | A | A | A |
| 33.8 | 2.9 | -22.1 | 9.43 | N | A | N |
| 51.22 | 6.27 | 18.0 | 12.85 | A | A | A |
| 39.4 | 3.5 | -9.2 | 9.71 | A | A | A |
| 67.6 | 6.6 | 55.8 | 10.52 | N | N | N |
| 46 | 5 | 6.0 | 11.55 | A | A | A |
| 44.67 | 6.15 | 2.9 | 14.31 | A | A | A |
| 39.95 | 3.95 | -8.0 | 10.63 | A | A | A |
| 58.3 | 8.3 | 34.3 | 14.77 | N | A | N |
| 42.5 | 1.1 | -2.1 | 4.69 | A | A | A |
| 48.9 | 8 | 12.7 | 16.82 | A | A | A |
| 32.8 | 5.3 | -24.4 | 16.63 | N | A | N |
| 38.4 | 7 | -11.5 | 18.65 | A | A | A |
| 42.5 | 8.8 | -2.1 | 21.07 | A | N | W |
| 50.4 | 3.9 | 16.1 | 8.67 | A | A | A |
| 31.7 | 9.94 | -27.0 | 31.60 | N | N | N |
| 51.71 | 7.74 | 19.1 | 15.47 | A | A | A |
| 71 | 6.1 | 63.6 | 9.44 | N | N | N |
| 36.9 | 3.2 | -15.0 | 9.52 | A | A | A |
| 39.6 | 2.8 | -8.8 | 8.08 | A | A | A |
| 46.3 | 3.3 | 6.7 | 8.13 | A | A | A |
| 26.6 | 5.3 | -38.7 | 20.31 | N | N | N |

Labcode

407

408

411

412

414

421

425

428

433

437

441

443

446

452

459

460

462

465

475

479

483

484

487

496

498

501

509

516

525

527

528

529

530

532

533

536

537

553

555

Bi-212 determination of activity concentration in Sample 3

Table 42. Target values

| Parameter | Value |
| --- | --- |
| Activity concentration [ Bq/kg ] | 39.6 |
| Uncertainty [ Bq/kg ] | 1.7 |
| Maximum Acceptable Relative Bias [%] | 20 |

Accepted 27%	Not Accepted 9%	Warning 3%	Not Reported 61%

50

55

Figure 42. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

25

30

35

37

429

500

400

102

283

496

444

446

530

464

21

32

218

426

463

57

224

172

131

417

330

516

369

303

522

41

268

491

319

19

437

34

459

165

204

72

293

421

233

249

555

100

452

272

441

483

395

478

328

357

484

25

159

287

316

358

154

260

288

114

40

2

217

376

445

527

479

403

344

398

111

202

341

384

537

379

382

90

184

339

107

332

408

178

364

410

78

267

128

55

64

140

460

150

198

13

15

68

109

325

348

450

462

506

381

83

254

253

304

425

318

276

430

465

58

247

54

110

52

191

300

223

258

281

315

291

49

273

299

85

313

81

28

113

197

63

226

124

433

66

262

133

155

284

11

179

278

97

414

495

203

326

286

333

Laboratory code

Figure 43. Reported results and evaluation parameters for Bi-212. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 43. Evaluation results for Bi-212 in Sample 3, Rep. Value and Rep. Unc. in Bq/kg

Labcode

Reported value

[Bq/kg ]

Reported uncertainty

(k = 1) [Bq/kg ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

Activity concentration  [ Bq/kg ]

40

45

2	36.5	2.1	-7.8	7.18	A	A	A

11	47.3	4.51	19.4	10.46	A	A	A

13	38.9	4.1	-1.8	11.38	A	A	A

| Reported value [Bq/kg ] 38.9 | Reported uncertainty (k = 1) [Bq/kg ] 2.1 | Relative bias [%] -1.8 | P-Test [%] 6.90 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 33.4 | 7 | -15.7 | 21.39 | A | N | W |
| 23 | 2 | -41.9 | 9.70 | N | N | N |
| 35.9 | 1.4 | -9.3 | 5.80 | A | A | A |
| 42.55 | 4.349 | 7.4 | 11.09 | A | A | A |
| 23.29 | 4.03 | -41.2 | 17.83 | N | A | N |
| 33.6 | 3.1 | -15.2 | 10.18 | A | A | A |
| 0.16 | 0.06 | -99.6 | 37.74 | N | N | N |
| 36.46 | 3.31 | -7.9 | 10.04 | A | A | A |
| 32.2 | 3.15 | -18.7 | 10.68 | A | A | A |
| 41 | 5 | 3.5 | 12.93 | A | A | A |
| 40.5 | 1.72 | 2.3 | 6.04 | A | A | A |
| 40.1 | 2.6 | 1.3 | 7.78 | A | A | A |
| 38.6 | 3.25 | -2.5 | 9.45 | A | A | A |
| 27.1 | 2.7 | -31.6 | 10.85 | N | N | N |
| 40 | 2.5 | 1.0 | 7.58 | A | A | A |
| 43.3 | 2.27 | 9.3 | 6.78 | A | A | A |
| 38.6 | 4.5 | -2.5 | 12.42 | A | A | A |
| 44.97 | 2.4045 | 13.6 | 6.86 | A | A | A |
| 38.9 | 4.16 | -1.8 | 11.52 | A | A | A |
| 34.4 | 1.89 | -13.1 | 6.97 | A | A | A |
| 38.34 | 9.3 | -3.2 | 24.63 | A | N | W |
| 42.45 | 11.75 | 7.2 | 28.01 | A | N | W |
| 39.1 | 5.7 | -1.3 | 15.20 | A | A | A |
| 41.533 | 1.787 | 4.9 | 6.08 | A | A | A |
| 37.7 | 3.5 | -4.8 | 10.23 | A | A | A |
| 51.3 | 5.9 | 29.6 | 12.28 | N | A | N |
| 34.9 | 2.1 | -11.9 | 7.39 | A | A | A |
| 19.4 | 2.1 | -51.0 | 11.64 | N | N | N |
| 38 | 3.1 | -4.0 | 9.22 | A | A | A |
| 39 | 2 | -1.5 | 6.69 | A | A | A |
| 40.2 | 2.6 | 1.5 | 7.76 | A | A | A |
| 37.11 | 4.7 | -6.3 | 13.37 | A | A | A |
| 42.55 | 10.64 | 7.4 | 25.37 | A | N | W |
| 36.4 | 3.6 | -8.1 | 10.78 | A | A | A |
| 44 | 5 | 11.1 | 12.15 | A | A | A |
| 38.56 | 2.04 | -2.6 | 6.81 | A | A | A |
| 28.36015 | 4.493641 | -28.4 | 16.42 | N | A | N |
| 46 | 11 | 16.2 | 24.30 | A | N | W |
| 38.6 | 3.6 | -2.5 | 10.27 | A | A | A |
| 38.7 | 2.8 | -2.3 | 8.41 | A | A | A |
| 36.1 | 2.1 | -8.8 | 7.23 | A | A | A |
| 46 | 5.4 | 16.2 | 12.50 | A | A | A |
| 36 | 0.936 | -9.1 | 5.02 | A | A | A |
| 34 | 2 | -14.1 | 7.28 | A | A | A |
| 28 | 6 | -29.3 | 21.85 | N | N | N |
| 38.1 | 4.6 | -3.8 | 12.81 | A | A | A |
| 50.16 | 0.56 | 26.7 | 4.44 | N | N | N |
| 37.8 | 5.7 | -4.6 | 15.68 | A | A | A |
| 40.5 | 5.4 | 2.3 | 14.01 | A | A | A |
| 42.62 | 6.82 | 7.6 | 16.57 | A | A | A |
| 38.8 | 4 | -2.0 | 11.17 | A | A | A |
| 37.15 | 1.15 | -6.2 | 5.29 | A | A | A |
| 58.4 | 9.9 | 47.5 | 17.49 | N | N | N |
| 34 | 4 | -14.1 | 12.52 | A | A | A |
| 36.5 | 5 | -7.8 | 14.36 | A | A | A |
| 23.40725 | 1.67791 | -40.9 | 8.36 | N | N | N |
| 40.6 | 4.1 | 2.5 | 10.97 | A | A | A |
| 27.8 | 2.5 | -29.8 | 9.96 | N | N | N |
| 43.49 | 2.88 | 9.8 | 7.89 | A | A | A |

Labcode

15

19

21

25

28

32

34

37

40

41

49

52

54

55

57

58

63

64

66

68

72

78

81

83

85

90

97

100

102

107

109

110

111

113

114

124

128

131

133

140

150

154

155

159

165

172

178

179

184

191

197

198

202

203

204

217

218

223

224

226

| Reported value [Bq/kg ] 34.5 | Reported uncertainty (k = 1) [Bq/kg ] 2.7 | Relative bias [%] -12.9 | P-Test [%] 8.93 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 40.087 | 3.867 | 1.2 | 10.56 | A | A | A |
| 34.57 | 0.743388 | -12.7 | 4.80 | A | N | W |
| 39.3 | 5.6 | -0.8 | 14.88 | A | A | A |
| 39.18 | 19.36 | -1.1 | 49.60 | A | N | W |
| 40.6 | 2.7 | 2.5 | 7.92 | A | A | A |
| 36.2 | 3.2 | -8.6 | 9.83 | A | A | A |
| 45.3 | 1.1 | 14.4 | 4.93 | A | N | W |
| 38.44 | 4.61 | -2.9 | 12.74 | A | A | A |
| 32.5 | 1.63 | -17.9 | 6.60 | A | N | W |
| 35.1 | 2.2 | -11.4 | 7.60 | A | A | A |
| 41 | 3 | 3.5 | 8.48 | A | A | A |
| 39.74 | 2.98 | 0.3 | 8.64 | A | A | A |
| 50.9 | 6 | 28.5 | 12.55 | N | A | N |
| 40.6 | 6.5 | 2.5 | 16.58 | A | A | A |
| 20.3 | 1.21 | -48.7 | 7.35 | N | N | N |
| 46.02 | 2.65 | 16.2 | 7.18 | A | A | A |
| 61.8 | 20.3 | 56.1 | 33.13 | N | N | N |
| 36 | 4.4 | -9.1 | 12.95 | A | A | A |
| 36.2 | 3.87 | -8.6 | 11.52 | A | A | A |
| 40.9 | 5.113 | 3.3 | 13.22 | A | A | A |
| 34.4 | 5.9 | -13.1 | 17.68 | A | A | A |
| 41 | 4 | 3.5 | 10.66 | A | A | A |
| 40.52 | 2.75 | 2.3 | 8.03 | A | A | A |
| 31 | 4 | -21.7 | 13.60 | N | A | N |
| 39.3 | 3.8 | -0.8 | 10.58 | A | A | A |
| 41.6 | 3.58 | 5.0 | 9.62 | A | A | A |
| 40.825 | 3.4 | 3.1 | 9.37 | A | A | A |
| 36 | 6 | -9.1 | 17.21 | A | A | A |
| 39.54 | 1.23 | -0.1 | 5.30 | A | A | A |
| 33.1 | 1.4 | -16.4 | 6.03 | A | N | W |
| 39 | 3.32 | -1.5 | 9.53 | A | A | A |
| 61.57 | 6.9 | 55.5 | 12.00 | N | N | N |
| 35.8 | 4 | -9.6 | 11.97 | A | A | A |
| 30.06 | 1.6 | -24.1 | 6.84 | N | N | N |
| 38 | 3.5 | -4.0 | 10.16 | A | A | A |
| 72.76 | 15.82 | 83.7 | 22.16 | N | N | N |
| 37.9 | 1.2 | -4.3 | 5.33 | A | A | A |
| 37.3 | 1.2 | -5.8 | 5.36 | A | A | A |
| 37.1 | 2.8 | -6.3 | 8.68 | A | A | A |
| 39 | 2 | -1.5 | 6.69 | A | A | A |
| 35.8 | 2 | -9.6 | 7.05 | A | A | A |
| 36 | 3.6 | -9.1 | 10.88 | A | A | A |
| 38.1 | 2.5 | -3.8 | 7.84 | A | A | A |
| 30.9 | 2.56 | -22.0 | 9.33 | N | A | N |
| 36.5 | 1.5 | -7.8 | 5.94 | A | A | A |
| 37.4 | 4.4 | -5.6 | 12.52 | A | A | A |
| 39.08 | 2.49 | -1.3 | 7.68 | A | A | A |
| 37.46 | 2.61 | -5.4 | 8.18 | A | A | A |
| 37.3 | 4.4 | -5.8 | 12.55 | A | A | A |
| 35.45 | 4.18 | -10.5 | 12.55 | A | A | A |
| 37.1 | 4.2 | -6.3 | 12.11 | A | A | A |
| 19 | 9 | -52.0 | 47.56 | N | N | N |
| 36.96 | 2.25 | -6.7 | 7.45 | A | A | A |
| 38 | 2 | -4.0 | 6.79 | A | A | A |
| 38.1 | 2.3 | -3.8 | 7.41 | A | A | A |
| 53.2 | 5.8 | 34.3 | 11.72 | N | N | N |
| 29.9 | 2.5 | -24.5 | 9.40 | N | N | N |
| 34.4 | 1.9 | -13.1 | 7.00 | A | A | A |
| 39.5 | 3.7 | -0.2 | 10.30 | A | A | A |

Labcode

233

247

249

253

254

258

260

262

267

268

272

273

276

278

281

283

284

286

287

288

291

293

299

300

303

304

313

315

316

318

319

325

326

328

330

332

333

339

341

344

348

357

358

364

369

376

379

381

382

384

395

398

400

403

408

410

414

417

421

425

| Reported value [Bq/kg ] 24.97 | Reported uncertainty (k = 1) [Bq/kg ] 3.97 | Relative bias [%] -36.9 | P-Test [%] 16.47 | Trueness evaluation N | Precision evaluation A | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 17.1 | 2.63 | -56.8 | 15.97 | N | N | N |
| 39.8 | 3.5 | 0.5 | 9.79 | A | A | A |
| 44.67 | 1.05 | 12.8 | 4.89 | A | N | W |
| 33.5 | 5.1 | -15.4 | 15.82 | A | A | A |
| 35.4 | 1.9 | -10.6 | 6.87 | A | A | A |
| 20.57 | 1.7 | -48.1 | 9.31 | N | N | N |
| 36.6 | 2.1 | -7.6 | 7.17 | A | A | A |
| 21.4 | 1.9 | -46.0 | 9.86 | N | N | N |
| 39 | 3 | -1.5 | 8.81 | A | A | A |
| 34.93 | 2.02594 | -11.8 | 7.22 | A | A | A |
| 33.97 | 3.01 | -14.2 | 9.85 | A | A | A |
| 38.6 | 5 | -2.5 | 13.65 | A | A | A |
| 39 | 6 | -1.5 | 15.97 | A | A | A |
| 27 | 3 | -31.8 | 11.91 | N | N | N |
| 22.73 | 2.93 | -42.6 | 13.59 | N | N | N |
| 39.8 | 4 | 0.5 | 10.93 | A | A | A |
| 35.7 | 1.9 | -9.8 | 6.84 | A | A | A |
| 36.89 | 5.48 | -6.8 | 15.46 | A | A | A |
| 35.4 | 1.9 | -10.6 | 6.87 | A | A | A |
| 35.8 | 5.5 | -9.6 | 15.95 | A | A | A |
| 32.756 | 2.726 | -17.3 | 9.36 | A | A | A |
| 54.06 | 6.79 | 36.5 | 13.27 | N | N | N |
| 20.41 | 2.57 | -48.5 | 13.30 | N | N | N |
| 17.88 | 8.16 | -54.9 | 45.84 | N | N | N |
| 39 | 4 | -1.5 | 11.12 | A | A | A |
| 30.7 | 4.4 | -22.5 | 14.96 | N | A | N |
| 31 | 1.2 | -21.7 | 5.78 | N | N | N |
| 36.6 | 4.7 | -7.6 | 13.54 | A | A | A |
| 21.56 | 3.21 | -45.6 | 15.50 | N | N | N |
| 37.3 | 1.9 | -5.8 | 6.66 | A | A | A |
| 34.8 | 7 | -12.1 | 20.57 | A | N | W |

Labcode

426

429

430

433

437

441

444

445

446

450

452

459

460

462

463

464

465

478

479

483

484

491

495

496

500

506

516

522

527

530

537

555

Pb-212 determination of activity concentration in Sample 3

Table 44. Target values

| Parameter | Value |
| --- | --- |
| Activity concentration [ Bq/kg ] | 36.8 |
| Uncertainty [ Bq/kg ] | 1.1 |
| Maximum Acceptable Relative Bias [%] | 20 |

Accepted 43%	Not Accepted 8%	Warning 3%	Not Reported 46%

45

50

Figure 44. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

25

30

37

94

84

283

464

400

496

355

224

145

417

34

314

412

166

185

448

330

369

453

530

303

499

522

293

25

151

437

268

131

384

168

232

319

24

41

45

429

444

484

391

50

140

68

72

233

408

249

358

459

40

452

105

132

479

122

290

487

272

64

441

483

395

102

270

278

328

55

111

478

276

78

58

159

316

508

256

154

79

260

288

312

410

465

104

114

235

411

197

2

217

445

213

398

242

156

258

357

364

403

204

462

527

371

202

60

110

150

344

116

90

341

537

555

215

450

382

379

247

425

184

356

430

83

339

421

63

107

506

347

170

178

347

32

393

393

315

286

516

81

128

13

426

62

428

71

198

15

376

211

299

325

348

460

381

109

119

373

318

402

191

261

532

267

291

304

49

253

446

205

227

281

310

226

52

223

467

21

124

172

480

65

187

53

28

85

254

313

161

495

273

113

284

11

287

307

19

338

433

66

501

262

300

155

133

97

218

203

220

326

500

20

311

414

333

491

Laboratory code

Figure 45. Reported results and evaluation parameters for Pb-212. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 45. Evaluation results for Pb-212 in Sample 3, Rep. Value and Rep. Unc. in Bq/kg

Labcode

Reported value

[Bq/kg ]

Reported uncertainty

(k = 1) [Bq/kg ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

Activity concentration  [ Bq/kg ]

35

40

2	36.5	2.1	-0.8	6.48	A	A	A

11	42.6	3.87	15.8	9.56	A	A	A

13	38.6	1.5	4.9	4.90	A	A	A

| Reported value [Bq/kg ] 38.9 | Reported uncertainty (k = 1) [Bq/kg ] 2.1 | Relative bias [%] 5.7 | P-Test [%] 6.17 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 44 | 9.2 | 19.6 | 21.12 | A | N | W |
| 51.57 | 1.29 | 40.1 | 3.90 | N | N | N |
| 41 | 10 | 11.4 | 24.57 | A | N | W |
| 33.3 | 2.69 | -9.5 | 8.61 | A | A | A |
| 31.6 | 1.9 | -14.1 | 6.71 | A | A | A |
| 41.25 | 3.627 | 12.1 | 9.29 | A | A | A |
| 38.3 | 1.42 | 4.1 | 4.76 | A | A | A |
| 28.7 | 2.2 | -22.0 | 8.23 | N | N | N |
| 0.18 | 0.02 | -99.5 | 11.51 | N | N | N |
| 34.72 | 2.54 | -5.7 | 7.90 | A | A | A |
| 33.44 | 2.8 | -9.1 | 8.89 | A | A | A |
| 33.47 | 1.49 | -9.1 | 5.36 | A | A | A |
| 40 | 5 | 8.7 | 12.85 | A | A | A |
| 34.2 | 1.3 | -7.1 | 4.84 | A | A | A |
| 40.5 | 1.72 | 10.1 | 5.19 | A | A | A |
| 41.23 | 3.69 | 12.0 | 9.44 | A | A | A |
| 35.59 | 1.99 | -3.3 | 6.34 | A | A | A |
| 35.9 | 1.8 | -2.5 | 5.84 | A | A | A |
| 37.1 | 2.1 | 0.8 | 6.40 | A | A | A |
| 38.7 | 7.2 | 5.2 | 18.84 | A | A | A |
| 37.92 | 0.99 | 3.0 | 3.97 | A | A | A |
| 35.4 | 1.5 | -3.8 | 5.19 | A | A | A |
| 41.2 | 3 | 12.0 | 7.87 | A | A | A |
| 44.97 | 2.4045 | 22.2 | 6.13 | N | N | N |
| 34.4 | 1.98 | -6.5 | 6.49 | A | A | A |
| 38.74 | 1.94 | 5.3 | 5.83 | A | A | A |
| 34.4 | 1.91 | -6.5 | 6.31 | A | A | A |
| 35.83 | 1.53 | -2.6 | 5.21 | A | A | A |
| 36.2 | 3 | -1.6 | 8.81 | A | A | A |
| 38.48 | 4.94 | 4.6 | 13.18 | A | A | A |
| 37.9 | 5.5 | 3.0 | 14.82 | A | A | A |
| 16.96 | 4.2 | -53.9 | 24.94 | N | N | N |
| 41.533 | 1.787 | 12.9 | 5.24 | A | A | A |
| 37.2 | 3.1 | 1.1 | 8.85 | A | A | A |
| 11.07 | 1.24 | -69.9 | 11.59 | N | N | N |
| 46.5 | 5.1 | 26.4 | 11.37 | N | A | N |
| 35.5 | 3.6 | -3.5 | 10.57 | A | A | A |
| 36.4 | 1.7 | -1.1 | 5.54 | A | A | A |
| 34.8 | 1.4 | -5.4 | 5.01 | A | A | A |
| 38 | 3.1 | 3.3 | 8.69 | A | A | A |
| 39.1 | 0.4 | 6.2 | 3.16 | A | A | A |
| 37.1 | 2.2 | 0.8 | 6.64 | A | A | A |
| 35.67 | 2.29 | -3.1 | 7.08 | A | A | A |
| 42.55 | 10.64 | 15.6 | 25.18 | A | N | W |
| 36.4 | 3.6 | -1.1 | 10.33 | A | A | A |
| 37.14 | 2.4 | 0.9 | 7.12 | A | A | A |
| 39.17 | 3.85 | 6.4 | 10.27 | A | A | A |
| 35 | 2.2 | -4.9 | 6.96 | A | A | A |
| 41 | 3 | 11.4 | 7.90 | A | A | A |
| 38.56 | 2.04 | 4.8 | 6.08 | A | A | A |
| 32.59043 | 1.081342 | -11.4 | 4.47 | A | A | A |
| 34.9 | 4.3 | -5.2 | 12.68 | A | A | A |
| 46 | 4 | 25.0 | 9.20 | N | N | N |
| 34.3 | 1.8 | -6.8 | 6.04 | A | A | A |
| 28.2 | 4.51 | -23.4 | 16.27 | N | A | N |
| 37.1 | 2.6 | 0.8 | 7.62 | A | A | A |
| 32.17 | 2.25 | -12.6 | 7.61 | A | A | A |
| 36.1 | 2.1 | -1.9 | 6.54 | A | A | A |
| 45.9 | 3 | 24.7 | 7.19 | N | N | N |

Labcode

15

19

20

21

24

25

28

32

34

37

40

41

45

49

50

52

53

55

58

60

62

63

64

65

66

68

71

72

78

79

81

83

84

85

90

94

97

102

104

105

107

109

110

111

113

114

116

119

122

124

128

131

132

133

140

145

150

151

154

155

| Reported value [Bq/kg ] 36.86 | Reported uncertainty (k = 1) [Bq/kg ] 0.8 | Relative bias [%] 0.2 | P-Test [%] 3.69 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 36 | 0.936 | -2.2 | 3.96 | A | A | A |
| 41.9 | 2.6 | 13.9 | 6.89 | A | A | A |
| 29.5 | 4.3 | -19.8 | 14.88 | A | A | A |
| 33 | 3.9 | -10.3 | 12.19 | A | A | A |
| 38.1 | 1.9 | 3.5 | 5.81 | A | A | A |
| 41 | 4 | 11.4 | 10.20 | A | A | A |
| 38.1 | 3.6 | 3.5 | 9.91 | A | A | A |
| 37.8 | 5.7 | 2.7 | 15.37 | A | A | A |
| 29.62 | 0.46 | -19.5 | 3.37 | A | N | W |
| 41.2 | 3 | 12.0 | 7.87 | A | A | A |
| 39.5 | 2.6 | 7.3 | 7.23 | A | A | A |
| 36.42 | 5.1 | -1.0 | 14.32 | A | A | A |
| 38.8 | 3.4 | 5.4 | 9.26 | A | A | A |
| 37.04 | 0.61 | 0.7 | 3.41 | A | A | A |
| 48.7 | 4.6 | 32.3 | 9.91 | N | N | N |
| 37 | 2 | 0.5 | 6.18 | A | A | A |
| 40.2 | 2.24 | 9.2 | 6.32 | A | A | A |
| 38.94 | 1.6 | 5.8 | 5.08 | A | A | A |
| 36.7 | 3.05 | -0.3 | 8.83 | A | A | A |
| 37.4 | 4.3 | 1.6 | 11.88 | A | A | A |
| 36.5 | 5 | -0.8 | 14.02 | A | A | A |
| 46.52533 | 2.29717 | 26.4 | 5.77 | N | N | N |
| 48.8 | 4.5 | 32.6 | 9.69 | N | N | N |
| 40.6 | 4.1 | 10.3 | 10.53 | A | A | A |
| 27.8 | 2.5 | -24.5 | 9.48 | N | N | N |
| 40.39 | 3.27 | 9.8 | 8.63 | A | A | A |
| 40.2 | 1.3 | 9.2 | 4.40 | A | A | A |
| 33.1 | 2.1 | -10.1 | 7.01 | A | A | A |
| 34.5 | 2.2 | -6.2 | 7.04 | A | A | A |
| 36.4 | 3.5 | -1.1 | 10.07 | A | A | A |
| 36.8 | 1.02 | 0.0 | 4.08 | A | A | A |
| 37.518 | 4.414 | 2.0 | 12.14 | A | A | A |
| 34.57 | 0.743388 | -6.1 | 3.68 | A | A | A |
| 40 | 2.8 | 8.7 | 7.61 | A | A | A |
| 41.69 | 2.68 | 13.3 | 7.09 | A | A | A |
| 36.02 | 1.99 | -2.1 | 6.28 | A | A | A |
| 36.9 | 1.7 | 0.3 | 5.49 | A | A | A |
| 36.2 | 2.5 | -1.6 | 7.53 | A | A | A |
| 39.5 | 1.5 | 7.3 | 4.83 | A | A | A |
| 45.3 | 1.1 | 23.1 | 3.85 | N | N | N |
| 39.61 | 3.76 | 7.6 | 9.95 | A | A | A |
| 32.5 | 1.63 | -11.7 | 5.84 | A | A | A |
| 35.5 | 1.3 | -3.5 | 4.73 | A | A | A |
| 35.1 | 2.2 | -4.6 | 6.94 | A | A | A |
| 42 | 3 | 14.1 | 7.74 | A | A | A |
| 35.79 | 2.36 | -2.7 | 7.24 | A | A | A |
| 35.5 | 1.2 | -3.5 | 4.51 | A | A | A |
| 40.3 | 1.5 | 9.5 | 4.77 | A | A | A |
| 19.7 | 0.5 | -46.5 | 3.92 | N | N | N |
| 42.59 | 2.45 | 15.7 | 6.48 | A | A | A |
| 38.4 | 4.4 | 4.3 | 11.84 | A | A | A |
| 43.1 | 2.1 | 17.1 | 5.72 | A | N | W |
| 36.2 | 2.56 | -1.6 | 7.68 | A | A | A |
| 35 | 1.3 | -4.9 | 4.77 | A | A | A |
| 39.797 | 4.975 | 8.1 | 12.85 | A | A | A |
| 31.1 | 2.6 | -15.5 | 8.88 | A | A | A |
| 39 | 3 | 6.0 | 8.25 | A | A | A |
| 45.7 | 0.47 | 24.2 | 3.16 | N | N | N |
| 31 | 4 | -15.8 | 13.24 | A | A | A |

Labcode

156

159

161

166

168

170

172

178

184

185

187

191

197

198

202

203

204

205

211

213

215

217

218

220

223

224

226

227

232

233

235

242

247

249

253

254

256

258

260

261

262

267

268

270

272

273

276

278

281

283

284

286

287

288

290

291

293

299

300

303

| Reported value [Bq/kg ] 39.9 | Reported uncertainty (k = 1) [Bq/kg ] 1.5 | Relative bias [%] 8.4 | P-Test [%] 4.80 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 43.6 | 3.5 | 18.5 | 8.57 | A | A | A |
| 40.3 | 2.9 | 9.5 | 7.79 | A | A | A |
| 52.2 | 4.03 | 41.9 | 8.28 | N | N | N |
| 36.2 | 3.8 | -1.6 | 10.91 | A | A | A |
| 41.7 | 3.46 | 13.3 | 8.82 | A | A | A |
| 28.7 | 0.46 | -22.0 | 3.39 | N | N | N |
| 38.325 | 1.873 | 4.1 | 5.73 | A | A | A |
| 36 | 6 | -2.2 | 16.93 | A | A | A |
| 39.4 | 0.94 | 7.1 | 3.82 | A | A | A |
| 33.1 | 1.4 | -10.1 | 5.18 | A | A | A |
| 39 | 3.32 | 6.0 | 9.02 | A | A | A |
| 51.04 | 3.8 | 38.7 | 8.02 | N | N | N |
| 35.5 | 4 | -3.5 | 11.66 | A | A | A |
| 30.06 | 1.6 | -18.3 | 6.10 | A | N | W |
| 59.52 | 1.51 | 61.7 | 3.92 | N | N | N |
| 44.32 | 2.71 | 20.4 | 6.81 | N | N | N |
| 37.9 | 1.2 | 3.0 | 4.35 | A | A | A |
| 37.3 | 1.2 | 1.4 | 4.39 | A | A | A |
| 37.1 | 2.8 | 0.8 | 8.12 | A | A | A |
| 38.21 | 3.2 | 3.8 | 8.89 | A | A | A |
| 38.05 | 1.9 | 3.4 | 5.82 | A | A | A |
| 39 | 2 | 6.0 | 5.94 | A | A | A |
| 26.04 | 0.32 | -29.2 | 3.23 | N | N | N |
| 37.8 | 4.7 | 2.7 | 12.79 | A | A | A |
| 36.9 | 2.7 | 0.3 | 7.90 | A | A | A |
| 34.6 | 3.5 | -6.0 | 10.55 | A | A | A |
| 36.9 | 2.1 | 0.3 | 6.43 | A | A | A |
| 30.1 | 0.87 | -18.2 | 4.16 | A | N | W |
| 37.03 | 3.7 | 0.6 | 10.43 | A | A | A |
| 39.3 | 3 | 6.8 | 8.20 | A | A | A |
| 38.9 | 1.6 | 5.7 | 5.08 | A | A | A |
| 37.5 | 4.4 | 1.9 | 12.11 | A | A | A |
| 39.08 | 2.49 | 6.2 | 7.04 | A | A | A |
| 37.46 | 2.61 | 1.8 | 7.58 | A | A | A |
| 32.7 | 3.6 | -11.1 | 11.41 | A | A | A |
| 33.88 | 1.03 | -7.9 | 4.26 | A | A | A |
| 38.3 | 3 | 4.1 | 8.38 | A | A | A |
| 38.3 | 2.1 | 4.1 | 6.24 | A | A | A |
| 35.45 | 4.18 | -3.7 | 12.16 | A | A | A |
| 36.7 | 4.2 | -0.3 | 11.83 | A | A | A |
| 22 | 2 | -40.2 | 9.57 | N | N | N |
| 39.4 | 2.2 | 7.1 | 6.33 | A | A | A |
| 36.96 | 2.25 | 0.4 | 6.78 | A | A | A |
| 34.5 | 1.1 | -6.2 | 4.37 | A | A | A |
| 36.2 | 2.3 | -1.6 | 7.02 | A | A | A |
| 36.4 | 1.7 | -1.1 | 5.54 | A | A | A |
| 29.08 | 2.35 | -21.0 | 8.62 | N | A | N |
| 53.9 | 4.3 | 46.5 | 8.52 | N | N | N |
| 28.5 | 1.8 | -22.6 | 6.99 | N | N | N |
| 37.9 | 2.8 | 3.0 | 7.97 | A | A | A |
| 37.7 | 2.6 | 2.5 | 7.52 | A | A | A |
| 38.6 | 4.16 | 4.9 | 11.18 | A | A | A |
| 38.7 | 4.1 | 5.2 | 11.01 | A | A | A |
| 33.5 | 0.63 | -9.0 | 3.53 | A | A | A |
| 37.8 | 2 | 2.7 | 6.08 | A | A | A |
| 44.67 | 1.26 | 21.4 | 4.11 | N | N | N |
| 32.2 | 5.3 | -12.5 | 16.73 | A | A | A |
| 35.4 | 1.9 | -3.8 | 6.14 | A | A | A |
| 33.5 | 2.27 | -9.0 | 7.41 | A | A | A |

Labcode

304

307

310

311

312

313

314

315

316

318

319

325

326

328

330

333

338

339

341

344

347

347

348

355

356

357

358

364

369

371

373

376

379

381

382

384

391

393

393

395

398

400

402

403

408

410

411

412

414

417

421

425

426

428

429

430

433

437

441

444

| Reported value [Bq/kg ] 36.6 | Reported uncertainty (k = 1) [Bq/kg ] 2.1 | Relative bias [%] -0.5 | P-Test [%] 6.47 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 40 | 3.7 | 8.7 | 9.72 | A | A | A |
| 29.62 | 0.46 | -19.5 | 3.37 | A | N | W |
| 37.4 | 1.1 | 1.6 | 4.19 | A | A | A |
| 34.76 | 2.0856 | -5.5 | 6.70 | A | A | A |
| 30.5 | 1.39 | -17.1 | 5.45 | A | N | W |
| 34.61 | 2.96 | -6.0 | 9.06 | A | A | A |
| 39 | 2 | 6.0 | 5.94 | A | A | A |
| 37 | 2 | 0.5 | 6.18 | A | A | A |
| 20.31 | 4.71 | -44.8 | 23.38 | N | N | N |
| 36.2 | 3.6 | -1.6 | 10.38 | A | A | A |
| 40.99 | 7.13 | 11.4 | 17.65 | A | A | A |
| 35.7 | 1.9 | -3.0 | 6.10 | A | A | A |
| 34.98 | 2.29 | -5.0 | 7.20 | A | A | A |
| 41 | 2.9 | 11.4 | 7.68 | A | A | A |
| 35.4 | 1.9 | -3.8 | 6.14 | A | A | A |
| 33.5 | 1.6 | -9.0 | 5.63 | A | A | A |
| 35 | 1.4 | -4.9 | 4.99 | A | A | A |
| 115.977 | 4.643 | 215.2 | 5.00 | N | N | N |
| 41.9 | 0.96 | 13.9 | 3.77 | A | N | W |
| 23.4 | 0.72 | -36.4 | 4.29 | N | N | N |
| 31 | 3 | -15.8 | 10.13 | A | A | A |
| 51.3 | 4.224 | 39.4 | 8.76 | N | N | N |
| 45 | 2.4 | 22.3 | 6.11 | N | N | N |
| 38 | 3 | 3.3 | 8.44 | A | A | A |
| 36 | 1.7 | -2.2 | 5.59 | A | A | A |
| 38.4 | 2 | 4.3 | 6.01 | A | A | A |
| 31 | 1 | -15.8 | 4.40 | A | N | W |
| 37 | 1.9 | 0.5 | 5.94 | A | A | A |
| 30.8 | 0.87 | -16.3 | 4.11 | A | N | W |
| 39.51 | 2.59 | 7.4 | 7.20 | A | A | A |
| 37.3 | 1.9 | 1.4 | 5.91 | A | A | A |
| 37.3 | 2.6 | 1.4 | 7.58 | A | A | A |

Labcode

445

446

448

450

452

453

459

460

462

464

465

467

478

479

480

483

484

487

491

495

496

499

500

501

506

508

516

522

527

530

532

537

555

Pb-214 determination of activity concentration in Sample 3

Table 46. Target values

| Parameter | Value |
| --- | --- |
| Activity concentration [ Bq/kg ] | 47.9 |
| Uncertainty [ Bq/kg ] | 1.3 |
| Maximum Acceptable Relative Bias [%] | 20 |

Accepted 27%	Not Accepted 21%	Warning 6%	Not Reported 46%

60

Figure 46. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

30

40

37

280

84

284

400

94

21

412

283

145

185

448

496

355

102

150

294

300

131

303

522

81

434

40

58

293

417

204

358

79

32

64

369

430

453

34

224

156

260

83

139

24

286

54

316

330

425

175

19

233

288

314

140

281

49

124

499

506

134

319

341

268

291

41

258

364

437

62

408

445

455

55

69

339

151

63

13

452

116

462

555

17

223

381

379

68

376

429

484

165

200

344

97

128

530

370

384

371

197

465

459

508

107

119

272

332

310

464

460

178

271

304

532

217

235

348

276

426

100

215

410

267

122

398

421

537

53

65

187

444

313

478

315

261

388

402

467

20

318

90

184

328

387

407

278

441

483

78

247

154

242

433

85

50

357

198

338

105

403

71

66

104

133

172

480

25

395

287

249

28

52

72

205

213

253

311

382

211

109

446

111

155

15

373

113

450

170

11

159

428

307

179

179

527

60

191

347

2

273

393

516

218

226

326

254

414

333

262

203

463

467

317

491

Laboratory code

Figure 47. Reported results and evaluation parameters for Pb-214. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 47. Evaluation results for Pb-214 in Sample 3, Rep. Value and Rep. Unc. in Bq/kg

Labcode

Reported value

[Bq/kg ]

Reported uncertainty

(k = 1) [Bq/kg ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

Activity concentration  [ Bq/kg ]

50

2	51	3.9	6.5	8.11	A	A	A

11	49	3.84	2.3	8.29	A	A	A

13	38.8	1.5	-19.0	4.72	A	N	W

| Reported value [Bq/kg ] 48.3 | Reported uncertainty (k = 1) [Bq/kg ] 2.8 | Relative bias [%] 0.8 | P-Test [%] 6.40 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 39.06 | 2.33 | -18.5 | 6.55 | A | N | W |
| 36.3 | 7.4 | -24.2 | 20.57 | N | N | N |
| 42.43 | 1.22 | -11.4 | 3.95 | A | N | W |
| 26 | 7 | -45.7 | 27.06 | N | N | N |
| 35.8 | 2.66 | -25.3 | 7.91 | N | N | N |
| 45.1 | 1.9 | -5.8 | 5.01 | A | A | A |
| 45.57 | 4.013 | -4.9 | 9.21 | A | A | A |
| 34.31 | 1.78 | -28.4 | 5.85 | N | N | N |
| 35.3 | 2.3 | -26.3 | 7.06 | N | N | N |
| 0.19 | 0.02 | -99.6 | 10.87 | N | N | N |
| 33.17 | 2.62 | -30.8 | 8.35 | N | N | N |
| 37.6 | 3.09 | -21.5 | 8.65 | N | A | N |
| 37 | 5 | -22.8 | 13.78 | N | A | N |
| 44.3 | 1.4 | -7.5 | 4.17 | A | A | A |
| 45.6 | 2.28 | -4.8 | 5.69 | A | A | A |
| 41.93 | 3.72 | -12.5 | 9.28 | A | A | A |
| 35.9 | 1.8 | -25.0 | 5.70 | N | N | N |
| 38.06 | 1.89 | -20.5 | 5.66 | N | N | N |
| 33.2 | 1.4 | -30.7 | 5.01 | N | N | N |
| 50.8 | 3.4 | 6.0 | 7.22 | A | A | A |
| 37.9 | 7.5 | -20.9 | 19.97 | N | A | N |
| 38.25 | 0.89 | -20.1 | 3.57 | N | N | N |
| 34.5 | 1.7 | -28.0 | 5.63 | N | N | N |
| 42 | 2.8 | -12.3 | 7.20 | A | A | A |
| 44.54 | 2.4445 | -7.0 | 6.12 | A | A | A |
| 39.3 | 1.48 | -18.0 | 4.64 | A | N | W |
| 38.1 | 1.7 | -20.5 | 5.22 | N | N | N |
| 44.52 | 2.23 | -7.1 | 5.70 | A | A | A |
| 45.6 | 1.74 | -4.8 | 4.68 | A | A | A |
| 43.01 | 3.2 | -10.2 | 7.92 | A | A | A |
| 34.3 | 1 | -28.4 | 3.98 | N | N | N |
| 32.28 | 4.06 | -32.6 | 12.87 | N | A | N |
| 35.7 | 5.2 | -25.5 | 14.82 | N | A | N |
| 15.5 | 2.68 | -67.6 | 17.50 | N | N | N |
| 43.726 | 3.11 | -8.7 | 7.61 | A | A | A |
| 42.5 | 2 | -11.3 | 5.43 | A | A | A |
| 24.35 | 1.28 | -49.2 | 5.92 | N | N | N |
| 39.6 | 4.3 | -17.3 | 11.19 | A | A | A |
| 41.3 | 1.1 | -13.8 | 3.80 | A | N | W |
| 30.4 | 3.2 | -36.5 | 10.87 | N | N | N |
| 44.6 | 1.5 | -6.9 | 4.32 | A | A | A |
| 44.4 | 1.5 | -7.3 | 4.33 | A | A | A |
| 40 | 2.8 | -16.5 | 7.51 | A | A | A |
| 47.1 | 0.5 | -1.7 | 2.91 | A | A | A |
| 47.94 | 3.89 | 0.1 | 8.56 | A | A | A |
| 48.44 | 12.11 | 1.1 | 25.15 | A | N | W |
| 38.96 | 1.75 | -18.7 | 5.25 | A | N | W |
| 40.1 | 4 | -16.3 | 10.34 | A | A | A |
| 41.65 | 3.9 | -13.1 | 9.75 | A | A | A |
| 37 | 2 | -22.8 | 6.05 | N | N | N |
| 39.61 | 2.87 | -17.3 | 7.74 | A | A | A |
| 31.99369 | 0.637759 | -33.2 | 3.37 | N | N | N |
| 45 | 3 | -6.0 | 7.20 | A | A | A |
| 37.1 | 0.5 | -22.6 | 3.03 | N | N | N |
| 35.7 | 1.59 | -25.5 | 5.22 | N | N | N |
| 36.8 | 1.9 | -23.2 | 5.83 | N | N | N |
| 29 | 4.64 | -39.5 | 16.23 | N | A | N |
| 31 | 1.9 | -35.3 | 6.70 | N | N | N |
| 38.11 | 3.81 | -20.4 | 10.36 | N | A | N |

Labcode

15

17

19

20

21

24

25

28

32

34

37

40

41

49

50

52

53

54

55

58

60

62

63

64

65

66

68

69

71

72

78

79

81

83

84

85

90

94

97

100

102

104

105

107

109

111

113

116

119

122

124

128

131

133

134

139

140

145

150

151

| Reported value [Bq/kg ] 43.5 | Reported uncertainty (k = 1) [Bq/kg ] 2.7 | Relative bias [%] -9.2 | P-Test [%] 6.77 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 48.1 | 4 | 0.4 | 8.75 | A | A | A |
| 35.5 | 1 | -25.9 | 3.91 | N | N | N |
| 49 | 2.55 | 2.3 | 5.87 | A | A | A |
| 39.5 | 2.4 | -17.5 | 6.65 | A | N | W |
| 48.8 | 1.8 | 1.9 | 4.58 | A | A | A |
| 45 | 4 | -6.0 | 9.29 | A | A | A |
| 36.2 | 1.5 | -24.4 | 4.95 | N | N | N |
| 40.8 | 3.3 | -14.8 | 8.53 | A | A | A |
| 50.13 | 0.56 | 4.7 | 2.93 | A | A | A |
| 50.13 | 0.59 | 4.7 | 2.96 | A | A | A |
| 42.5 | 6.4 | -11.3 | 15.30 | A | A | A |
| 29.05 | 0.94 | -39.4 | 4.22 | N | N | N |
| 42 | 2.8 | -12.3 | 7.20 | A | A | A |
| 50.8 | 4.2 | 6.0 | 8.70 | A | A | A |
| 39.75 | 6.76 | -17.0 | 17.22 | A | A | A |
| 44.3 | 3.7 | -7.5 | 8.78 | A | A | A |
| 39.5 | 2 | -17.5 | 5.74 | A | N | W |
| 58.8 | 5.1 | 22.8 | 9.09 | N | A | N |
| 34 | 1 | -29.0 | 4.00 | N | N | N |
| 46.1 | 2.58 | -3.8 | 6.22 | A | A | A |
| 46.93 | 1.64 | -2.0 | 4.42 | A | A | A |
| 46.2 | 4.28 | -3.5 | 9.65 | A | A | A |
| 41.35 | 4.1 | -13.7 | 10.28 | A | A | A |
| 41 | 6 | -14.4 | 14.88 | A | A | A |
| 52.33706 | 2.47697 | 9.3 | 5.46 | A | A | A |
| 39.1 | 3.5 | -18.4 | 9.35 | A | A | A |
| 35.3 | 3.2 | -26.3 | 9.46 | N | N | N |
| 52.99 | 4.25 | 10.6 | 8.47 | A | A | A |
| 36.6 | 2.3 | -23.6 | 6.85 | N | N | N |
| 41 | 4 | -14.4 | 10.13 | A | A | A |
| 43.56 | 1.2 | -9.1 | 3.87 | A | A | A |
| 43.117 | 4.298 | -10.0 | 10.33 | A | A | A |
| 45.54 | 1.240693 | -4.9 | 3.85 | A | A | A |
| 46.2 | 3.2 | -3.5 | 7.44 | A | A | A |
| 53.32 | 2.99 | 11.3 | 6.23 | A | A | A |
| 37.7 | 1.3 | -21.3 | 4.39 | N | N | N |
| 35.6 | 1.9 | -25.7 | 5.99 | N | N | N |
| 42.3 | 1.6 | -11.7 | 4.66 | A | A | A |
| 58.4 | 1.1 | 21.9 | 3.30 | N | N | N |
| 41.51 | 4.07 | -13.3 | 10.17 | A | A | A |
| 37.4 | 1.87 | -21.9 | 5.69 | N | N | N |
| 40.8 | 3.3 | -14.8 | 8.53 | A | A | A |
| 40.1 | 2.3 | -16.3 | 6.35 | A | A | A |
| 51 | 3 | 6.5 | 6.48 | A | A | A |
| 41.04 | 2.67 | -14.3 | 7.05 | A | A | A |
| 42.9 | 1.4 | -10.4 | 4.24 | A | A | A |
| 11.93 | 2.14 | -75.1 | 18.14 | N | N | N |
| 36.9 | 1.3 | -23.0 | 4.45 | N | N | N |
| 28.3 | 1.12 | -40.9 | 4.80 | N | N | N |
| 22.92 | 1.1 | -52.1 | 5.51 | N | N | N |
| 35.8 | 4.7 | -25.3 | 13.41 | N | A | N |
| 45.5 | 0.9 | -5.0 | 3.36 | A | A | A |
| 36.6 | 1.49 | -23.6 | 4.89 | N | N | N |
| 37.536 | 4.692 | -21.6 | 12.79 | N | A | N |
| 33.7 | 2.5 | -29.6 | 7.90 | N | N | N |
| 31 | 2.6 | -35.3 | 8.82 | N | N | N |
| 31.06 | 0.48 | -35.2 | 3.12 | N | N | N |
| 32 | 4 | -33.2 | 12.79 | N | N | N |
| 40.8 | 1.4 | -14.8 | 4.37 | A | N | W |

Labcode

154

155

156

159

165

170

172

175

178

179

179

184

185

187

191

197

198

200

203

204

205

211

213

215

217

218

223

224

226

233

235

242

247

249

253

254

258

260

261

262

267

268

271

272

273

276

278

280

281

283

284

286

287

288

291

293

294

300

303

304

| Reported value [Bq/kg ] 49.9 | Reported uncertainty (k = 1) [Bq/kg ] 2.9 | Relative bias [%] 4.2 | P-Test [%] 6.41 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 40.4 | 2.9 | -15.7 | 7.67 | A | A | A |
| 46.3 | 5.9 | -3.3 | 13.03 | A | A | A |
| 42.1 | 2.84 | -12.1 | 7.27 | A | A | A |
| 36.6 | 1.47 | -23.6 | 4.85 | N | N | N |
| 42.2 | 2.04 | -11.9 | 5.54 | A | A | A |
| 36.1 | 5.3 | -24.6 | 14.93 | N | A | N |
| 64.6 | 4.5 | 34.9 | 7.48 | N | N | N |
| 42.46 | 1.07 | -11.4 | 3.70 | A | N | W |
| 37.3 | 1.2 | -22.1 | 4.21 | N | N | N |
| 53.25 | 3.1 | 11.2 | 6.42 | A | A | A |
| 42.5 | 3 | -11.3 | 7.56 | A | A | A |
| 36.1 | 1.63 | -24.6 | 5.27 | N | N | N |
| 40.2 | 3.2 | -16.1 | 8.41 | A | A | A |
| 57.88 | 2.52 | 20.8 | 5.13 | N | N | N |
| 44.32 | 2.71 | -7.5 | 6.69 | A | A | A |
| 38.1 | 1.1 | -20.5 | 3.96 | N | N | N |
| 37.3 | 1 | -22.1 | 3.81 | N | N | N |
| 39.5 | 3.1 | -17.5 | 8.30 | A | A | A |
| 50.83 | 2.2 | 6.1 | 5.11 | A | A | A |
| 41 | 2 | -14.4 | 5.58 | A | N | W |
| 29.46 | 0.34 | -38.5 | 2.95 | N | N | N |
| 44.3 | 2.3 | -7.5 | 5.86 | A | A | A |
| 34.1 | 3.4 | -28.8 | 10.33 | N | N | N |
| 37.7 | 2.2 | -21.3 | 6.44 | N | N | N |
| 34.7 | 0.78 | -27.6 | 3.52 | N | N | N |
| 39.7 | 2.8 | -17.1 | 7.56 | A | A | A |
| 39.71 | 4.16 | -17.1 | 10.82 | A | A | A |
| 48.3 | 2.1 | 0.8 | 5.13 | A | A | A |
| 39.4 | 1.6 | -17.8 | 4.88 | A | N | W |
| 39.2 | 6.2 | -18.2 | 16.05 | A | A | A |
| 39.17 | 2.37 | -18.2 | 6.63 | A | N | W |
| 46.52 | 3.25 | -2.9 | 7.49 | A | A | A |
| 39.7 | 4.1 | -17.1 | 10.68 | A | A | A |
| 42.7 | 1.2 | -10.9 | 3.91 | A | N | W |
| 42.3 | 1.2 | -11.7 | 3.93 | A | N | W |
| 51 | 2.1 | 6.5 | 4.93 | A | A | A |
| 45.3 | 5.25 | -5.4 | 11.90 | A | A | A |
| 41.7 | 4.1 | -12.9 | 10.20 | A | A | A |
| 23 | 4 | -52.0 | 17.60 | N | N | N |
| 42.3 | 2.5 | -11.7 | 6.50 | A | A | A |
| 44.41 | 3.6 | -7.3 | 8.55 | A | A | A |
| 42.7 | 2.1 | -10.9 | 5.62 | A | A | A |
| 38 | 4 | -20.7 | 10.87 | N | A | N |
| 41.5 | 2.4 | -13.4 | 6.39 | A | A | A |
| 26.55 | 2.72 | -44.6 | 10.60 | N | N | N |
| 57.4 | 4.6 | 19.8 | 8.46 | A | A | A |
| 33.8 | 2.1 | -29.4 | 6.78 | N | N | N |
| 41.9 | 3.1 | -12.5 | 7.88 | A | A | A |
| 36.1 | 2.1 | -24.6 | 6.42 | N | N | N |
| 41.26 | 4.06 | -13.9 | 10.21 | A | A | A |
| 49.65 | 5 | 3.6 | 10.43 | A | A | A |
| 39.4 | 0.805 | -17.8 | 3.40 | A | N | W |
| 34.9 | 1.8 | -27.1 | 5.83 | N | N | N |
| 43.61 | 0.75 | -9.0 | 3.21 | A | N | W |
| 33.05 | 1.9 | -31.0 | 6.36 | N | N | N |
| 37.7 | 5.5 | -21.3 | 14.84 | N | A | N |
| 43 | 2.5 | -10.2 | 6.42 | A | A | A |
| 42.06 | 2.42 | -12.2 | 6.36 | A | A | A |
| 38 | 2.6 | -20.7 | 7.36 | N | N | N |

Labcode

307

310

311

313

314

315

316

317

318

319

326

328

330

332

333

338

339

341

344

347

348

355

357

358

364

369

370

371

373

376

379

381

382

384

387

388

393

395

398

400

402

403

407

408

410

412

414

417

421

425

426

428

429

430

433

434

437

441

444

445

| Reported value [Bq/kg ] 47.7 | Reported uncertainty (k = 1) [Bq/kg ] 3.3 | Relative bias [%] -0.4 | P-Test [%] 7.43 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 29.05 | 0.94 | -39.4 | 4.22 | N | N | N |
| 48.7 | 1.5 | 1.7 | 4.11 | A | A | A |
| 38.81 | 2.40622 | -19.0 | 6.77 | A | N | W |
| 34.9 | 1.11 | -27.1 | 4.18 | N | N | N |
| 38 | 2.4 | -20.7 | 6.87 | N | N | N |
| 39.89 | 3.99 | -16.7 | 10.36 | A | A | A |
| 40.7 | 3 | -15.0 | 7.85 | A | A | A |
| 39 | 2 | -18.6 | 5.80 | A | N | W |
| 60 | 3 | 25.3 | 5.69 | N | N | N |
| 40.46 | 2.38 | -15.5 | 6.48 | A | A | A |
| 39.8 | 4 | -16.9 | 10.41 | A | A | A |
| 61.48 | 13.7 | 28.4 | 22.45 | N | N | N |
| 42.35 | 8.4 | -11.6 | 20.02 | A | N | W |
| 42.1 | 2.9 | -12.1 | 7.40 | A | A | A |
| 45 | 3.8 | -6.0 | 8.87 | A | A | A |
| 43 | 2.5 | -10.2 | 6.42 | A | A | A |
| 39.4 | 1.5 | -17.8 | 4.68 | A | N | W |
| 102.118 | 2.395 | 113.2 | 3.59 | N | N | N |
| 29.23 | 1.01 | -39.0 | 4.39 | N | N | N |
| 37 | 3 | -22.8 | 8.55 | N | N | N |
| 37 | 3 | -22.8 | 8.55 | N | N | N |
| 39.99 | 0.92 | -16.5 | 3.56 | A | N | W |
| 51.1 | 3 | 6.7 | 6.47 | A | A | A |
| 32 | 1.1 | -33.2 | 4.38 | N | N | N |
| 50.2 | 3 | 4.8 | 6.56 | A | A | A |
| 39.61 | 1.11 | -17.3 | 3.90 | A | N | W |
| 40.97 | 2.73 | -14.5 | 7.19 | A | A | A |
| 41.9 | 1.7 | -12.5 | 4.88 | A | A | A |
| 39 | 2.7 | -18.6 | 7.44 | A | A | A |

Labcode

446

448

450

452

453

455

459

460

462

463

464

465

467

467

478

480

483

484

491

496

499

506

508

516

522

527

530

532

537

555

Bi-214 determination of activity concentration in Sample 3

Table 48. Target values

| Parameter | Value |
| --- | --- |
| Activity concentration [ Bq/kg ] | 47.9 |
| Uncertainty [ Bq/kg ] | 1.3 |
| Maximum Acceptable Relative Bias [%] | 20 |

Accepted 25%	Not Accepted 22%	Warning 8%	Not Reported 45%

60

Figure 48. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

30

40

37

400

19

412

284

94

185

448

355

294

145

131

102

495

21

283

58

81

300

496

522

434

303

364

34

453

156

430

40

358

45

293

310

32

134

386

140

369

384

79

462

506

224

281

150

260

139

467

464

64

204

233

437

316

330

13

83

71

408

371

429

555

280

62

288

55

63

319

341

268

258

286

402

41

417

452

9

49

271

445

499

339

460

291

97

8

54

425

116

197

426

223

410

381

376

17

165

200

344

68

205

128

370

379

444

508

107

124

172

119

459

332

272

287

276

178

151

532

65

187

217

235

313

348

530

90

421

100

315

267

28

249

278

537

247

318

208

478

484

122

261

388

446

482

215

184

292

387

53

398

214

441

483

154

265

407

242

433

85

357

67

328

480

66

304

50

198

338

105

403

155

104

84

133

211

11

395

52

72

60

218

213

311

25

382

109

111

191

450

15

113

2

516

159

170

307

273

428

179

78

253

347

393

373

226

527

57

326

203

414

254

333

262

463

317

491

Laboratory code

Figure 49. Reported results and evaluation parameters for Bi-214. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 49. Evaluation results for Bi-214 in Sample 3, Rep. Value and Rep. Unc. in Bq/kg

Labcode

Reported value

[Bq/kg ]

Reported uncertainty

(k = 1) [Bq/kg ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

Activity concentration  [ Bq/kg ]

50

2	48.5	5.5	1.2	11.66	A	A	A

8	38.6	2.85	-19.4	7.87	A	A	A

9	37.8	4.5	-21.1	12.21	N	A	N

| Reported value [Bq/kg ] 45.2 | Reported uncertainty (k = 1) [Bq/kg ] 2.8 | Relative bias [%] -5.6 | P-Test [%] 6.76 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 36.2 | 1.5 | -24.4 | 4.95 | N | N | N |
| 48.3 | 2.8 | 0.8 | 6.40 | A | A | A |
| 39.38 | 1.82 | -17.8 | 5.36 | A | N | W |
| 22.9 | 4.5 | -52.2 | 19.84 | N | N | N |
| 30 | 2 | -37.4 | 7.20 | N | N | N |
| 46.5 | 1.2 | -2.9 | 3.75 | A | A | A |
| 41.52 | 3.662 | -13.3 | 9.23 | A | A | A |
| 34.4 | 2.27 | -28.2 | 7.14 | N | N | N |
| 32.5 | 2 | -32.1 | 6.73 | N | N | N |
| 0.2 | 0.03 | -99.6 | 15.24 | N | N | N |
| 33.42 | 2.42 | -30.2 | 7.73 | N | N | N |
| 37.6 | 1.2 | -21.5 | 4.19 | N | N | N |
| 33.61 | 1.48 | -29.8 | 5.17 | N | N | N |
| 38 | 5 | -20.7 | 13.43 | N | A | N |
| 44.3 | 1.4 | -7.5 | 4.17 | A | A | A |
| 45.6 | 2.28 | -4.8 | 5.69 | A | A | A |
| 42.68 | 4.32 | -10.9 | 10.48 | A | A | A |
| 38.7 | 1.9 | -19.2 | 5.61 | A | N | W |
| 36.66 | 1.09 | -23.5 | 4.03 | N | N | N |
| 52 | 5.2 | 8.6 | 10.36 | A | A | A |
| 30.2 | 1.3 | -37.0 | 5.09 | N | N | N |
| 45.8 | 4.9 | -4.4 | 11.04 | A | A | A |
| 36.6 | 7.1 | -23.6 | 19.59 | N | A | N |
| 36.92 | 0.88 | -22.9 | 3.61 | N | N | N |
| 35.9 | 1.5 | -25.0 | 4.98 | N | N | N |
| 41 | 2.8 | -14.4 | 7.35 | A | A | A |
| 44.07 | 2.382 | -8.0 | 6.05 | A | A | A |
| 44 | 4.7 | -8.1 | 11.02 | A | A | A |
| 39.7 | 1.48 | -17.1 | 4.61 | A | N | W |
| 36.3 | 1.81 | -24.2 | 5.68 | N | N | N |
| 45.6 | 1.74 | -4.8 | 4.68 | A | A | A |
| 50.4 | 7.4 | 5.2 | 14.93 | A | A | A |
| 35 | 2.4 | -26.9 | 7.37 | N | N | N |
| 30.38 | 3.42 | -36.6 | 11.58 | N | N | N |
| 36.2 | 5.25 | -24.4 | 14.75 | N | A | N |
| 44.65 | 11.98 | -6.8 | 26.97 | A | N | W |
| 43.726 | 3.11 | -8.7 | 7.61 | A | A | A |
| 41.1 | 1.8 | -14.2 | 5.15 | A | N | W |
| 24.35 | 1.28 | -49.2 | 5.92 | N | N | N |
| 38.57 | 4.2 | -19.5 | 11.22 | A | A | A |
| 41.3 | 1.1 | -13.8 | 3.80 | A | N | W |
| 29.5 | 3 | -38.4 | 10.53 | N | N | N |
| 44.6 | 1.5 | -6.9 | 4.32 | A | A | A |
| 44.4 | 1.5 | -7.3 | 4.33 | A | A | A |
| 40 | 2.8 | -16.5 | 7.51 | A | A | A |
| 46.8 | 0.5 | -2.3 | 2.92 | A | A | A |
| 47 | 3.85 | -1.9 | 8.63 | A | A | A |
| 48.44 | 12.11 | 1.1 | 25.15 | A | N | W |
| 38.81 | 2.72 | -19.0 | 7.52 | A | A | A |
| 40.1 | 4 | -16.3 | 10.34 | A | A | A |
| 42.25 | 4 | -11.8 | 9.85 | A | A | A |
| 40 | 7 | -16.5 | 17.71 | A | A | A |
| 39.76 | 2.03 | -17.0 | 5.78 | A | N | W |
| 29.24424 | 0.615928 | -39.0 | 3.44 | N | N | N |
| 45 | 4 | -6.0 | 9.29 | A | A | A |
| 34.4 | 0.5 | -28.2 | 3.08 | N | N | N |
| 35.7 | 1.59 | -25.5 | 5.22 | N | N | N |
| 34.6 | 1.8 | -27.8 | 5.87 | N | N | N |
| 29.2 | 4.67 | -39.0 | 16.22 | N | A | N |

Labcode

11

13

15

17

19

21

25

28

32

34

37

40

41

45

49

50

52

53

54

55

57

58

60

62

63

64

65

66

67

68

71

72

78

79

81

83

84

85

90

94

97

100

102

104

105

107

109

111

113

116

119

122

124

128

131

133

134

139

140

145

| Reported value [Bq/kg ] 35.4 | Reported uncertainty (k = 1) [Bq/kg ] 1.8 | Relative bias [%] -26.1 | P-Test [%] 5.76 | Trueness evaluation N | Precision evaluation N | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 40.96 | 4.1 | -14.5 | 10.37 | A | A | A |
| 43.5 | 2.7 | -9.2 | 6.77 | A | A | A |
| 44.5 | 3.7 | -7.1 | 8.75 | A | A | A |
| 32.74 | 0.86 | -31.6 | 3.78 | N | N | N |
| 49 | 2.55 | 2.3 | 5.87 | A | A | A |
| 39.5 | 2.4 | -17.5 | 6.65 | A | N | W |
| 49.8 | 1.7 | 4.0 | 4.36 | A | A | A |
| 40 | 4 | -16.5 | 10.36 | A | A | A |
| 40.8 | 3.3 | -14.8 | 8.53 | A | A | A |
| 50.16 | 0.59 | 4.7 | 2.96 | A | A | A |
| 42.5 | 6.4 | -11.3 | 15.30 | A | A | A |
| 27.27 | 0.92 | -43.1 | 4.33 | N | N | N |
| 41 | 2.8 | -14.4 | 7.35 | A | A | A |
| 47.6 | 4.2 | -0.6 | 9.23 | A | A | A |
| 38.93 | 6.62 | -18.7 | 17.22 | A | A | A |
| 44.3 | 3.7 | -7.5 | 8.78 | A | A | A |
| 39.5 | 2 | -17.5 | 5.74 | A | N | W |
| 54.5 | 9 | 13.8 | 16.74 | A | A | A |
| 36 | 2 | -24.8 | 6.18 | N | N | N |
| 39.7 | 2.23 | -17.1 | 6.24 | A | N | W |
| 42.1 | 1.33 | -12.1 | 4.16 | A | N | W |
| 45.12 | 1.4 | -5.8 | 4.12 | A | A | A |
| 46.1 | 3.35 | -3.8 | 7.76 | A | A | A |
| 43 | 3 | -10.2 | 7.49 | A | A | A |
| 42.45 | 4.2 | -11.4 | 10.26 | A | A | A |
| 41 | 6 | -14.4 | 14.88 | A | A | A |
| 45.83091 | 1.26166 | -4.3 | 3.87 | A | A | A |
| 39.1 | 3.5 | -18.4 | 9.35 | A | A | A |
| 35.3 | 3.2 | -26.3 | 9.46 | N | N | N |
| 51.48 | 3.12 | 7.5 | 6.64 | A | A | A |
| 36 | 2.9 | -24.8 | 8.50 | N | N | N |
| 41 | 4 | -14.4 | 10.13 | A | A | A |
| 43.53 | 1.16 | -9.1 | 3.80 | A | A | A |
| 42.029 | 3.556 | -12.3 | 8.89 | A | A | A |
| 41.85 | 0.735266 | -12.6 | 3.23 | A | N | W |
| 50.8 | 3.8 | 6.0 | 7.96 | A | A | A |
| 55.49 | 4.24 | 15.8 | 8.11 | A | A | A |
| 37.5 | 1.2 | -21.7 | 4.20 | N | N | N |
| 35.6 | 1.7 | -25.7 | 5.49 | N | N | N |
| 42.3 | 1.6 | -11.7 | 4.66 | A | A | A |
| 58.4 | 1.1 | 21.9 | 3.30 | N | N | N |
| 43.5 | 5.2 | -9.2 | 12.26 | A | A | A |
| 41.42 | 4.06 | -13.5 | 10.17 | A | A | A |
| 37.4 | 1.87 | -21.9 | 5.69 | N | N | N |
| 38 | 5.6 | -20.7 | 14.98 | N | A | N |
| 40.3 | 2.1 | -15.9 | 5.88 | A | N | W |
| 50 | 3 | 4.4 | 6.59 | A | A | A |
| 40.74 | 2.73 | -14.9 | 7.23 | A | A | A |
| 41.9 | 1.4 | -12.5 | 4.30 | A | N | W |
| 36.57 | 6.41 | -23.6 | 17.74 | N | A | N |
| 35.3 | 1.5 | -26.3 | 5.04 | N | N | N |
| 30 | 2.21 | -37.4 | 7.85 | N | N | N |
| 24.13 | 1.16 | -49.6 | 5.52 | N | N | N |
| 37.5 | 3.9 | -21.7 | 10.75 | N | A | N |
| 40.4 | 2.6 | -15.7 | 6.98 | A | A | A |
| 36.6 | 1.56 | -23.6 | 5.05 | N | N | N |
| 38.44 | 4.956 | -19.8 | 13.18 | A | A | A |
| 42.5 | 1.2 | -11.3 | 3.92 | A | N | W |
| 34.3 | 2.6 | -28.4 | 8.05 | N | N | N |

Labcode

150

151

154

155

156

159

165

170

172

178

179

184

185

187

191

197

198

200

203

204

205

208

211

213

214

215

217

218

223

224

226

233

235

242

247

249

253

254

258

260

261

262

265

267

268

271

272

273

276

278

280

281

283

284

286

287

288

291

292

293

| Reported value [Bq/kg ] 28.8 | Reported uncertainty (k = 1) [Bq/kg ] 2.1 | Relative bias [%] -39.9 | P-Test [%] 7.78 | Trueness evaluation N | Precision evaluation N | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 30.42 | 0.53 | -36.5 | 3.23 | N | N | N |
| 32 | 4 | -33.2 | 12.79 | N | N | N |
| 44.2 | 2.4 | -7.7 | 6.07 | A | A | A |
| 49.9 | 2.9 | 4.2 | 6.41 | A | A | A |
| 34.3 | 2.5 | -28.4 | 7.78 | N | N | N |
| 46.1 | 1.25 | -3.8 | 3.84 | A | A | A |
| 41 | 2.4 | -14.4 | 6.45 | A | A | A |
| 41.3 | 1.98 | -13.8 | 5.51 | A | A | A |
| 36.1 | 5.3 | -24.6 | 14.93 | N | A | N |
| 65.8 | 3.8 | 37.4 | 6.38 | N | N | N |
| 42.09 | 0.97 | -12.1 | 3.56 | A | N | W |
| 37.3 | 1.2 | -22.1 | 4.21 | N | N | N |
| 52.75 | 2.7 | 10.1 | 5.79 | A | A | A |
| 44 | 3.1 | -8.1 | 7.55 | A | A | A |
| 36.1 | 1.63 | -24.6 | 5.27 | N | N | N |
| 40.2 | 3.2 | -16.1 | 8.41 | A | A | A |
| 57.2 | 2.97 | 19.4 | 5.86 | A | N | W |
| 44.32 | 2.71 | -7.5 | 6.69 | A | A | A |
| 38.1 | 1.1 | -20.5 | 3.96 | N | N | N |
| 37.3 | 1 | -22.1 | 3.81 | N | N | N |
| 39.5 | 3.1 | -17.5 | 8.30 | A | A | A |
| 50.87 | 2.55 | 6.2 | 5.70 | A | A | A |
| 41 | 2 | -14.4 | 5.58 | A | N | W |
| 28.36 | 0.55 | -40.8 | 3.34 | N | N | N |
| 43.9 | 2.3 | -8.3 | 5.90 | A | A | A |
| 33.5 | 3.4 | -30.1 | 10.51 | N | N | N |
| 32 | 1.8 | -33.2 | 6.25 | N | N | N |
| 34.8 | 0.93 | -27.4 | 3.81 | N | N | N |
| 39.8 | 3.2 | -16.9 | 8.49 | A | A | A |
| 36.38 | 4.62 | -24.0 | 12.99 | N | A | N |
| 51.2 | 2 | 6.9 | 4.76 | A | A | A |
| 39.3 | 1.6 | -18.0 | 4.89 | A | N | W |
| 39.8 | 6.2 | -16.9 | 15.81 | A | A | A |
| 39.17 | 2.37 | -18.2 | 6.63 | A | N | W |
| 46.52 | 3.25 | -2.9 | 7.49 | A | A | A |
| 34.9 | 3.5 | -27.1 | 10.39 | N | N | N |
| 34.47 | 3.55 | -28.0 | 10.65 | N | N | N |
| 42.6 | 1.2 | -11.1 | 3.91 | A | N | W |
| 42.3 | 1.2 | -11.7 | 3.93 | A | N | W |
| 51.1 | 2.6 | 6.7 | 5.77 | A | A | A |
| 45.3 | 5.25 | -5.4 | 11.90 | A | A | A |
| 42.7 | 4 | -10.9 | 9.75 | A | A | A |
| 20 | 5 | -58.2 | 25.15 | N | N | N |
| 37.5 | 2.2 | -21.7 | 6.46 | N | N | N |
| 44.41 | 3.6 | -7.3 | 8.55 | A | A | A |
| 43.5 | 2.2 | -9.2 | 5.74 | A | A | A |
| 36.3 | 0.8 | -24.2 | 3.50 | N | N | N |
| 39.1 | 2.4 | -18.4 | 6.71 | A | N | W |
| 23.17 | 3.67 | -51.6 | 16.07 | N | N | N |
| 54.5 | 3.9 | 13.8 | 7.65 | A | A | A |
| 37.6 | 2.3 | -21.5 | 6.69 | N | N | N |
| 41.1 | 3 | -14.2 | 7.79 | A | A | A |
| 38.7 | 2.1 | -19.2 | 6.07 | A | N | W |
| 39 | 2.25 | -18.6 | 6.38 | A | N | W |
| 50.05 | 4.9 | 4.5 | 10.16 | A | A | A |
| 36.4 | 0.82 | -24.0 | 3.53 | N | N | N |
| 33.3 | 1.5 | -30.5 | 5.26 | N | N | N |
| 43.61 | 0.81 | -9.0 | 3.29 | A | N | W |
| 31.14 | 2.1 | -35.0 | 7.27 | N | N | N |

Labcode

294

300

303

304

307

310

311

313

315

316

317

318

319

326

328

330

332

333

338

339

341

344

347

348

355

357

358

364

369

370

371

373

376

379

381

382

384

386

387

388

393

395

398

400

402

403

407

408

410

412

414

417

421

425

426

428

429

430

433

434

| Reported value [Bq/kg ] 36 | Reported uncertainty (k = 1) [Bq/kg ] 4.5 | Relative bias [%] -24.8 | P-Test [%] 12.79 | Trueness evaluation N | Precision evaluation A | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 43 | 2.5 | -10.2 | 6.42 | A | A | A |
| 39.94 | 2.14 | -16.6 | 6.01 | A | N | W |
| 38 | 2.6 | -20.7 | 7.36 | N | N | N |
| 42.4 | 1.9 | -11.5 | 5.24 | A | A | A |
| 27.27 | 0.92 | -43.1 | 4.33 | N | N | N |
| 47.6 | 1.4 | -0.6 | 4.00 | A | A | A |
| 37.65 | 2.22135 | -21.4 | 6.49 | N | N | N |
| 32.7 | 0.897 | -31.7 | 3.86 | N | N | N |
| 40.12 | 4.01 | -16.2 | 10.36 | A | A | A |
| 38.4 | 3 | -19.8 | 8.27 | A | A | A |
| 35 | 2 | -26.9 | 6.33 | N | N | N |
| 59 | 3 | 23.2 | 5.76 | N | N | N |
| 35.78 | 1.63 | -25.3 | 5.30 | N | N | N |
| 35.71 | 8.5 | -25.4 | 23.96 | N | N | N |
| 42.1 | 2.9 | -12.1 | 7.40 | A | A | A |
| 44 | 4.8 | -8.1 | 11.24 | A | A | A |
| 42.4 | 1.8 | -11.5 | 5.04 | A | A | A |
| 43 | 2.5 | -10.2 | 6.42 | A | A | A |
| 42.1 | 1.9 | -12.1 | 5.27 | A | A | A |
| 114.665 | 2.877 | 139.4 | 3.70 | N | N | N |
| 29.61 | 1.44 | -38.2 | 5.57 | N | N | N |
| 30.58 | 0.85 | -36.2 | 3.88 | N | N | N |
| 38 | 4 | -20.7 | 10.87 | N | A | N |
| 35 | 2 | -26.9 | 6.33 | N | N | N |
| 39.99 | 0.92 | -16.5 | 3.56 | A | N | W |
| 48.5 | 2.9 | 1.2 | 6.57 | A | A | A |
| 31 | 1 | -35.3 | 4.22 | N | N | N |
| 51.5 | 3.1 | 7.5 | 6.60 | A | A | A |
| 41.06 | 1.32 | -14.3 | 4.21 | A | N | W |
| 40.98 | 2.77 | -14.4 | 7.28 | A | A | A |
| 41.9 | 1.7 | -12.5 | 4.88 | A | A | A |
| 36.4 | 2.5 | -24.0 | 7.38 | N | N | N |

Labcode

437

441

444

445

446

448

450

452

453

459

460

462

463

464

467

478

480

482

483

484

491

495

496

499

506

508

516

522

527

530

532

537

555

Ra-226 determination of activity concentration in Sample 3

Table 50. Target values

| Parameter | Value |
| --- | --- |
| Activity concentration [ Bq/kg ] | 47.9 |
| Uncertainty [ Bq/kg ] | 1.3 |
| Maximum Acceptable Relative Bias [%] | 20 |

Accepted 31%	Not Accepted 18%	Warning 7%	Not Reported 44%

60

Figure 50. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

30

40

266

218

397

94

355

145

21

81

498

283

58

303

38

323

346

350

79

462

506

31

180

183

260

139

286

316

330

83

233

256

435

288

417

408

62

13

35

84

319

341

55

258

425

134

525

163

455

63

535

57

125

40

124

358

116

252

276

165

200

344

379

97

410

128

370

384

465

332

272

532

178

304

217

235

238

348

398

520

363

100

150

315

267

544

78

507

90

536

328

274

402

537

65

155

172

187

514

278

478

261

388

184

405

482

247

147

387

219

158

441

483

127

41

311

407

281

154

357

433

249

484

85

492

192

50

198

338

105

287

403

528

66

104

475

133

214

443

446

487

318

395

52

72

313

391

122

553

25

477

119

450

312

215

382

523

109

53

161

207

111

290

299

15

356

113

232

159

509

60

529

549

11

373

307

428

270

2

102

179

110

191

347

295

273

533

393

411

156

401

421

203

227

414

220

501

322

181

400

262

132

204

326

555

152

463

314

280

34

495

530

453

333

426

500

71

19

491

377

Laboratory code

Figure 51. Reported results and evaluation parameters for Ra-226. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 51. Evaluation results for Ra-226 in Sample 3, Rep. Value and Rep. Unc. in Bq/kg

Labcode

Reported value

[Bq/kg ]

Reported uncertainty

(k = 1) [Bq/kg ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

Activity concentration  [ Bq/kg ]

50

2	50.1	4.5	4.6	9.38	A	A	A

11	49.6	4.5	3.5	9.47	A	A	A

13	37.2	4.7	-22.3	12.92	N	A	N

| Reported value [Bq/kg ] 48.3 | Reported uncertainty (k = 1) [Bq/kg ] 2.8 | Relative bias [%] 0.8 | P-Test [%] 6.40 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 93 | 24 | 94.2 | 25.95 | N | N | N |
| 28 | 7 | -41.5 | 25.15 | N | N | N |
| 45.8 | 7 | -4.4 | 15.52 | A | A | A |
| 35.1 | 1.21 | -26.7 | 4.39 | N | N | N |
| 76.7 | 7 | 60.1 | 9.52 | N | N | N |
| 37.3 | 3.5 | -22.1 | 9.77 | N | A | N |
| 32.19 | 3.37 | -32.8 | 10.82 | N | N | N |
| 38.89 | 3.97 | -18.8 | 10.56 | A | A | A |
| 43.15 | 13.6 | -9.9 | 31.63 | A | N | W |
| 44.3 | 1.4 | -7.5 | 4.17 | A | A | A |
| 45.6 | 2.28 | -4.8 | 5.69 | A | A | A |
| 47.2 | 4.72 | -1.5 | 10.36 | A | A | A |
| 37.36 | 1.49 | -22.0 | 4.82 | N | N | N |
| 38.7 | 3.9 | -19.2 | 10.44 | A | A | A |
| 31.7 | 1.3 | -33.8 | 4.92 | N | N | N |
| 49.1 | 4 | 2.5 | 8.59 | A | A | A |
| 37.1 | 6.5 | -22.6 | 17.73 | N | A | N |
| 38.25 | 0.89 | -20.1 | 3.57 | N | N | N |
| 42 | 3 | -12.3 | 7.64 | A | A | A |
| 44.54 | 2.4445 | -7.0 | 6.12 | A | A | A |
| 86.83 | 4.34 | 81.3 | 5.69 | N | N | N |
| 45.6 | 1.74 | -4.8 | 4.68 | A | A | A |
| 41.52 | 4.88 | -13.3 | 12.06 | A | A | A |
| 35 | 2.4 | -26.9 | 7.37 | N | N | N |
| 30.38 | 3.42 | -36.6 | 11.58 | N | N | N |
| 36.2 | 5.25 | -24.4 | 14.75 | N | A | N |
| 37.3 | 16.51 | -22.1 | 44.35 | N | N | N |
| 43.726 | 3.11 | -8.7 | 7.61 | A | A | A |
| 41.7 | 1.9 | -12.9 | 5.30 | A | A | A |
| 24.35 | 1.28 | -49.2 | 5.92 | N | N | N |
| 39.6 | 4.3 | -17.3 | 11.19 | A | A | A |
| 41.2 | 3 | -14.0 | 7.77 | A | A | A |
| 50.2 | 4.7 | 4.8 | 9.75 | A | A | A |
| 44.6 | 1.5 | -6.9 | 4.32 | A | A | A |
| 44.4 | 1.5 | -7.3 | 4.33 | A | A | A |
| 46.9 | 0.3 | -2.1 | 2.79 | A | A | A |
| 50.7 | 3.1 | 5.8 | 6.69 | A | A | A |
| 47.94 | 3.89 | 0.1 | 8.56 | A | A | A |
| 48.44 | 12.11 | 1.1 | 25.15 | A | N | W |
| 39.16 | 2.74 | -18.2 | 7.50 | A | A | A |
| 46 | 5.3 | -4.0 | 11.84 | A | A | A |
| 45.7 | 4.6 | -4.6 | 10.43 | A | A | A |
| 39 | 4 | -18.6 | 10.61 | A | A | A |
| 38.8 | 2.34 | -19.0 | 6.61 | A | N | W |
| 43.1 | 2.6 | -10.0 | 6.61 | A | A | A |
| 39.68 | 2.94 | -17.2 | 7.89 | A | A | A |
| 59.2 | 3.7 | 23.6 | 6.81 | N | N | N |
| 45 | 3 | -6.0 | 7.20 | A | A | A |
| 37.5 | 2.7 | -21.7 | 7.69 | N | N | N |
| 35.7 | 1.59 | -25.5 | 5.22 | N | N | N |
| 25.5 | 4.08 | -46.8 | 16.23 | N | N | N |
| 42.594 | 0.1428 | -11.1 | 2.73 | A | N | W |
| 41.4 | 3.9 | -13.6 | 9.80 | A | A | A |
| 65.4 | 6.7 | 36.5 | 10.60 | N | N | N |
| 43.5 | 2.7 | -9.2 | 6.77 | A | A | A |
| 42 | 11 | -12.3 | 26.33 | A | N | W |
| 51.4 | 2.6 | 7.3 | 5.74 | A | A | A |
| 43 | 5 | -10.2 | 11.94 | A | A | A |
| 49 | 2.55 | 2.3 | 5.87 | A | A | A |

Labcode

15

19

21

25

31

34

35

38

40

41

50

52

53

55

57

58

60

62

63

65

66

71

72

78

79

81

83

84

85

90

94

97

100

102

104

105

109

110

111

113

116

119

122

124

125

127

128

132

133

134

139

145

147

150

152

154

155

156

158

159

| Reported value [Bq/kg ] 47.4 | Reported uncertainty (k = 1) [Bq/kg ] 3.2 | Relative bias [%] -1.0 | P-Test [%] 7.28 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 38 | 6 | -20.7 | 16.02 | N | A | N |
| 39.5 | 2.4 | -17.5 | 6.65 | A | N | W |
| 42 | 4 | -12.3 | 9.90 | A | A | A |
| 40.8 | 3.3 | -14.8 | 8.53 | A | A | A |
| 50.67 | 1.5 | 5.8 | 4.02 | A | A | A |
| 35.27 | 1.01 | -26.4 | 3.95 | N | N | N |
| 58 | 4 | 21.1 | 7.41 | N | N | N |
| 35.4 | 1 | -26.1 | 3.92 | N | N | N |
| 42.5 | 6.4 | -11.3 | 15.30 | A | A | A |
| 42 | 3 | -12.3 | 7.64 | A | A | A |
| 50.8 | 4.2 | 6.0 | 8.70 | A | A | A |
| 44 | 8 | -8.1 | 18.38 | A | A | A |
| 44.3 | 3.7 | -7.5 | 8.78 | A | A | A |
| 39.5 | 2 | -17.5 | 5.74 | A | N | W |
| 54.1 | 12 | 12.9 | 22.35 | A | N | W |
| 60 | 5 | 25.3 | 8.76 | N | N | N |
| 47.5 | 2.2 | -0.8 | 5.37 | A | A | A |
| 45 | 3 | -6.0 | 7.20 | A | A | A |
| 46.5 | 5.2 | -2.9 | 11.51 | A | A | A |
| 41 | 6 | -14.4 | 14.88 | A | A | A |
| 21.25897 | 11.15741 | -55.6 | 52.55 | N | N | N |
| 42.9 | 1.8 | -10.4 | 5.00 | A | A | A |
| 56.5 | 4.9 | 18.0 | 9.09 | A | A | A |
| 54.5 | 1.7 | 13.8 | 4.13 | A | N | W |
| 48.6 | 9.9 | 1.5 | 20.55 | A | N | W |
| 36.3 | 3.2 | -24.2 | 9.22 | N | N | N |
| 41 | 4 | -14.4 | 10.13 | A | A | A |
| 41 | 3.18 | -14.4 | 8.22 | A | A | A |
| 42.573 | 3.927 | -11.1 | 9.62 | A | A | A |
| 43.7 | 0.823984 | -8.8 | 3.30 | A | N | W |
| 39.25 | 2.07 | -18.1 | 5.93 | A | N | W |
| 36.31 | 2.33 | -24.2 | 6.97 | N | N | N |
| 37.4 | 1.2 | -21.9 | 4.20 | N | N | N |
| 35.6 | 4.7 | -25.7 | 13.48 | N | A | N |
| 42.3 | 1.6 | -11.7 | 4.66 | A | A | A |
| 58.4 | 1.1 | 21.9 | 3.30 | N | N | N |
| 2.22 | 0.36 | -95.4 | 16.44 | N | N | N |
| 41.46 | 4.07 | -13.4 | 10.18 | A | A | A |
| 50 | 4.5 | 4.4 | 9.40 | A | A | A |
| 40.3 | 2.4 | -15.9 | 6.54 | A | A | A |
| 51 | 3 | 6.5 | 6.48 | A | A | A |
| 41.9 | 3.327558 | -12.5 | 8.39 | A | A | A |
| 39.45 | 5.58 | -17.6 | 14.40 | A | A | A |
| 42.1 | 1.6 | -12.1 | 4.67 | A | N | W |
| 69 | 6.14 | 44.0 | 9.30 | N | N | N |
| 43.4 | 8 | -9.4 | 18.63 | A | A | A |
| 31.6 | 1.39 | -34.0 | 5.17 | N | N | N |
| 36.1 | 3.5 | -24.6 | 10.07 | N | A | N |
| 44.4 | 3 | -7.3 | 7.28 | A | A | A |
| 36.6 | 2.1 | -23.6 | 6.35 | N | N | N |
| 48 | 4.5 | 0.2 | 9.76 | A | A | A |
| 50.94 | 1.94 | 6.3 | 4.68 | A | A | A |
| 48 | 5 | 0.2 | 10.76 | A | A | A |
| 32 | 4 | -33.2 | 12.79 | N | N | N |
| 40.8 | 1.4 | -14.8 | 4.37 | A | N | W |
| 49.9 | 2.9 | 4.2 | 6.41 | A | A | A |
| 43.2 | 15.4 | -9.8 | 35.75 | A | N | W |
| 46.4 | 6.2 | -3.1 | 13.63 | A | A | A |
| 45.6 | 5 | -4.8 | 11.30 | A | A | A |

Labcode

161

163

165

172

178

179

180

181

183

184

187

191

192

198

200

203

204

207

214

215

217

218

219

220

227

232

233

235

238

247

249

252

256

258

260

261

262

266

267

270

272

273

274

276

278

280

281

283

286

287

288

290

295

299

303

304

307

311

312

313

| Reported value [Bq/kg ] 68.07 | Reported uncertainty (k = 1) [Bq/kg ] 4.36 | Relative bias [%] 42.1 | P-Test [%] 6.96 | Trueness evaluation N | Precision evaluation N | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 41.4 | 2.1 | -13.6 | 5.75 | A | A | A |
| 36.1 | 5.3 | -24.6 | 14.93 | N | A | N |
| 45.19 | 1.56 | -5.7 | 4.39 | A | A | A |
| 37.3 | 1.2 | -22.1 | 4.21 | N | N | N |
| 57.3 | 3.4 | 19.6 | 6.52 | A | N | W |
| 32.9 | 1.5 | -31.3 | 5.31 | N | N | N |
| 60.18 | 20.4 | 25.6 | 34.01 | N | N | N |
| 41.8 | 4.5 | -12.7 | 11.10 | A | A | A |
| 36.1 | 1.63 | -24.6 | 5.27 | N | N | N |
| 40.2 | 3.2 | -16.1 | 8.41 | A | A | A |
| 84.49 | 13.59 | 76.4 | 16.31 | N | N | N |
| 44.32 | 2.71 | -7.5 | 6.69 | A | A | A |
| 37.3 | 1 | -22.1 | 3.81 | N | N | N |
| 39.5 | 3.1 | -17.5 | 8.30 | A | A | A |
| 32.9 | 2 | -31.3 | 6.66 | N | N | N |
| 50.92 | 2.6 | 6.3 | 5.78 | A | A | A |
| 41 | 2 | -14.4 | 5.58 | A | N | W |
| 33.5 | 1 | -30.1 | 4.03 | N | N | N |
| 25.14 | 0.73 | -47.5 | 3.97 | N | N | N |
| 48.4 | 4.6 | 1.0 | 9.88 | A | A | A |
| 43.6 | 3.2 | -9.0 | 7.83 | A | A | A |
| 39.1 | 3.9 | -18.4 | 10.34 | A | A | A |
| 41.1 | 1.3 | -14.2 | 4.17 | A | N | W |
| 39.7 | 3.9 | -17.1 | 10.19 | A | A | A |
| 49.7 | 1.5 | 3.8 | 4.06 | A | A | A |
| 115.0968 | 12.06731 | 140.3 | 10.83 | N | N | N |
| 39.5 | 6.2 | -17.5 | 15.93 | A | A | A |
| 46.52 | 3.25 | -2.9 | 7.49 | A | A | A |
| 39.7 | 4.1 | -17.1 | 10.68 | A | A | A |
| 42.7 | 1.2 | -10.9 | 3.91 | A | N | W |
| 42.3 | 1.2 | -11.7 | 3.93 | A | N | W |
| 45.61 | 2.1 | -4.8 | 5.34 | A | A | A |
| 51.2 | 2.5 | 6.9 | 5.59 | A | A | A |
| 45.3 | 5.25 | -5.4 | 11.90 | A | A | A |
| 23.13 | 7.5 | -51.7 | 32.54 | N | N | N |
| 41 | 4.8 | -14.4 | 12.02 | A | A | A |
| 58 | 3 | 21.1 | 5.84 | N | N | N |
| 51.6 | 4.2 | 7.7 | 8.58 | A | A | A |
| 41.9 | 2.2 | -12.5 | 5.91 | A | A | A |
| 44.41 | 3.6 | -7.3 | 8.55 | A | A | A |
| 42.5 | 1.8 | -11.3 | 5.03 | A | A | A |
| 43.3 | 2.2 | -9.6 | 5.76 | A | A | A |
| 36.8 | 0.6 | -23.2 | 3.17 | N | N | N |
| 39.6 | 2.4 | -17.3 | 6.64 | A | N | W |
| 51.2 | 0.3 | 6.9 | 2.78 | A | A | A |
| 55.8 | 3 | 16.5 | 6.02 | A | N | W |
| 36.7 | 3.3 | -23.4 | 9.39 | N | A | N |
| 52.4 | 4.2 | 9.4 | 8.46 | A | A | A |
| 37.4 | 5 | -21.9 | 13.64 | N | A | N |
| 84.66 | 16.89 | 76.7 | 20.13 | N | N | N |
| 49.9 | 3.5 | 4.2 | 7.52 | A | A | A |
| 43.61 | 0.89 | -9.0 | 3.40 | A | N | W |
| 36.5 | 2 | -23.8 | 6.11 | N | N | N |
| 43 | 2.5 | -10.2 | 6.42 | A | A | A |
| 45 | 2 | -6.0 | 5.21 | A | A | A |
| 45 | 3.8 | -6.0 | 8.87 | A | A | A |
| 46 | 3 | -4.0 | 7.06 | A | A | A |
| 83.4 | 17.6 | 74.1 | 21.28 | N | N | N |
| 38 | 2.4 | -20.7 | 6.87 | N | N | N |

Labcode

314

315

316

318

319

322

323

326

328

330

332

333

338

341

344

346

347

348

350

355

356

357

358

363

370

373

377

379

382

384

387

388

391

393

395

397

398

400

401

402

403

405

407

408

410

411

414

417

421

425

426

428

433

435

441

443

446

450

453

455

| Reported value [Bq/kg ] 35 | Reported uncertainty (k = 1) [Bq/kg ] 3 | Relative bias [%] -26.9 | P-Test [%] 8.99 | Trueness evaluation N | Precision evaluation N | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 68 | 8 | 42.0 | 12.07 | N | N | N |
| 39.8 | 4 | -16.9 | 10.41 | A | A | A |
| 44.9 | 3.6 | -6.3 | 8.46 | A | A | A |
| 45.8 | 1.6 | -4.4 | 4.42 | A | A | A |
| 42.1 | 2.9 | -12.1 | 7.40 | A | A | A |
| 42.5 | 1.8 | -11.3 | 5.03 | A | A | A |
| 43 | 2.5 | -10.2 | 6.42 | A | A | A |
| 43.7 | 3.5 | -8.8 | 8.46 | A | A | A |
| 45 | 4.5 | -6.0 | 10.36 | A | A | A |
| 108.392 | 1.841 | 126.3 | 3.20 | N | N | N |
| 43.75 | 1.7 | -8.7 | 4.74 | A | A | A |
| 79.72 | 6.5 | 66.4 | 8.59 | N | N | N |
| 30.85 | 2.31 | -35.6 | 7.96 | N | N | N |
| 85.13 | 40.67 | 77.7 | 47.85 | N | N | N |
| 56.7 | 10.8 | 18.4 | 19.24 | A | A | A |
| 35 | 3 | -26.9 | 8.99 | N | N | N |
| 41.643 | 0.064 | -13.1 | 2.72 | A | N | W |
| 49 | 0.8 | 2.3 | 3.17 | A | A | A |
| 42 | 3 | -12.3 | 7.64 | A | A | A |
| 41.08 | 2.59 | -14.2 | 6.86 | A | A | A |
| 46.7 | 1.8 | -2.5 | 4.71 | A | A | A |
| 37.6 | 2.3 | -21.5 | 6.69 | N | N | N |
| 44.5 | 2 | -7.1 | 5.25 | A | A | A |
| 49.1 | 2.7 | 2.5 | 6.13 | A | A | A |
| 81.59 | 10.54 | 70.3 | 13.20 | N | N | N |
| 40.63 | 2.75 | -15.2 | 7.29 | A | A | A |
| 51 | 2.3 | 6.5 | 5.26 | A | A | A |
| 38.4 | 1.4 | -19.8 | 4.55 | A | N | W |
| 41.7 | 2 | -12.9 | 5.51 | A | A | A |
| 41.9 | 1.7 | -12.5 | 4.88 | A | A | A |
| 41.51 | 1.41 | -13.3 | 4.35 | A | N | W |
| 49.4 | 2.5 | 3.1 | 5.74 | A | A | A |
| 45.7 | 2.7 | -4.6 | 6.50 | A | A | A |
| 63.4 | 7 | 32.4 | 11.37 | N | N | N |

Labcode

462

463

465

475

477

478

482

483

484

487

491

492

495

498

500

501

506

507

509

514

520

523

525

528

529

530

532

533

535

536

537

544

549

553

555

Activity concentration  [ Bq/kg ]

35

40

Ac-228 determination of activity concentration in Sample 3

Table 52. Target values

| Parameter | Value |
| --- | --- |
| Activity concentration [ Bq/kg ] | 36.4 |
| Uncertainty [ Bq/kg ] | 1.7 |
| Maximum Acceptable Relative Bias [%] | 20 |

Accepted 52%	Not Accepted 6%	Warning 2%	Not Reported 40%

45

50

Figure 52. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

25

30

94

400

283

355

185

102

189

369

303

496

271

172

268

437

249

151

121

233

165

40

535

230

45

453

444

495

480

32

358

100

9

395

314

421

478

155

21

156

316

154

482

110

288

364

408

276

530

104

200

278

425

344

417

300

134

479

67

371

506

62

64

132

441

537

384

17

55

339

140

184

175

410

107

462

20

178

393

256

70

315

128

402

170

527

15

516

325

487

370

202

83

19

122

465

49

187

290

119

226

161

318

223

53

387

78

97

356

261

124

536

273

338

253

311

433

203

133

326

333

Laboratory code

Figure 53. Reported results and evaluation parameters for Ac-228. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 53. Evaluation results for Ac-228 in Sample 3, Rep. Value and Rep. Unc. in Bq/kg

| Reported value | Reported uncertainty | Relative bias | P-Test | Trueness | Precision | Final |
| --- | --- | --- | --- | --- | --- | --- |
| [Bq/kg ] | (k = 1) [Bq/kg ] | [%] | [%] | evaluation | evaluation | Score |
| 37.8 | 3.9 | 3.9 | 11.33 | A | A | A |
| 35.87 | 2.33 | -1.5 | 8.00 | A | A | A |
| 35.3 | 4.2 | -3.0 | 12.78 | A | A | A |

Labcode

2

8

9

| Reported value [Bq/kg ] 42.4 | Reported uncertainty (k = 1) [Bq/kg ] 2.61 | Relative bias [%] 16.5 | P-Test [%] 7.73 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 38.1 | 2.3 | 4.7 | 7.63 | A | A | A |
| 38.9 | 2.1 | 6.9 | 7.14 | A | A | A |
| 37.43 | 1.88 | 2.8 | 6.86 | A | A | A |
| 39.3 | 8.1 | 8.0 | 21.13 | A | N | W |
| 38.05 | 1.77 | 4.5 | 6.59 | A | A | A |
| 36 | 2 | -1.1 | 7.26 | A | A | A |
| 33.4 | 173 | -8.2 | 517.99 | A | N | W |
| 35 | 1 | -3.9 | 5.47 | A | A | A |
| 38.78 | 3.49 | 6.5 | 10.14 | A | A | A |
| 35.09 | 1.86 | -3.6 | 7.06 | A | A | A |
| 31.9 | 2 | -12.4 | 7.82 | A | A | A |
| 34.17 | 2.57 | -6.1 | 8.85 | A | A | A |
| 36.4 | 2.03 | 0.0 | 7.27 | A | A | A |
| 34.53 | 1.38 | -5.1 | 6.15 | A | A | A |
| 39.2 | 2.3 | 7.7 | 7.50 | A | A | A |
| 40 | 5 | 9.9 | 13.34 | A | A | A |
| 34.2 | 1.3 | -6.0 | 6.02 | A | A | A |
| 40.5 | 1.72 | 11.3 | 6.31 | A | A | A |
| 40.68 | 3.14 | 11.8 | 9.02 | A | A | A |
| 37.5 | 1.9 | 3.0 | 6.89 | A | A | A |
| 37.49 | 0.98 | 3.0 | 5.35 | A | A | A |
| 27.1 | 2.7 | -25.6 | 11.00 | N | A | N |
| 34 | 1.7 | -6.6 | 6.84 | A | A | A |
| 38 | 4 | 4.4 | 11.52 | A | A | A |
| 37.1 | 6.5 | 1.9 | 18.13 | A | A | A |
| 36.77 | 1.54 | 1.0 | 6.27 | A | A | A |
| 37.2 | 2.2 | 2.2 | 7.54 | A | A | A |
| 40 | 3 | 9.9 | 8.84 | A | A | A |
| 36.9 | 4.5 | 1.4 | 13.06 | A | A | A |
| 40.3 | 1.9 | 10.7 | 6.64 | A | A | A |
| 38.3 | 1.9 | 5.2 | 6.81 | A | A | A |
| 34.4 | 1.91 | -5.5 | 7.26 | A | A | A |
| 41.5 | 4.23 | 14.0 | 11.21 | A | A | A |
| 38.5 | 1.4 | 5.8 | 5.92 | A | A | A |
| 35.13 | 4.6 | -3.5 | 13.90 | A | A | A |
| 39.2 | 5.7 | 7.7 | 15.27 | A | A | A |
| 14.74 | 3.74 | -59.5 | 25.80 | N | N | N |
| 41.533 | 1.787 | 14.1 | 6.35 | A | A | A |
| 37 | 2.6 | 1.6 | 8.44 | A | A | A |
| 11.07 | 1.24 | -69.6 | 12.14 | N | N | N |
| 41.7 | 4.6 | 14.6 | 11.98 | A | A | A |
| 35.2 | 1.8 | -3.3 | 6.93 | A | A | A |
| 29.3 | 2.9 | -19.5 | 10.94 | A | A | A |
| 36.4 | 1.7 | 0.0 | 6.60 | A | A | A |
| 34.8 | 1.4 | -4.4 | 6.16 | A | A | A |
| 38 | 3.1 | 4.4 | 9.40 | A | A | A |
| 36.6 | 0.7 | 0.6 | 5.05 | A | A | A |
| 36.2 | 2.4 | -0.6 | 8.11 | A | A | A |
| 36.36 | 3.04 | -0.1 | 9.58 | A | A | A |
| 42.55 | 10.64 | 16.9 | 25.44 | A | N | W |
| 36.4 | 3.1 | 0.0 | 9.71 | A | A | A |
| 40.28 | 3.6 | 10.7 | 10.08 | A | A | A |
| 33.3 | 3 | -8.5 | 10.15 | A | A | A |
| 39.5 | 3.7 | 8.5 | 10.47 | A | A | A |
| 42 | 6 | 15.4 | 15.03 | A | A | A |
| 38.56 | 2.04 | 5.9 | 7.06 | A | A | A |
| 30.54639 | 0.67258 | -16.1 | 5.16 | A | N | W |
| 37.3 | 3.2 | 2.5 | 9.77 | A | A | A |
| 48 | 4 | 31.9 | 9.55 | N | N | N |

Labcode

11

13

15

17

19

20

21

24

25

28

32

34

40

41

45

48

49

50

52

53

54

55

57

58

60

62

63

64

65

67

69

70

72

78

79

81

83

84

85

90

94

97

100

102

104

105

107

109

110

111

113

114

119

121

122

124

128

131

132

133

| Reported value [Bq/kg ] 36.7 | Reported uncertainty (k = 1) [Bq/kg ] 0.9 | Relative bias [%] 0.8 | P-Test [%] 5.28 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 36 | 1.51 | -1.1 | 6.28 | A | A | A |
| 37.7 | 2.2 | 3.6 | 7.47 | A | A | A |
| 28.5 | 4.56 | -21.7 | 16.67 | N | A | N |
| 38.4 | 2.2 | 5.5 | 7.39 | A | A | A |
| 33.22 | 2.33 | -8.7 | 8.43 | A | A | A |
| 36.1 | 2.1 | -0.8 | 7.46 | A | A | A |
| 35.9 | 2.6 | -1.4 | 8.62 | A | A | A |
| 36 | 0.69 | -1.1 | 5.05 | A | A | A |
| 36 | 0.936 | -1.1 | 5.35 | A | A | A |
| 40.4 | 2.6 | 11.0 | 7.95 | A | A | A |
| 34 | 2 | -6.6 | 7.51 | A | A | A |
| 35.2 | 2.8 | -3.3 | 9.22 | A | A | A |
| 38.8 | 1.8 | 6.6 | 6.58 | A | A | A |
| 32 | 4 | -12.1 | 13.34 | A | A | A |
| 37.9 | 2.1 | 4.1 | 7.25 | A | A | A |
| 38.1 | 3.6 | 4.7 | 10.54 | A | A | A |
| 57.72 | 0.88 | 58.6 | 4.91 | N | N | N |
| 37.8 | 5.7 | 3.9 | 15.79 | A | A | A |
| 28.61 | 0.34 | -21.4 | 4.82 | N | N | N |
| 40 | 3 | 9.9 | 8.84 | A | A | A |
| 30.2 | 2.2 | -17.0 | 8.65 | A | A | A |
| 39.3 | 3.4 | 8.0 | 9.83 | A | A | A |
| 37.03 | 5.18 | 1.7 | 14.75 | A | A | A |
| 38.8 | 3.1 | 6.6 | 9.25 | A | A | A |
| 36.4 | 1.5 | 0.0 | 6.23 | A | A | A |
| 39.11 | 0.49 | 7.4 | 4.84 | A | A | A |
| 45.8 | 5 | 25.8 | 11.87 | N | A | N |
| 31 | 3 | -14.8 | 10.75 | A | A | A |
| 37.8 | 2.51 | 3.9 | 8.12 | A | A | A |
| 33.25 | 1.32 | -8.7 | 6.13 | A | A | A |
| 35.5 | 2.76 | -2.5 | 9.07 | A | A | A |
| 38.2 | 4.3 | 5.0 | 12.19 | A | A | A |
| 36.5 | 5 | 0.3 | 14.47 | A | A | A |
| 34.50817 | 1.19534 | -5.2 | 5.81 | A | A | A |
| 40.6 | 4.1 | 11.5 | 11.13 | A | A | A |
| 27.8 | 2.5 | -23.6 | 10.13 | N | A | N |
| 40.3 | 2.37 | 10.7 | 7.51 | A | A | A |
| 34.4 | 1.2 | -5.5 | 5.83 | A | A | A |
| 37.4 | 2.1 | 2.8 | 7.30 | A | A | A |
| 33.7 | 2 | -7.4 | 7.55 | A | A | A |
| 38.8 | 3.1 | 6.6 | 9.25 | A | A | A |
| 36.14 | 1.01 | -0.7 | 5.44 | A | A | A |
| 36.308 | 3.008 | -0.2 | 9.51 | A | A | A |
| 33.1 | 0.606403 | -9.1 | 5.02 | A | A | A |
| 43.7 | 3.6 | 20.0 | 9.47 | N | A | N |
| 35.9 | 1.34 | -1.4 | 5.98 | A | A | A |
| 38.2 | 2.37 | 5.0 | 7.77 | A | A | A |
| 38.9 | 1.1 | 6.9 | 5.46 | A | A | A |
| 41.7 | 2.4 | 14.6 | 7.41 | A | A | A |
| 41.8 | 1.4 | 14.8 | 5.75 | A | N | W |
| 45.3 | 1.1 | 24.4 | 5.26 | N | N | N |
| 36.2 | 4.3 | -0.6 | 12.76 | A | A | A |
| 39.87 | 3.79 | 9.5 | 10.59 | A | A | A |
| 32.5 | 1.63 | -10.7 | 6.85 | A | A | A |
| 40 | 2 | 9.9 | 6.84 | A | A | A |
| 31.9 | 10.4 | -12.4 | 32.93 | A | N | W |
| 35.4 | 2.3 | -2.8 | 8.00 | A | A | A |
| 43 | 3 | 18.1 | 8.40 | A | A | A |
| 36.33 | 2.72 | -0.2 | 8.82 | A | A | A |

Labcode

134

139

140

145

150

151

154

155

156

159

161

165

168

170

172

175

178

179

184

185

187

189

191

197

198

200

202

203

204

208

211

213

215

217

218

223

224

226

230

232

233

235

242

247

249

253

254

256

258

260

261

262

265

267

268

270

271

272

273

276

| Reported value [Bq/kg ] 36.43 | Reported uncertainty (k = 1) [Bq/kg ] 1.8 | Relative bias [%] 0.1 | P-Test [%] 6.80 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 40.4 | 2.2 | 11.0 | 7.17 | A | A | A |
| 27.7 | 1.31 | -23.9 | 6.65 | N | N | N |
| 43.04 | 2.48 | 18.2 | 7.42 | A | A | A |
| 37.9 | 6.2 | 4.1 | 17.01 | A | A | A |
| 39.5 | 1.2 | 8.5 | 5.57 | A | A | A |
| 36.2 | 2.16 | -0.6 | 7.58 | A | A | A |
| 40 | 2 | 9.9 | 6.84 | A | A | A |
| 39.106 | 4.888 | 7.4 | 13.34 | A | A | A |
| 32.5 | 2.1 | -10.7 | 7.97 | A | A | A |
| 35.6 | 3.5 | -2.2 | 10.88 | A | A | A |
| 40 | 3 | 9.9 | 8.84 | A | A | A |
| 36.66 | 1.05 | 0.7 | 5.48 | A | A | A |
| 31 | 4 | -14.8 | 13.72 | A | A | A |
| 36.1 | 1.7 | -0.8 | 6.63 | A | A | A |
| 43.6 | 3.5 | 19.8 | 9.29 | A | A | A |
| 40.7 | 3.2 | 11.8 | 9.14 | A | A | A |
| 44.16 | 1.35 | 21.3 | 5.58 | N | N | N |
| 40.3 | 5.2 | 10.7 | 13.72 | A | A | A |
| 38.1 | 2.4 | 4.7 | 7.84 | A | A | A |
| 35.51 | 1.34 | -2.5 | 6.00 | A | A | A |
| 38.45 | 1.84 | 5.6 | 6.69 | A | A | A |
| 36 | 6 | -1.1 | 17.31 | A | A | A |
| 53.4 | 2.2 | 46.7 | 6.23 | N | N | N |
| 40.44 | 1.6 | 11.1 | 6.12 | A | A | A |
| 33.1 | 1.4 | -9.1 | 6.30 | A | A | A |
| 39 | 3.32 | 7.1 | 9.71 | A | A | A |
| 54 | 2.7 | 48.4 | 6.84 | N | N | N |
| 36.7 | 3.5 | 0.8 | 10.62 | A | A | A |
| 31.9 | 4 | -12.4 | 13.38 | A | A | A |
| 30.06 | 1.6 | -17.4 | 7.08 | A | A | A |
| 38 | 3.5 | 4.4 | 10.33 | A | A | A |
| 68.95 | 5.53 | 89.4 | 9.28 | N | N | N |
| 43.13 | 2.63 | 18.5 | 7.68 | A | A | A |
| 37.5 | 1.3 | 3.0 | 5.82 | A | A | A |
| 37.3 | 1.2 | 2.5 | 5.67 | A | A | A |
| 36.6 | 2.6 | 0.6 | 8.50 | A | A | A |
| 38.02 | 1.91 | 4.4 | 6.86 | A | A | A |
| 39 | 2 | 7.1 | 6.94 | A | A | A |
| 27.96 | 0.66 | -23.2 | 5.23 | N | N | N |
| 41.7 | 2.9 | 14.6 | 8.38 | A | A | A |
| 34.7 | 1.9 | -4.7 | 7.20 | A | A | A |
| 35.1 | 3.5 | -3.6 | 11.01 | A | A | A |
| 36.3 | 2.1 | -0.3 | 7.44 | A | A | A |
| 30.6 | 0.89 | -15.9 | 5.50 | A | N | W |
| 39.1 | 2.5 | 7.4 | 7.92 | A | A | A |
| 36.99 | 3.7 | 1.6 | 11.04 | A | A | A |
| 36.6 | 2.5 | 0.6 | 8.27 | A | A | A |
| 40.6 | 1.6 | 11.5 | 6.11 | A | A | A |
| 37.5 | 4.8 | 3.0 | 13.63 | A | A | A |
| 39.08 | 2.49 | 7.4 | 7.90 | A | A | A |
| 37.46 | 2.61 | 2.9 | 8.39 | A | A | A |
| 37.4 | 3.8 | 2.8 | 11.18 | A | A | A |
| 33.05 | 3.43 | -9.2 | 11.38 | A | A | A |
| 40.9 | 1.2 | 12.4 | 5.52 | A | A | A |
| 41 | 1.2 | 12.6 | 5.51 | A | A | A |
| 34.54 | 0.7 | -5.1 | 5.09 | A | A | A |
| 38.1 | 1.9 | 4.7 | 6.83 | A | A | A |
| 35.45 | 4.18 | -2.6 | 12.68 | A | A | A |
| 36.3 | 4 | -0.3 | 11.97 | A | A | A |

Labcode

278

281

283

284

286

287

288

290

291

293

294

299

300

303

304

307

310

311

312

313

314

315

316

317

318

319

325

326

328

329

330

332

333

338

339

341

344

347

348

355

356

357

358

364

369

370

371

373

376

379

381

382

384

386

387

388

391

393

395

398

| Reported value [Bq/kg ] 19 | Reported uncertainty (k = 1) [Bq/kg ] 5 | Relative bias [%] -47.8 | P-Test [%] 26.73 | Trueness evaluation N | Precision evaluation N | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 38.7 | 3.1 | 6.3 | 9.27 | A | A | A |
| 36.96 | 2.25 | 1.5 | 7.67 | A | A | A |
| 38.6 | 2 | 6.0 | 6.98 | A | A | A |
| 36.3 | 1.1 | -0.3 | 5.57 | A | A | A |
| 37.9 | 2.3 | 4.1 | 7.66 | A | A | A |
| 36.68 | 4.08 | 0.8 | 12.06 | A | A | A |
| 43.9 | 2.2 | 20.6 | 6.85 | N | N | N |
| 36.6 | 2.2 | 0.6 | 7.61 | A | A | A |
| 35.6 | 2.8 | -2.2 | 9.15 | A | A | A |
| 36.5 | 2.3 | 0.3 | 7.84 | A | A | A |
| 36.29 | 2.09 | -0.3 | 7.41 | A | A | A |
| 37.2 | 1.39 | 2.2 | 5.98 | A | A | A |
| 39.2 | 1.9 | 7.7 | 6.73 | A | A | A |
| 44.67 | 2.51 | 22.7 | 7.31 | N | N | N |
| 32.6 | 4.2 | -10.4 | 13.70 | A | A | A |
| 37.3 | 2.6 | 2.5 | 8.39 | A | A | A |
| 34.77 | 1.85 | -4.5 | 7.08 | A | A | A |
| 36.6 | 2.6 | 0.6 | 8.50 | A | A | A |
| 35.6 | 1.6 | -2.2 | 6.48 | A | A | A |
| 28.61 | 0.34 | -21.4 | 4.82 | N | N | N |
| 37.4 | 1.1 | 2.8 | 5.52 | A | A | A |
| 35.09 | 2.1054 | -3.6 | 7.60 | A | A | A |
| 34.6 | 0.955 | -5.0 | 5.42 | A | A | A |
| 37.1 | 1.6 | 1.9 | 6.36 | A | A | A |
| 35.01 | 2.93 | -3.8 | 9.58 | A | A | A |
| 41.9 | 3 | 15.1 | 8.55 | A | A | A |
| 38 | 3 | 4.4 | 9.17 | A | A | A |
| 34 | 2 | -6.6 | 7.51 | A | A | A |
| 38.25 | 1.54 | 5.1 | 6.17 | A | A | A |
| 39.8 | 4 | 9.3 | 11.08 | A | A | A |
| 35.7 | 1.9 | -1.9 | 7.08 | A | A | A |
| 36.71 | 3.16 | 0.8 | 9.79 | A | A | A |
| 35 | 4 | -3.9 | 12.35 | A | A | A |
| 36.1 | 1.3 | -0.8 | 5.90 | A | A | A |
| 37.3 | 2.6 | 2.5 | 8.39 | A | A | A |
| 36.4 | 1.9 | 0.0 | 7.00 | A | A | A |
| 39 | 2 | 7.1 | 6.94 | A | A | A |
| 82.397 | 1.914 | 126.4 | 5.22 | N | N | N |
| 34.98 | 2.39 | -3.9 | 8.28 | A | A | A |
| 31.11 | 1.1 | -14.5 | 5.86 | A | A | A |
| 32 | 3 | -12.1 | 10.47 | A | A | A |
| 44.6 | 1.7 | 22.5 | 6.03 | N | N | N |
| 37 | 2 | 1.6 | 7.14 | A | A | A |
| 36 | 1.7 | -1.1 | 6.64 | A | A | A |
| 38.9 | 2.7 | 6.9 | 8.37 | A | A | A |
| 31 | 1.2 | -14.8 | 6.07 | A | A | A |
| 38.8 | 2.7 | 6.6 | 8.38 | A | A | A |
| 36.38 | 1.23 | -0.0 | 5.77 | A | A | A |
| 41.77 | 3.31 | 14.8 | 9.20 | A | A | A |
| 47.5 | 3.1 | 30.5 | 8.03 | N | N | N |
| 34.2 | 1.7 | -6.0 | 6.82 | A | A | A |
| 42.5 | 1.4 | 16.8 | 5.72 | A | N | W |
| 37.3 | 1.9 | 2.5 | 6.91 | A | A | A |
| 38.9 | 3.1 | 6.9 | 9.24 | A | A | A |

Labcode

400

402

403

407

408

410

412

414

417

421

425

426

429

430

433

437

441

444

445

446

448

450

452

453

455

459

460

462

463

464

465

478

479

480

482

483

484

487

491

495

496

499

501

506

508

516

522

527

530

532

533

535

536

537

555

Th-232 determination of activity concentration in Sample 3

Table 54. Target values

| Parameter | Value |
| --- | --- |
| Activity concentration [ Bq/kg ] | 36.4 |
| Uncertainty [ Bq/kg ] | 1.7 |
| Maximum Acceptable Relative Bias [%] | 20 |

Accepted 18%	Not Accepted 3%	Warning 1%	Not Reported 78%

45

50

Figure 54. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

25

30

131

84

71

252

57

145

330

314

183

206

319

125

477

539

397

350

463

50

443

346

72

391

79

218

525

105

498

238

363

506

100

272

274

292

405

411

159

31

482

90

410

508

104

200

344

2

403

373

341

537

384

382

55

523

286

116

178

215

315

407

128

235

198

15

122

214

348

428

83

267

465

119

402

52

53

507

446

261

147

338

133

323

193

179

87

491

Laboratory code

Figure 55. Reported results and evaluation parameters for Th-232. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 55. Evaluation results for Th-232 in Sample 3, Rep. Value and Rep. Unc. in Bq/kg

Labcode

Reported value

[Bq/kg ]

Reported uncertainty

(k = 1) [Bq/kg ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

Activity concentration  [ Bq/kg ]

35

40

2	36.7	1.5	0.8	6.21	A	A	A

15	38.9	2.1	6.9	7.14	A	A	A

31	36.2	1.35	-0.6	5.98	A	A	A

| Reported value [Bq/kg ] 34.2 | Reported uncertainty (k = 1) [Bq/kg ] 1.3 | Relative bias [%] -6.0 | P-Test [%] 6.02 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 40.5 | 1.72 | 11.3 | 6.31 | A | A | A |
| 40.68 | 3.14 | 11.8 | 9.02 | A | A | A |
| 37.49 | 0.98 | 3.0 | 5.35 | A | A | A |
| 27.1 | 2.7 | -25.6 | 11.00 | N | A | N |
| 18.33 | 0.92 | -49.6 | 6.86 | N | N | N |
| 34.4 | 1.91 | -5.5 | 7.26 | A | A | A |
| 34.5 | 3.9 | -5.2 | 12.23 | A | A | A |
| 39.2 | 5.7 | 7.7 | 15.27 | A | A | A |
| 14.93 | 4.01 | -59.0 | 27.26 | N | N | N |
| 36.3 | 2.9 | -0.3 | 9.25 | A | A | A |
| 35.1 | 1.6 | -3.6 | 6.53 | A | A | A |
| 36.4 | 1.7 | 0.0 | 6.60 | A | A | A |
| 34.8 | 1.4 | -4.4 | 6.16 | A | A | A |
| 37.87 | 2.65 | 4.0 | 8.41 | A | A | A |
| 40.28 | 3.6 | 10.7 | 10.08 | A | A | A |
| 39 | 3.9 | 7.1 | 11.04 | A | A | A |
| 33.3 | 2.26 | -8.5 | 8.24 | A | A | A |
| 38.56 | 2.04 | 5.9 | 7.06 | A | A | A |
| 7.712536 | 2.177005 | -78.8 | 28.61 | N | N | N |
| 45 | 3 | 23.6 | 8.14 | N | N | N |
| 28.5 | 4.56 | -21.7 | 16.67 | N | A | N |
| 42.295 | 0.1305 | 16.2 | 4.68 | A | N | W |
| 36 | 0.936 | -1.1 | 5.35 | A | A | A |
| 38.1 | 3.6 | 4.7 | 10.54 | A | A | A |
| 57.72 | 0.88 | 58.6 | 4.91 | N | N | N |
| 32.3 | 2.2 | -11.3 | 8.26 | A | A | A |
| 48.34 | 2.42 | 32.8 | 6.85 | N | N | N |
| 38.8 | 3.1 | 6.6 | 9.25 | A | A | A |
| 36.4 | 1.5 | 0.0 | 6.23 | A | A | A |
| 33 | 3 | -9.3 | 10.22 | A | A | A |
| 39 | 2 | 7.1 | 6.94 | A | A | A |
| 38.2 | 4.3 | 5.0 | 12.19 | A | A | A |
| 34.50817 | 1.19534 | -5.2 | 5.81 | A | A | A |
| 38.8 | 3.1 | 6.6 | 9.25 | A | A | A |
| 34.9 | 2.63 | -4.1 | 8.87 | A | A | A |
| 26.91 | 0.93 | -26.1 | 5.81 | N | N | N |
| 41.8 | 1.4 | 14.8 | 5.75 | A | N | W |
| 39.72 | 4.61 | 9.1 | 12.51 | A | A | A |
| 35.4 | 2.3 | -2.8 | 8.00 | A | A | A |
| 35.6 | 2.199256 | -2.2 | 7.74 | A | A | A |
| 37.6 | 5.5 | 3.3 | 15.36 | A | A | A |
| 35.6 | 1.6 | -2.2 | 6.48 | A | A | A |
| 30.22 | 0.94 | -17.0 | 5.61 | A | N | W |
| 38.45 | 1.84 | 5.6 | 6.69 | A | A | A |
| 33.1 | 1.4 | -9.1 | 6.30 | A | A | A |
| 47.3 | 2.5 | 30.0 | 7.05 | N | N | N |
| 30.06 | 1.6 | -17.4 | 7.08 | A | A | A |
| 43.13 | 2.63 | 18.5 | 7.68 | A | A | A |
| 37.3 | 1.2 | 2.5 | 5.67 | A | A | A |
| 36.6 | 2.6 | 0.6 | 8.50 | A | A | A |
| 34.3 | 1.3 | -5.8 | 6.01 | A | A | A |
| 39 | 2 | 7.1 | 6.94 | A | A | A |
| 34 | 1 | -6.6 | 5.52 | A | A | A |
| 34.9 | 1.1 | -4.1 | 5.63 | A | A | A |
| 37.2 | 2.5 | 2.2 | 8.18 | A | A | A |
| 37.46 | 2.61 | 2.9 | 8.39 | A | A | A |
| 37.4 | 3.8 | 2.8 | 11.18 | A | A | A |
| 34.47 | 0.47 | -5.3 | 4.87 | A | A | A |
| 33.98 | 1.86 | -6.7 | 7.20 | A | A | A |

Labcode

50

52

53

55

57

71

72

79

83

84

90

100

104

105

116

119

122

125

128

131

133

145

147

159

178

179

183

193

198

200

206

214

215

218

235

238

252

261

267

272

274

286

292

314

315

319

323

330

338

341

344

346

348

350

363

373

382

384

391

397

| Reported value [Bq/kg ] 40.3 | Reported uncertainty (k = 1) [Bq/kg ] 2.1 | Relative bias [%] 10.7 | P-Test [%] 7.00 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 36.96 | 2.25 | 1.5 | 7.67 | A | A | A |
| 35.8 | 1.7 | -1.6 | 6.66 | A | A | A |
| 38.5 | 2.3 | 5.8 | 7.58 | A | A | A |
| 36.3 | 2.2 | -0.3 | 7.65 | A | A | A |
| 35.9 | 1.8 | -1.4 | 6.85 | A | A | A |
| 39 | 2.8 | 7.1 | 8.56 | A | A | A |
| 34.2 | 1.7 | -6.0 | 6.82 | A | A | A |
| 41.6 | 6.5 | 14.3 | 16.31 | A | A | A |
| 34 | 2 | -6.6 | 7.51 | A | A | A |
| 39.8 | 4 | 9.3 | 11.08 | A | A | A |
| 33.3 | 1.5 | -8.5 | 6.49 | A | A | A |
| 36.2 | 1.3 | -0.6 | 5.89 | A | A | A |
| 76.373 | 1.254 | 109.8 | 4.95 | N | N | N |
| 34.85 | 3.56 | -4.3 | 11.23 | A | A | A |
| 35 | 2 | -3.9 | 7.38 | A | A | A |
| 41.006 | 0.891 | 12.7 | 5.15 | A | A | A |
| 37.5 | 2.1 | 3.0 | 7.29 | A | A | A |
| 34.7 | 1.8 | -4.7 | 6.98 | A | A | A |
| 37.3 | 1.9 | 2.5 | 6.91 | A | A | A |
| 33.88 | 2.2 | -6.9 | 8.00 | A | A | A |

Labcode

402

403

405

407

410

411

428

443

446

463

465

477

482

491

498

506

507

523

525

537

539

Pa-234m determination of activity concentration in Sample 3

Table 56. Target values

| Parameter | Value |
| --- | --- |
| Activity concentration [ Bq/kg ] | 27.0 |
| Uncertainty [ Bq/kg ] | 0.4 |
| Maximum Acceptable Relative Bias [%] | 20 |

Accepted 2%	Not Accepted 3%	Warning 1%	Not Reported 94%

35

Figure 56. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

20

25

25

226

348

537

15

421

452

315

110

150

220

276

414

78

19

116

341

90

395

403

384

326

52

491

Laboratory code

Figure 57. Reported results and evaluation parameters for Pa-234m. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 57. Evaluation results for Pa-234m in Sample 3, Rep. Value and Rep. Unc. in Bq/kg

Labcode

Reported value

[Bq/kg ]

Reported uncertainty

(k = 1) [Bq/kg ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

Activity concentration  [ Bq/kg ]

30

15	25.5	2.4	-5.6	9.53	A	A	A

19	37.9	7.6	40.4	20.11	N	N	N

25	21.7	5.3	-19.6	24.47	A	N	W

Table 57. Evaluation results for Pa-234m in Sample 3, Rep. Value and Rep. Unc. in Bq/kg

| Reported value [Bq/kg ] 67.6 | Reported uncertainty (k = 1) [Bq/kg ] 3.03 | Relative bias [%] 150.4 | P-Test [%] 4.72 | Trueness evaluation N | Precision evaluation N | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 37.2 | 8.52 | 37.8 | 22.95 | N | N | N |
| 40.4 | 1.8 | 49.6 | 4.70 | N | N | N |
| 28.2 | 7.2 | 4.4 | 25.57 | A | N | W |
| 38.63 | 2.71 | 43.1 | 7.17 | N | N | N |
| 30 | 4.6 | 11.1 | 15.40 | A | A | A |
| 30.8 | 3.8 | 14.1 | 12.43 | A | A | A |
| 21.83 | 8.02 | -19.1 | 36.77 | A | N | W |
| 31.21 | 8.98 | 15.6 | 28.81 | A | N | W |
| 28 | 3.37 | 3.7 | 12.13 | A | A | A |
| 62.59 | 11.3 | 131.8 | 18.11 | N | N | N |
| 39 | 5.9 | 44.4 | 15.20 | N | N | N |
| 25 | 3 | -7.4 | 12.09 | A | A | A |
| 50.9 | 6.2 | 88.5 | 12.27 | N | N | N |
| 45.3 | 5.25 | 67.8 | 11.68 | N | N | N |
| 46.22 | 4.6 | 71.2 | 10.06 | N | N | N |
| 32.4 | 4.4 | 20.0 | 13.66 | A | A | A |
| 26.5 | 3.7 | -1.9 | 14.04 | A | A | A |
| 27.97 | 5.95761 | 3.6 | 21.35 | A | N | W |
| 99.991 | 2.446 | 270.3 | 2.86 | N | N | N |
| 25 | 1.1 | -7.4 | 4.64 | A | A | A |

Labcode

52

78

90

110

116

150

220

226

276

315

326

341

348

384

395

403

414

421

452

491

537

Th-234 determination of activity concentration in Sample 3

Table 58. Target values

| Parameter | Value |
| --- | --- |
| Activity concentration [ Bq/kg ] | 27.0 |
| Uncertainty [ Bq/kg ] | 0.4 |
| Maximum Acceptable Relative Bias [%] | 20 |

Accepted 12%	Not Accepted 9%	Warning 2%	Not Reported 77%

35

Figure 58. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

20

25

155

64

294

57

19

71

533

159

288

347

393

463

272

379

25

41

276

348

537

226

15

421

66

72

425

45

60

499

161

151

202

128

322

426

2

371

295

21

170

315

247

78

373

484

110

184

204

452

63

150

319

131

97

407

242

156

178

364

55

501

330

278

414

358

446

256

291

49

532

79

230

299

116

402

430

480

58

341

382

286

90

408

203

317

68

124

462

326

400

403

384

555

495

Laboratory code

Figure 59. Reported results and evaluation parameters for Th-234. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 59. Evaluation results for Th-234 in Sample 3, Rep. Value and Rep. Unc. in Bq/kg

Labcode

Reported value

[Bq/kg ]

Reported uncertainty

(k = 1) [Bq/kg ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

Activity concentration  [ Bq/kg ]

30

2	27.7	2.9	2.6	10.57	A	A	A

15	25.5	2.4	-5.6	9.53	A	A	A

19	20.1	3.8	-25.6	18.96	N	A	N

| Reported value [Bq/kg ] 28 | Reported uncertainty (k = 1) [Bq/kg ] 3 | Relative bias [%] 3.7 | P-Test [%] 10.82 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 24.4 | 1.7 | -9.6 | 7.12 | A | A | A |
| 24.65 | 2.74 | -8.7 | 11.21 | A | A | A |
| 25.86 | 3.74 | -4.2 | 14.54 | A | A | A |
| 35 | 5 | 29.6 | 14.36 | N | A | N |
| 31.38 | 6.36 | 16.2 | 20.32 | A | N | W |
| 20 | 2 | -25.9 | 10.11 | N | A | N |
| 38.7 | 5.5 | 43.3 | 14.29 | N | N | N |
| 26 | 3.2 | -3.7 | 12.40 | A | A | A |
| 29.2 | 1.65 | 8.2 | 5.84 | A | A | A |
| 17.05 | 2.3 | -36.9 | 13.57 | N | N | N |
| 25.67 | 3.358 | -4.9 | 13.17 | A | A | A |
| 41.8 | 6.76 | 54.8 | 16.24 | N | N | N |
| 21.04 | 1.05 | -22.1 | 5.21 | N | N | N |
| 25.8 | 1.15 | -4.4 | 4.70 | A | A | A |
| 28.1 | 3.84 | 4.1 | 13.75 | A | A | A |
| 35.8 | 2.4 | 32.6 | 6.87 | N | N | N |
| 39.9 | 3.6 | 47.8 | 9.14 | N | N | N |
| 29.9 | 3.6 | 10.7 | 12.13 | A | A | A |
| 28.5 | 2.1 | 5.6 | 7.52 | A | A | A |
| 37.14 | 2.6 | 37.6 | 7.16 | N | N | N |
| 43 | 5 | 59.3 | 11.72 | N | N | N |
| 27.18 | 1.73 | 0.7 | 6.54 | A | A | A |
| 29.51942 | 8.355257 | 9.3 | 28.34 | A | N | W |
| 29.3 | 3.2 | 8.5 | 11.02 | A | A | A |
| 26.72 | 2.3 | -1.0 | 8.73 | A | A | A |
| 16.4 | 2.9 | -39.3 | 17.74 | N | A | N |
| 30.1 | 3.5 | 11.5 | 11.72 | A | A | A |
| 22.4 | 1.83 | -17.0 | 8.30 | A | A | A |
| 26.3 | 3.3 | -2.6 | 12.63 | A | A | A |
| 28 | 3 | 3.7 | 10.82 | A | A | A |
| 30.7 | 3.1 | 13.7 | 10.21 | A | A | A |
| 28.8 | 4.3 | 6.7 | 15.00 | A | A | A |
| 27.13 | 1.31 | 0.5 | 5.05 | A | A | A |
| 40.6 | 14.5 | 50.4 | 35.74 | N | N | N |
| 29 | 5 | 7.4 | 17.30 | A | A | A |
| 25.35 | 3.55 | -6.1 | 14.08 | A | A | A |
| 36.9 | 3.8 | 36.7 | 10.40 | N | N | N |
| 29.98 | 1.03 | 11.0 | 3.74 | A | N | W |
| 28.042 | 3.705 | 3.9 | 13.30 | A | A | A |
| 34.25 | 5.59 | 26.9 | 16.39 | N | A | N |
| 24.2 | 1.7 | -10.4 | 7.18 | A | A | A |
| 24.95 | 2.13 | -7.6 | 8.66 | A | A | A |
| 32.3 | 8.2 | 19.6 | 25.43 | A | N | W |
| 39.6 | 17.8 | 46.7 | 44.97 | N | N | N |
| 22.6 | 2.7 | -16.3 | 12.04 | A | A | A |
| 34.88 | 6.76 | 29.2 | 19.44 | N | A | N |
| 19.4 | 2 | -28.1 | 10.42 | N | N | N |
| 27.81 | 2.71 | 3.0 | 9.86 | A | A | A |
| 37 | 9 | 37.0 | 24.37 | N | N | N |
| 28 | 3.37 | 3.7 | 12.13 | A | A | A |
| 41.4 | 2.9 | 53.3 | 7.16 | N | N | N |
| 29.4 | 1.7 | 8.9 | 5.97 | A | A | A |
| 27.2 | 6.8 | 0.7 | 25.04 | A | N | W |
| 43.66 | 4 | 61.7 | 9.28 | N | N | N |
| 31.95 | 0.95 | 18.3 | 3.32 | A | N | W |
| 39 | 5.9 | 44.4 | 15.20 | N | N | N |
| 23.98 | 1.6 | -11.2 | 6.83 | A | A | A |
| 25 | 3 | -7.4 | 12.09 | A | A | A |
| 33 | 3.3 | 22.2 | 10.11 | N | A | N |

Labcode

21

25

41

45

49

55

57

58

60

63

64

66

68

71

72

78

79

90

97

110

116

124

128

131

150

151

155

156

159

161

170

178

184

202

203

204

226

230

242

247

256

272

276

278

286

288

291

294

295

299

315

317

319

322

326

330

341

347

348

358

| Reported value [Bq/kg ] 30.7 | Reported uncertainty (k = 1) [Bq/kg ] 4.6 | Relative bias [%] 13.7 | P-Test [%] 15.06 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 27.79 | 4.48 | 2.9 | 16.19 | A | A | A |
| 28.2 | 2.5 | 4.4 | 8.99 | A | A | A |
| 24.2 | 3.7 | -10.4 | 15.36 | A | A | A |
| 39.23 | 6.65 | 45.3 | 17.02 | N | N | N |
| 50.9 | 6.2 | 88.5 | 12.27 | N | N | N |
| 24 | 1.5 | -11.1 | 6.42 | A | A | A |
| 44 | 12 | 63.0 | 27.31 | N | N | N |
| 37.6 | 3.3 | 39.3 | 8.90 | N | N | N |
| 46.22 | 4.6 | 71.2 | 10.06 | N | N | N |
| 29.9 | 2.2 | 10.7 | 7.51 | A | A | A |
| 40 | 2 | 48.1 | 5.21 | N | N | N |
| 32.5 | 3.4 | 20.4 | 10.57 | N | A | N |
| 25.6 | 2.7 | -5.2 | 10.65 | A | A | A |
| 25.8 | 2.9 | -4.4 | 11.34 | A | A | A |
| 27.25 | 9.1 | 0.9 | 33.43 | A | N | W |
| 37.8 | 11.8 | 40.0 | 31.25 | N | N | N |
| 33.7 | 7.8 | 24.8 | 23.19 | N | N | N |
| 29.18 | 5.77764 | 8.1 | 19.86 | A | A | A |
| 43 | 6 | 59.3 | 14.03 | N | N | N |
| 24 | 3 | -11.1 | 12.59 | A | A | A |
| 38 | 10 | 40.7 | 26.36 | N | N | N |
| 28.4 | 4.4 | 5.2 | 15.56 | A | A | A |
| 153.69 | 1.37 | 469.2 | 1.73 | N | N | N |
| 26 | 3 | -3.7 | 11.63 | A | A | A |
| 31.8 | 5 | 17.8 | 15.79 | A | A | A |
| 35.13 | 6.84 | 30.1 | 19.53 | N | A | N |
| 21.9 | 2.1 | -18.9 | 9.70 | A | A | A |
| 25 | 1.1 | -7.4 | 4.64 | A | A | A |
| 51.1 | 8.7 | 89.3 | 17.09 | N | N | N |

Labcode

364

371

373

379

382

384

393

400

402

403

407

408

414

421

425

426

430

446

452

462

463

480

484

495

499

501

532

533

537

555

U-235 determination of activity concentration in Sample 3

Table 60. Target values

| Parameter | Value |
| --- | --- |
| Activity concentration [ Bq/kg ] | 1.24 |
| Uncertainty [ Bq/kg ] | 0.02 |
| Maximum Acceptable Relative Bias [%] | 20 |

Accepted 4%	Not Accepted 15%	Warning 1%	Not Reported 80%

3

Figure 60. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

−1

0

308

475

261

393

110

347

508

184

391

537

2

272

97

421

25

227

433

357

191

299

414

131

200

328

313

159

398

267

509

116

303

64

175

258

341

319

41

529

276

315

348

410

462

287

452

440

64

344

128

403

87

19

119

260

384

127

203

397

150

411

13

326

147

333

218

124

346

249

197

143

34

284

78

55

491

412

254

Laboratory code

Figure 61. Reported results and evaluation parameters for U-235. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 61. Evaluation results for U-235 in Sample 3, Rep. Value and Rep. Unc. in Bq/kg

Labcode

Reported value

[Bq/kg ]

Reported uncertainty

(k = 1) [Bq/kg ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

Activity concentration  [ Bq/kg ]

1

2

2	1.27	0.13	2.4	10.36	A	A	A

13	2.9	1.3	133.9	44.86	N	N	N

19	2.28	0.35	83.9	15.44	N	N	N

| Reported value [Bq/kg ] 1.37 | Reported uncertainty (k = 1) [Bq/kg ] 0.19 | Relative bias [%] 10.5 | P-Test [%] 13.96 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 4.8 | 0.5 | 287.1 | 10.54 | N | N | N |
| 1.9 | 0.725 | 53.2 | 38.19 | N | N | N |
| 5.21 | 0.52 | 320.2 | 10.11 | N | N | N |
| 2.1 | 0.9 | 69.3 | 42.89 | N | N | N |
| 1.8 | 0.07 | 45.2 | 4.21 | N | N | N |
| 4.85 | 0.63 | 291.1 | 13.09 | N | N | N |
| 1.3 | 0.2 | 4.8 | 15.47 | A | A | A |
| 1.11 | 0.21 | -10.5 | 18.99 | A | A | A |
| 1.72 | 0.16 | 38.7 | 9.44 | N | N | N |
| 2.31 | 0.32 | 86.3 | 13.95 | N | N | N |
| 4 | 1 | 222.6 | 25.05 | N | N | N |
| 2.46 | 0.47 | 98.4 | 19.17 | N | N | N |
| 2.12 | 0.32 | 71.0 | 15.18 | N | N | N |
| 1.512841 | 0.07514 | 22.0 | 5.22 | N | N | N |
| 4.561303 | 0.408749 | 267.9 | 9.11 | N | N | N |
| 3.1735 | 0.3527 | 155.9 | 11.23 | N | N | N |
| 2.71 | 0.34 | 118.6 | 12.65 | N | N | N |
| 1.59 | 0.0856 | 28.2 | 5.62 | N | N | N |
| 1.8 | 0.29 | 45.2 | 16.19 | N | N | N |
| 1.2 | 0.18 | -3.2 | 15.09 | A | A | A |
| 1.47 | 0.43 | 18.6 | 29.30 | A | N | W |
| 4.43 | 1.11 | 257.3 | 25.11 | N | N | N |
| 1.55 | 0.3 | 25.0 | 19.42 | N | A | N |
| 2.5 | 0.6 | 101.6 | 24.05 | N | N | N |
| 3.4162 | 0.55982 | 175.5 | 16.47 | N | N | N |
| 1.4 | 0.2 | 12.9 | 14.38 | A | A | A |
| 4.39 | 0.06988 | 254.0 | 2.27 | N | N | N |
| 6.21 | 2.06 | 400.8 | 33.21 | N | N | N |
| 1.8 | 0.1 | 45.2 | 5.78 | N | N | N |
| 2.4 | 0.25 | 93.6 | 10.54 | N | N | N |
| 1 | 0.3 | -19.4 | 30.04 | A | N | W |
| 1.638 | 0.663 | 32.1 | 40.51 | N | N | N |
| 1.28 | 0.11 | 3.2 | 8.74 | A | A | A |
| 1.95 | 0.33 | 57.3 | 17.00 | N | N | N |
| 4.805 | 0.59 | 287.5 | 12.38 | N | N | N |
| 2.05 | 0.14 | 65.3 | 7.02 | N | N | N |
| 1.5 | 0.2 | 21.0 | 13.43 | N | A | N |
| 1.74 | 0.31 | 40.3 | 17.89 | N | A | N |
| 0.36 | 0.07 | -71.0 | 19.51 | N | N | N |
| 1.57 | 0.194 | 26.6 | 12.46 | N | A | N |
| 1.98 | 0.27 | 59.7 | 13.73 | N | N | N |
| 1.88 | 0.11 | 51.6 | 6.07 | N | N | N |
| 3.1 | 1.1 | 150.0 | 35.52 | N | N | N |
| 1.55 | 0.41 | 25.0 | 26.50 | N | N | N |
| 3.31 | 0.84 | 166.9 | 25.43 | N | N | N |
| 1.8 | 0.13 | 45.2 | 7.40 | N | N | N |
| 2.1 | 0.4 | 69.3 | 19.12 | N | N | N |
| 4.29 | 0.6 | 246.0 | 14.08 | N | N | N |
| 1.11 | 0.1 | -10.5 | 9.15 | A | A | A |
| 1.98 | 0.17 | 59.7 | 8.74 | N | N | N |
| 1.44 | 0.1 | 16.1 | 7.13 | A | A | A |
| 2.4 | 0.4 | 93.6 | 16.74 | N | N | N |
| 1.2 | 0.22 | -3.2 | 18.40 | A | A | A |
| 1.1 | 0.1 | -11.3 | 9.23 | A | A | A |
| 2.64 | 0.03 | 112.9 | 1.97 | N | N | N |
| 1.6 | 0.45 | 29.0 | 28.17 | N | N | N |
| 2.17 | 0.44 | 75.0 | 20.34 | N | N | N |
| 2 | 0.4 | 61.3 | 20.06 | N | N | N |
| 2.8 | 0.9 | 125.8 | 32.18 | N | N | N |

Labcode

25

34

41

55

64

64

78

97

110

116

119

124

127

128

131

143

147

150

159

175

184

191

197

200

203

218

227

249

254

258

260

261

267

272

276

284

287

299

303

308

313

315

319

326

328

333

341

344

346

347

348

357

384

391

393

397

398

403

410

411

| Reported value [Bq/kg ] 5.6 | Reported uncertainty (k = 1) [Bq/kg ] 1.39 | Relative bias [%] 351.6 | P-Test [%] 24.87 | Trueness evaluation N | Precision evaluation N | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 1.5 | 0.23 | 21.0 | 15.42 | N | A | N |
| 1.36 | 0.14 | 9.7 | 10.42 | A | A | A |
| 1.42 | 0.76 | 14.5 | 53.55 | A | N | W |
| 2.082 | 0.454 | 67.9 | 21.87 | N | N | N |
| 2.066 | 0.427662 | 66.6 | 20.76 | N | N | N |
| 2 | 1 | 61.3 | 50.03 | N | N | N |
| 0.87 | 0.08 | -29.8 | 9.34 | N | N | N |
| 5.256 | 1.115 | 323.9 | 21.28 | N | N | N |
| 1.69 | 0.11 | 36.3 | 6.71 | N | N | N |
| 1.92 | 0.17 | 54.8 | 9.00 | N | N | N |
| 1.2 | 0.1 | -3.2 | 8.49 | A | A | A |

Labcode

412

414

421

433

440

452

462

475

491

509

529

537

U-238 determination of activity concentration in Sample 3

Table 62. Target values

| Parameter | Value |
| --- | --- |
| Activity concentration [ Bq/kg ] | 27.0 |
| Uncertainty [ Bq/kg ] | 0.4 |
| Maximum Acceptable Relative Bias [%] | 20 |

Accepted 11%	Not Accepted 10%	Warning 1%	Not Reported 78%

60

80

Figure 62. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

−20

0

408

377

308

293

475

57

303

261

553

232

525

405

159

31

508

463

363

523

348

537

391

127

15

72

163

290

312

125

235

370

270

487

128

539

122

356

2

192

411

215

152

184

119

465

319

433

227

158

509

407

178

53

207

357

55

191

330

414

328

398

200

267

529

38

79

116

78

193

258

341

506

238

382

206

344

100

410

443

214

403

87

315

384

323

249

35

52

491

397

Laboratory code

Figure 63. Reported results and evaluation parameters for U-238. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 63. Evaluation results for U-238 in Sample 3, Rep. Value and Rep. Unc. in Bq/kg

Labcode

Reported value

[Bq/kg ]

Reported uncertainty

(k = 1) [Bq/kg ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

Activity concentration  [ Bq/kg ]

20

40

2	27.7	2.9	2.6	10.57	A	A	A

15	25.5	2.4	-5.6	9.53	A	A	A

31	23.3	2.55	-13.7	11.04	A	A	A

| Reported value [Bq/kg ] 59 | Reported uncertainty (k = 1) [Bq/kg ] 15.1 | Relative bias [%] 118.5 | P-Test [%] 25.64 | Trueness evaluation N | Precision evaluation N | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 35.4 | 6.91 | 31.1 | 19.58 | N | A | N |
| 67.6 | 3.03 | 150.4 | 4.72 | N | N | N |
| 30.83 | 3.17 | 14.2 | 10.39 | A | A | A |
| 31.38 | 6.36 | 16.2 | 20.32 | A | N | W |
| 20 | 2 | -25.9 | 10.11 | N | A | N |
| 25.8 | 4.9 | -4.4 | 19.05 | A | A | A |
| 37.2 | 8.52 | 37.8 | 22.95 | N | N | N |
| 35.8 | 2.4 | 32.6 | 6.87 | N | N | N |
| 40.8 | 6.8 | 51.1 | 16.73 | N | N | N |
| 37.14 | 2.6 | 37.6 | 7.16 | N | N | N |
| 29 | 3.35 | 7.4 | 11.65 | A | A | A |
| 27.6 | 3.7 | 2.2 | 13.49 | A | A | A |
| 26.6 | 2.34 | -1.5 | 8.92 | A | A | A |
| 25.4 | 2 | -5.9 | 8.01 | A | A | A |
| 27.18 | 1.73 | 0.7 | 6.54 | A | A | A |
| 28.7 | 2.8 | 6.3 | 9.87 | A | A | A |
| 30 | 16 | 11.1 | 53.35 | A | N | W |
| 22.4 | 1.83 | -17.0 | 8.30 | A | A | A |
| 26 | 3 | -3.7 | 11.63 | A | A | A |
| 30.7 | 3.1 | 13.7 | 10.21 | A | A | A |
| 28.8 | 4.3 | 6.7 | 15.00 | A | A | A |
| 31.9 | 9.3 | 18.1 | 29.19 | A | N | W |
| 28 | 5 | 3.7 | 17.92 | A | A | A |
| 37.92 | 2.32 | 40.4 | 6.29 | N | N | N |
| 34.4 | 3 | 27.4 | 8.85 | N | N | N |
| 40 | 3 | 48.1 | 7.64 | N | N | N |
| 31.3 | 2.8 | 15.9 | 9.07 | A | A | A |
| 43 | 3 | 59.3 | 7.13 | N | N | N |
| 28.4 | 3.4 | 5.2 | 12.06 | A | A | A |
| 29.9 | 2.3 | 10.7 | 7.83 | A | A | A |
| 22.2 | 3.9 | -17.8 | 17.63 | A | A | A |
| 26.6 | 2.4 | -1.5 | 9.14 | A | A | A |
| 39.2 | 5.98 | 45.2 | 15.33 | N | N | N |
| 53.8 | 0.858891 | 99.3 | 2.18 | N | N | N |
| 38.1 | 2.7 | 41.1 | 7.24 | N | N | N |
| 21.4 | 4.2 | -20.7 | 19.68 | N | A | N |
| 34.62 | 1.563 | 28.2 | 4.75 | N | N | N |
| 27 | 3 | 0.0 | 11.21 | A | A | A |
| 26 | 3 | -3.7 | 11.63 | A | A | A |
| 18.4 | 2.8 | -31.9 | 15.29 | N | A | N |
| 20.6 | 1.3 | -23.7 | 6.48 | N | N | N |
| 11.2 | 0.4 | -58.5 | 3.87 | N | N | N |
| 26.3 | 3.1 | -2.6 | 11.88 | A | A | A |
| 49.5 | 6.75 | 83.3 | 13.72 | N | N | N |
| 29.4 | 1.7 | 8.9 | 5.97 | A | A | A |
| 53.5 | 2.6 | 98.2 | 5.08 | N | N | N |
| 32.7 | 4 | 21.1 | 12.32 | N | A | N |
| 31.95 | 0.95 | 18.3 | 3.32 | A | N | W |
| 39 | 5.9 | 44.4 | 15.20 | N | N | N |
| 40.2 | 3.1 | 48.9 | 7.85 | N | N | N |
| 25 | 3 | -7.4 | 12.09 | A | A | A |
| 27.6 | 4.4 | 2.2 | 16.01 | A | A | A |
| 31.3 | 3.9 | 15.9 | 12.55 | A | A | A |
| 24.7 | 4.3 | -8.5 | 17.47 | A | A | A |
| 26.9 | 2.2 | -0.4 | 8.31 | A | A | A |
| 2.018346 | 0.245445 | -92.5 | 12.25 | N | N | N |
| 39.23 | 6.65 | 45.3 | 17.02 | N | N | N |
| 50.9 | 6.2 | 88.5 | 12.27 | N | N | N |
| 25.11 | 1.7 | -7.0 | 6.93 | A | A | A |

Labcode

35

38

52

53

55

57

72

78

79

100

116

119

122

125

127

128

152

158

159

163

178

184

191

192

193

200

206

207

214

215

227

232

235

238

249

258

261

267

270

290

293

303

308

312

315

319

323

328

330

341

344

348

356

357

363

370

377

382

384

391

| Reported value [Bq/kg ] 194.25 | Reported uncertainty (k = 1) [Bq/kg ] 27.98 | Relative bias [%] 619.4 | P-Test [%] 14.48 | Trueness evaluation N | Precision evaluation N | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 33.9 | 4.5 | 25.6 | 13.36 | N | A | N |
| 46.22 | 4.6 | 71.2 | 10.06 | N | N | N |
| 22.3 | 1.2 | -17.4 | 5.58 | A | N | W |
| 30.5 | 2.2 | 13.0 | 7.36 | A | A | A |
| 1.84 | 0.11 | -93.2 | 6.16 | N | N | N |
| 41.3 | 2.7 | 53.0 | 6.70 | N | N | N |
| 28 | 2.3 | 3.7 | 8.35 | A | A | A |
| 32.5 | 2.7 | 20.4 | 8.44 | N | A | N |
| 29.48 | 4.25 | 9.2 | 14.49 | A | A | A |
| 42.4 | 2.9 | 57.0 | 7.00 | N | N | N |
| 24 | 3 | -11.1 | 12.59 | A | A | A |
| 29.2 | 2.4 | 8.2 | 8.35 | A | A | A |
| 18.9 | 0.8 | -30.0 | 4.48 | N | N | N |
| 27 | 3 | 0.0 | 11.21 | A | A | A |
| 106.292 | 1.149 | 293.7 | 1.83 | N | N | N |
| 39 | 3 | 44.4 | 7.83 | N | N | N |
| 30 | 5 | 11.1 | 16.73 | A | A | A |
| 24.9 | 2.2 | -7.8 | 8.96 | A | A | A |
| 22.2 | 3.1 | -17.8 | 14.04 | A | A | A |
| 34.8 | 3.6 | 28.9 | 10.45 | N | N | N |
| 25 | 1.1 | -7.4 | 4.64 | A | A | A |
| 27.44 | 3.32 | 1.6 | 12.19 | A | A | A |
| 21.6 | 2.1 | -20.0 | 9.83 | A | A | A |

Labcode

397

398

403

405

407

408

410

411

414

433

443

463

465

475

487

491

506

509

523

525

529

537

539

553

K-40 determination of activity concentration in Sample 4

Table 64. Target values

| Parameter | Value |
| --- | --- |
| Activity concentration [ Bq/kg ] | 2029 |
| Uncertainty [ Bq/kg ] | 69 |
| Maximum Acceptable Relative Bias [%] | 20 |

Accepted 66%	Not Accepted 8%	Warning 7%	Not Reported 19%

2500

Figure 64. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

1500

37

300

467

397

24

80

284

184

172

446

395

288

276

287

452

539

323

145

40

73

500

185

66

230

453

294

218

100

213

109

75

247

523

460

233

314

134

223

391

319

346

492

403

81

508

450

111

316

241

178

303

528

143

256

55

514

405

393

307

191

471

235

151

532

121

140

464

67

154

315

475

270

65

8

396

127

152

290

267

312

537

119

175

344

363

53

310

499

306

425

274

437

208

25

291

49

192

388

35

466

509

260

444

387

302

2

198

441

398

60

522

417

156

83

525

339

227

94

286

82

13

332

411

181

222

516

484

69

54

420

536

498

386

443

64

330

402

105

317

48

295

429

58

377

78

501

130

331

299

338

357

495

530

31

364

507

262

462

455

333

Laboratory code

Figure 65. Reported results and evaluation parameters for K-40. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 65. Evaluation results for K-40 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

Labcode

Reported value

[Bq/kg ]

Reported uncertainty

(k = 1) [Bq/kg ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

Activity concentration  [ Bq/kg ]

2000

2	2140	200	5.5	9.95	A	A	A

5	2133	140	5.1	7.39	A	A	A

8	2064	100	1.7	5.92	A	A	A

| Reported value [Bq/kg ] 2065 | Reported uncertainty (k = 1) [Bq/kg ] 248 | Relative bias [%] 1.8 | P-Test [%] 12.48 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 2190 | 149 | 7.9 | 7.61 | A | A | A |
| 2108.59 | 136.57 | 3.9 | 7.32 | A | A | A |
| 2194.5 | 83.8 | 8.2 | 5.11 | A | A | A |
| 1949 | 127 | -3.9 | 7.35 | A | A | A |
| 2102.36 | 87.94 | 3.6 | 5.39 | A | A | A |
| 2182 | 283 | 7.5 | 13.41 | A | A | A |
| 2329.09 | 32.04 | 14.8 | 3.67 | A | N | W |
| 2260 | 46 | 11.4 | 3.96 | A | N | W |
| 1500 | 62 | -26.1 | 5.35 | N | N | N |
| 2110 | 87 | 4.0 | 5.34 | A | A | A |
| 2136 | 185.2 | 5.3 | 9.31 | A | A | A |
| 2460 | 85.5 | 21.2 | 4.86 | N | N | N |
| 1814 | 70 | -10.6 | 5.14 | A | A | A |
| 2420 | 255 | 19.3 | 11.07 | A | A | A |
| 2127 | 178 | 4.8 | 9.03 | A | A | A |
| 16.97 | 1.43 | -99.2 | 9.09 | N | N | N |
| 1907.22 | 131.54 | -6.0 | 7.69 | A | A | A |
| 1849 | 133 | -8.9 | 7.96 | A | A | A |
| 2034.88 | 93.1 | 0.3 | 5.70 | A | A | A |
| 1954 | 64 | -3.7 | 4.72 | A | A | A |
| 1916 | 93 | -5.6 | 5.93 | A | A | A |
| 2302 | 82 | 13.4 | 4.92 | A | N | W |
| 2120 | 210 | 4.5 | 10.47 | A | A | A |
| 2205.7 | 44.6 | 8.7 | 3.96 | A | A | A |
| 2150 | 92.6 | 6.0 | 5.49 | A | A | A |
| 2083 | 167 | 2.7 | 8.71 | A | A | A |
| 2225 | 112 | 9.7 | 6.07 | A | A | A |
| 2011.4 | 84.17 | -0.9 | 5.39 | A | A | A |
| 2019.9 | 202 | -0.4 | 10.56 | A | A | A |
| 2328 | 118 | 14.7 | 6.10 | A | A | A |
| 2150 | 210 | 6.0 | 10.34 | A | A | A |
| 1920 | 150 | -5.4 | 8.52 | A | A | A |
| 2060.04 | 38.85 | 1.5 | 3.89 | A | A | A |
| 2254.8 | 197.9 | 11.1 | 9.41 | A | A | A |
| 2061.5 | 104.7 | 1.6 | 6.11 | A | A | A |
| 1875 | 98.9 | -7.6 | 6.28 | A | A | A |
| 2049 | 254 | 1.0 | 12.85 | A | A | A |
| 2220 | 97 | 9.4 | 5.54 | A | A | A |
| 1950 | 64 | -3.9 | 4.73 | A | A | A |
| 2191.5 | 109.6 | 8.0 | 6.05 | A | A | A |
| 1826 | 80 | -10.0 | 5.55 | A | A | A |
| 1860 | 78 | -8.3 | 5.40 | A | A | A |
| 1921.9 | 27.4 | -5.3 | 3.69 | A | A | A |
| 2347.9 | 18.73 | 15.7 | 3.49 | A | N | W |
| 2023 | 22 | -0.3 | 3.57 | A | A | A |
| 1545.1 | 163.95 | -23.9 | 11.14 | N | A | N |
| 1976 | 206 | -2.6 | 10.97 | A | A | A |
| 2190.82 | 20.59 | 8.0 | 3.53 | A | A | A |
| 2162 | 313.5 | 6.6 | 14.89 | A | A | A |
| 1600 | 131 | -21.1 | 8.87 | N | A | N |
| 2119 | 139 | 4.4 | 7.39 | A | A | A |
| 2185.1 | 143.17 | 7.7 | 7.38 | A | A | A |
| 2732.2 | 213.59 | 34.7 | 8.53 | N | N | N |
| 1998 | 86 | -1.5 | 5.49 | A | A | A |
| 1915 | 141 | -5.6 | 8.11 | A | A | A |
| 1921 | 187 | -5.3 | 10.31 | A | A | A |
| 2128.4 | 43.4 | 4.9 | 3.97 | A | A | A |
| 2274.1 | 47.5 | 12.1 | 3.99 | A | N | W |
| 2050 | 115 | 1.0 | 6.56 | A | A | A |

Labcode

9

11

12

13

15

17

19

20

21

24

25

28

31

32

34

35

37

38

40

41

42

45

48

49

50

52

53

54

55

57

58

60

62

63

64

65

66

67

69

70

71

72

73

75

78

79

80

81

82

83

84

90

94

95

97

100

102

104

105

107

| Reported value [Bq/kg ] 1920 | Reported uncertainty (k = 1) [Bq/kg ] 10 | Relative bias [%] -5.4 | P-Test [%] 3.44 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 2027 | 115 | -0.1 | 6.61 | A | A | A |
| 1983 | 120 | -2.3 | 6.94 | A | A | A |
| 2090 | 209 | 3.0 | 10.56 | A | A | A |
| 1939 | 130 | -4.4 | 7.52 | A | A | A |
| 1861.3 | 112.1 | -8.3 | 6.92 | A | A | A |
| 2080 | 170 | 2.5 | 8.85 | A | A | A |
| 2040 | 151 | 0.5 | 8.15 | A | A | A |
| 2075 | 135 | 2.3 | 7.34 | A | A | A |
| 2256 | 134 | 11.2 | 6.84 | A | A | A |
| 2068 | 180 | 1.9 | 9.34 | A | A | A |
| 2168.23 | 81.85 | 6.9 | 5.08 | A | A | A |
| 2383 | 91 | 17.4 | 5.11 | A | N | W |
| 1946.496 | 45.96987 | -4.1 | 4.14 | A | A | A |
| 2100 | 107 | 3.5 | 6.13 | A | A | A |
| 2449 | 218 | 20.7 | 9.53 | N | A | N |
| 1949.3 | 9.9 | -3.9 | 3.44 | A | A | A |
| 2040 | 84.6 | 0.5 | 5.36 | A | A | A |
| 2040 | 130 | 0.5 | 7.22 | A | A | A |
| 2002.071 | 83.146214 | -1.3 | 5.37 | A | A | A |
| 1842.6 | 276.39 | -9.2 | 15.38 | A | A | A |
| 2539.378 | 102.28 | 25.1 | 5.27 | N | N | N |
| 1638 | 115 | -19.3 | 7.80 | A | A | A |
| 2035 | 163 | 0.3 | 8.70 | A | A | A |
| 2070 | 110 | 2.0 | 6.31 | A | A | A |
| 2050 | 110 | 1.0 | 6.35 | A | A | A |
| 2037 | 114 | 0.4 | 6.55 | A | A | A |
| 2156 | 47 | 6.3 | 4.04 | A | A | A |
| 2078 | 279 | 2.4 | 13.85 | A | A | A |
| 2210 | 95 | 8.9 | 5.48 | A | A | A |
| 2080 | 85 | 2.5 | 5.32 | A | A | A |
| 2243.7 | 93.6 | 10.6 | 5.38 | A | A | A |
| 1980 | 126 | -2.4 | 7.22 | A | A | A |
| 2120 | 210 | 4.5 | 10.47 | A | A | A |
| 2000 | 300 | -1.4 | 15.38 | A | A | A |
| 1660 | 120 | -18.2 | 7.99 | A | A | A |
| 2080 | 93 | 2.5 | 5.62 | A | A | A |
| 2000 | 160 | -1.4 | 8.69 | A | A | A |
| 1679.52 | 24.97 | -17.2 | 3.71 | A | N | W |
| 1876.48 | 80.21 | -7.5 | 5.46 | A | A | A |
| 2206.6 | 98.08 | 8.8 | 5.60 | A | A | A |
| 2051 | 100 | 1.1 | 5.94 | A | A | A |
| 1650 | 230 | -18.7 | 14.35 | A | A | A |
| 1870.94 | 19.01 | -7.8 | 3.55 | A | A | A |
| 2061.5 | 104.7 | 1.6 | 6.11 | A | A | A |
| 1653 | 73 | -18.5 | 5.57 | A | N | W |
| 2030 | 130 | 0.0 | 7.25 | A | A | A |
| 2120 | 212 | 4.5 | 10.56 | A | A | A |
| 1852 | 21 | -8.7 | 3.58 | A | A | A |
| 2029.7 | 284.2 | 0.0 | 14.41 | A | A | A |
| 2140 | 180 | 5.5 | 9.07 | A | A | A |
| 2320 | 140 | 14.3 | 6.93 | A | A | A |
| 215.23 | 3.79 | -89.4 | 3.83 | N | N | N |
| 2459 | 229 | 21.2 | 9.91 | N | A | N |
| 2000 | 50 | -1.4 | 4.22 | A | A | A |
| 2360 | 130 | 16.3 | 6.47 | A | A | A |
| 1881 | 205 | -7.3 | 11.42 | A | A | A |
| 2410 | 160 | 18.8 | 7.46 | A | A | A |
| 2104 | 93.4 | 3.7 | 5.59 | A | A | A |
| 1847.42 | 60.4 | -8.9 | 4.72 | A | A | A |

Labcode

109

110

111

113

114

116

119

121

122

124

127

128

130

131

132

133

134

139

140

143

145

147

150

151

152

154

155

156

158

159

161

165

166

168

170

172

175

178

179

180

181

183

184

185

187

189

191

192

194

197

198

200

202

203

204

205

206

207

208

211

| Reported value [Bq/kg ] 1918 | Reported uncertainty (k = 1) [Bq/kg ] 89.8 | Relative bias [%] -5.5 | P-Test [%] 5.79 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 1964 | 32 | -3.2 | 3.77 | A | A | A |
| 2070 | 135 | 2.0 | 7.36 | A | A | A |
| 1958.1 | 313.3 | -3.5 | 16.36 | A | A | A |
| 1900.4176 | 78.95298 | -6.3 | 5.37 | A | A | A |
| 2040.3 | 122.1 | 0.6 | 6.88 | A | A | A |
| 2151 | 108 | 6.0 | 6.06 | A | A | A |
| 2208 | 139 | 8.8 | 7.16 | A | A | A |
| 1950 | 190 | -3.9 | 10.32 | A | A | A |
| 2001.7 | 180.2 | -1.4 | 9.62 | A | A | A |
| 1542.61 | 64.78 | -24.0 | 5.40 | N | N | N |
| 2180 | 70 | 7.4 | 4.68 | A | A | A |
| 1486.7 | 69.2 | -26.7 | 5.76 | N | N | N |
| 1880 | 110 | -7.3 | 6.77 | A | A | A |
| 2054 | 120 | 1.2 | 6.76 | A | A | A |
| 1945 | 110 | -4.1 | 6.60 | A | A | A |
| 2034 | 200 | 0.2 | 10.40 | A | A | A |
| 1943 | 82.6 | -4.2 | 5.44 | A | A | A |
| 1999 | 89 | -1.5 | 5.60 | A | A | A |
| 2013.99 | 57.7 | -0.7 | 4.45 | A | A | A |
| 1929.773 | 152.031 | -4.9 | 8.58 | A | A | A |
| 2160.504133 | 68.920082 | 6.5 | 4.66 | A | A | A |
| 2127.9 | 31.5 | 4.9 | 3.71 | A | A | A |
| 1933 | 118 | -4.7 | 6.99 | A | A | A |
| 2004.56 | 203.36 | -1.2 | 10.70 | A | A | A |
| 1668.94 | 43.71 | -17.8 | 4.29 | A | N | W |
| 2265 | 103 | 11.6 | 5.68 | A | A | A |
| 2130 | 110 | 5.0 | 6.18 | A | A | A |
| 2126 | 39 | 4.8 | 3.86 | A | A | A |
| 2658.3 | 85.9 | 31.0 | 4.69 | N | N | N |
| 2253 | 270 | 11.0 | 12.46 | A | A | A |
| 2072 | 249 | 2.1 | 12.49 | A | A | A |
| 1900 | 76 | -6.4 | 5.25 | A | A | A |
| 2056 | 120 | 1.3 | 6.76 | A | A | A |
| 2031.7 | 59.3 | 0.1 | 4.48 | A | A | A |
| 1926 | 120 | -5.1 | 7.10 | A | A | A |
| 2280 | 100 | 12.4 | 5.55 | A | A | A |
| 2100 | 100.674384 | 3.5 | 5.88 | A | A | A |
| 1790.07 | 125.98 | -11.8 | 7.82 | A | A | A |
| 2150 | 54 | 6.0 | 4.23 | A | A | A |
| 2084 | 45.2 | 2.7 | 4.03 | A | A | A |
| 19.07 | 0.64 | -99.1 | 4.78 | N | N | N |
| 2120 | 59 | 4.5 | 4.39 | A | A | A |
| 2207 | 51.4 | 8.8 | 4.12 | A | A | A |
| 1623 | 55 | -20.0 | 4.80 | N | N | N |
| 2190 | 96 | 7.9 | 5.55 | A | A | A |
| 1802 | 38 | -11.2 | 4.00 | A | N | W |
| 1741 | 181 | -14.2 | 10.94 | A | A | A |
| 2070 | 135 | 2.0 | 7.36 | A | A | A |
| 2118.3 | 211.83 | 4.4 | 10.56 | A | A | A |
| 2312 | 44 | 13.9 | 3.90 | A | N | W |
| 1952 | 187 | -3.8 | 10.17 | A | A | A |
| 1884 | 179 | -7.2 | 10.09 | A | A | A |
| 2319.2 | 92.68 | 14.3 | 5.25 | A | N | W |
| 2400 | 200 | 18.3 | 9.00 | A | A | A |
| 193 | 3 | -90.5 | 3.74 | N | N | N |
| 2136 | 171 | 5.3 | 8.70 | A | A | A |
| 2000 | 400 | -1.4 | 20.29 | A | N | W |
| 2230 | 72 | 9.9 | 4.69 | A | A | A |
| 2088 | 209 | 2.9 | 10.57 | A | A | A |

Labcode

213

214

215

217

218

219

220

222

223

224

226

227

228

230

232

233

235

238

241

242

247

249

252

253

256

257

258

260

261

262

265

267

268

270

271

272

273

274

276

277

278

280

281

283

284

286

287

288

290

291

292

293

294

295

299

300

302

303

304

306

| Reported value [Bq/kg ] 2027 | Reported uncertainty (k = 1) [Bq/kg ] 67 | Relative bias [%] -0.1 | P-Test [%] 4.74 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 2084 | 103 | 2.7 | 6.00 | A | A | A |
| 2130 | 63 | 5.0 | 4.51 | A | A | A |
| 2075 | 130 | 2.3 | 7.13 | A | A | A |
| 2050 | 135 | 1.0 | 7.41 | A | A | A |
| 1947 | 24 | -4.0 | 3.62 | A | A | A |
| 2050 | 281 | 1.0 | 14.12 | A | A | A |
| 1990 | 180 | -1.9 | 9.66 | A | A | A |
| 2286 | 50 | 12.7 | 4.04 | A | N | W |
| 1957 | 98 | -3.5 | 6.05 | A | A | A |
| 1957 | 107 | -3.5 | 6.44 | A | A | A |
| 1831 | 91 | -9.8 | 6.02 | A | A | A |
| 2080 | 146 | 2.5 | 7.80 | A | A | A |
| 2081.55 | 85 | 2.6 | 5.31 | A | A | A |
| 2070 | 94 | 2.0 | 5.67 | A | A | A |
| 2195 | 188.9 | 8.2 | 9.25 | A | A | A |
| 2259.47 | 124.78 | 11.4 | 6.49 | A | A | A |
| 2390 | 120 | 17.8 | 6.06 | A | N | W |
| 2200 | 180 | 8.4 | 8.86 | A | A | A |
| 3500.3 | 57.78 | 72.5 | 3.78 | N | N | N |
| 2419 | 215 | 19.2 | 9.52 | A | A | A |
| 2174 | 77 | 7.2 | 4.91 | A | A | A |
| 2041 | 57 | 0.6 | 4.40 | A | A | A |
| 2080.5 | 160.4 | 2.5 | 8.43 | A | A | A |
| 1957 | 83 | -3.5 | 5.44 | A | A | A |
| 2011 | 143 | -0.9 | 7.88 | A | A | A |
| 2100 | 140 | 3.5 | 7.48 | A | A | A |
| 2395.2 | 15.8 | 18.0 | 3.46 | A | N | W |
| 2000 | 70 | -1.4 | 4.88 | A | A | A |
| 1707.18 | 10.25 | -15.9 | 3.45 | A | N | W |
| 2127 | 70 | 4.8 | 4.73 | A | A | A |
| 2430 | 120 | 19.8 | 6.00 | A | N | W |
| 2250 | 22.5 | 10.9 | 3.54 | A | N | W |
| 2082 | 129 | 2.6 | 7.07 | A | A | A |
| 2521 | 107 | 24.2 | 5.44 | N | N | N |
| 1762 | 70.2 | -13.2 | 5.24 | A | A | A |
| 2200 | 100 | 8.4 | 5.68 | A | A | A |
| 2218 | 111 | 9.3 | 6.05 | A | A | A |
| 2241 | 90 | 10.4 | 5.26 | A | A | A |
| 2329.862 | 103.6831 | 14.8 | 5.60 | A | N | W |
| 1870 | 300 | -7.8 | 16.40 | A | A | A |
| 2336 | 104 | 15.1 | 5.60 | A | N | W |
| 1880 | 60 | -7.3 | 4.66 | A | A | A |
| 2086 | 208 | 2.8 | 10.54 | A | A | A |
| 2248.49 | 277.02 | 10.8 | 12.78 | A | A | A |
| 2134 | 120 | 5.2 | 6.57 | A | A | A |
| 2120 | 120 | 4.5 | 6.60 | A | A | A |
| 1953.72 | 61.55 | -3.7 | 4.64 | A | A | A |
| 2023 | 150 | -0.3 | 8.16 | A | A | A |
| 1704.5 | 170 | -16.0 | 10.54 | A | A | A |
| 2065 | 90.63 | 1.8 | 5.55 | A | A | A |
| 1239.7 | 14.15 | -38.9 | 3.59 | N | N | N |
| 2145 | 105 | 5.7 | 5.96 | A | A | A |
| 1800 | 84.8 | -11.3 | 5.81 | A | A | A |
| 2261 | 119 | 11.4 | 6.27 | A | A | A |
| 1972 | 79.21 | -2.8 | 5.26 | A | A | A |
| 2022 | 89 | -0.3 | 5.56 | A | A | A |
| 2068 | 93 | 1.9 | 5.64 | A | A | A |
| 1816 | 26 | -10.5 | 3.69 | A | N | W |
| 2110 | 140 | 4.0 | 7.46 | A | A | A |

Labcode

307

310

311

312

313

314

315

316

317

319

322

323

325

326

328

329

330

331

332

333

338

339

341

344

346

347

348

350

354

355

356

357

358

363

364

369

370

371

376

377

379

382

383

384

386

387

388

391

393

395

396

397

398

401

402

403

405

407

408

410

| Reported value [Bq/kg ] 2202 | Reported uncertainty (k = 1) [Bq/kg ] 72 | Relative bias [%] 8.5 | P-Test [%] 4.72 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 2590 | 86.8 | 27.6 | 4.77 | N | N | N |
| 2231 | 98 | 10.0 | 5.56 | A | A | A |
| 2151 | 108 | 6.0 | 6.06 | A | A | A |
| 2230 | 145 | 9.9 | 7.34 | A | A | A |
| 2092 | 140 | 3.1 | 7.51 | A | A | A |
| 2190 | 180.5 | 7.9 | 8.92 | A | A | A |
| 2137 | 250 | 5.3 | 12.18 | A | A | A |
| 2320 | 15.4 | 14.3 | 3.46 | A | N | W |
| 2389.36 | 61.1 | 17.8 | 4.25 | A | N | W |
| 1981.4 | 84.7 | -2.4 | 5.46 | A | A | A |
| 2380 | 100 | 17.3 | 5.41 | A | N | W |
| 2102 | 137 | 3.6 | 7.35 | A | A | A |
| 2140 | 91 | 5.5 | 5.44 | A | A | A |
| 3498.84 | 30.26 | 72.4 | 3.51 | N | N | N |
| 2144 | 111 | 5.7 | 6.19 | A | A | A |
| 2325.9255 | 153.196716 | 14.6 | 7.41 | A | A | A |
| 2250 | 150 | 10.9 | 7.48 | A | A | A |
| 2131 | 180 | 5.0 | 9.11 | A | A | A |
| 2142.3 | 99.9 | 5.6 | 5.77 | A | A | A |
| 1675 | 101 | -17.4 | 6.92 | A | A | A |
| 1870.94 | 19.01 | -7.8 | 3.55 | A | A | A |
| 1980 | 60 | -2.4 | 4.55 | A | A | A |
| 1814 | 90.7 | -10.6 | 6.05 | A | A | A |
| 1880 | 67.1 | -7.3 | 4.93 | A | A | A |
| 2299 | 123 | 13.3 | 6.34 | A | A | A |
| 3393 | 98 | 67.2 | 4.46 | N | N | N |
| 2165 | 198 | 6.7 | 9.76 | A | A | A |
| 1940 | 118 | -4.4 | 6.97 | A | A | A |
| 2954 | 127 | 45.6 | 5.48 | N | N | N |
| 2220 | 110 | 9.4 | 6.01 | A | A | A |
| 2040.85 | 189.16 | 0.6 | 9.87 | A | A | A |
| 2004 | 192 | -1.2 | 10.17 | A | A | A |
| 2127 | 47 | 4.8 | 4.06 | A | A | A |
| 511 | 2.7 | -74.8 | 3.44 | N | N | N |
| 2030 | 97 | 0.0 | 5.86 | A | A | A |
| 2030 | 52 | 0.0 | 4.26 | A | A | A |
| 2051 | 100 | 1.1 | 5.94 | A | A | A |
| 1984 | 40 | -2.2 | 3.95 | A | A | A |
| 2152 | 112 | 6.1 | 6.22 | A | A | A |
| 1972 | 120 | -2.8 | 6.97 | A | A | A |
| 2082 | 135 | 2.6 | 7.32 | A | A | A |
| 2065 | 140 | 1.8 | 7.58 | A | A | A |
| 2144 | 111 | 5.7 | 6.19 | A | A | A |
| 2213 | 98 | 9.1 | 5.58 | A | A | A |
| 2075 | 145 | 2.3 | 7.77 | A | A | A |
| 3197.681 | 83.767 | 57.6 | 4.29 | N | N | N |
| 1960 | 60 | -3.4 | 4.58 | A | A | A |
| 2434 | 46 | 20.0 | 3.89 | A | N | W |
| 2174 | 91.66 | 7.2 | 5.42 | A | A | A |
| 2242 | 266 | 10.5 | 12.34 | A | A | A |
| 2086 | 143 | 2.8 | 7.65 | A | A | A |
| 1865 | 114.7 | -8.1 | 7.03 | A | A | A |
| 2370 | 100 | 16.8 | 5.42 | A | N | W |
| 1978 | 127 | -2.5 | 7.27 | A | A | A |
| 2550.973 | 0.285 | 25.7 | 3.40 | N | N | N |
| 1979 | 82 | -2.5 | 5.36 | A | A | A |
| 2128 | 37 | 4.9 | 3.82 | A | A | A |
| 2015 | 140 | -0.7 | 7.74 | A | A | A |
| 2210 | 130 | 8.9 | 6.79 | A | A | A |

Labcode

411

412

414

417

420

425

426

428

429

433

434

435

437

439

440

441

442

443

444

445

446

448

450

452

453

454

455

459

460

462

463

464

465

466

467

469

471

475

477

478

479

480

482

483

484

487

491

492

495

496

498

499

500

501

506

507

508

509

514

516

| Reported value [Bq/kg ] 2430.95 | Reported uncertainty (k = 1) [Bq/kg ] 100 | Relative bias [%] 19.8 | P-Test [%] 5.34 | Trueness evaluation A | Precision evaluation N | Final Score W |
| --- | --- | --- | --- | --- | --- | --- |
| 2150 | 60 | 6.0 | 4.40 | A | A | A |
| 1936 | 55 | -4.6 | 4.43 | A | A | A |
| 2167 | 122 | 6.8 | 6.58 | A | A | A |
| 2210 | 130 | 8.9 | 6.79 | A | A | A |
| 2000 | 80 | -1.4 | 5.25 | A | A | A |
| 3515 | 18 | 73.2 | 3.44 | N | N | N |
| 2454.5 | 122.85 | 21.0 | 6.05 | N | N | N |
| 2037.01 | 164 | 0.4 | 8.74 | A | A | A |
| 2470 | 127 | 21.7 | 6.16 | N | N | N |
| 2038 | 85 | 0.4 | 5.38 | A | A | A |
| 2240 | 160 | 10.4 | 7.91 | A | A | A |
| 2075 | 98 | 2.3 | 5.82 | A | A | A |
| 1824.49 | 82.45 | -10.1 | 5.66 | A | A | A |
| 1141.5 | 55.24 | -43.7 | 5.91 | N | N | N |
| 1831 | 68 | -9.8 | 5.04 | A | A | A |
| 2087 | 167 | 2.9 | 8.69 | A | A | A |

Labcode

520

522

523

525

527

528

529

530

532

533

535

536

537

539

544

553

555

Massic activity  [ Bq/sample ]

3

4

Sr-90 determination of activity per sample in Sample 5

Table 66. Target values

| Parameter | Value |
| --- | --- |
| Massic activity [ Bq/sample ] | 3.22 |
| Uncertainty [ Bq/sample ] | 0.16 |
| Maximum Acceptable Relative Bias [%] | 30 |

Accepted 33%	Not Accepted 9%	Warning 3%	Not Reported 55%

5

Figure 66. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

2

64

44

114

145

192

278

299

49

242

183

291

204

156

51

110

260

100

194

52

344

523

191

12

66

97

127

381

55

432

384

527

181

252

264

393

152

258

439

287

369

113

58

231

241

371

159

249

288

222

283

319

347

28

341

437

33

42

132

155

293

379

443

474

276

414

536

466

5

7

81

72

178

250

128

317

79

107

240

410

13

311

45

304

346

391

405

262

428

166

516

63

267

90

208

533

170

454

508

2

19

85

463

60

91

302

555

15

38

382

75

220

207

470

445

70

214

223

553

285

257

140

420

65

133

187

286

395

265

281

27

303

158

168

370

484

433

82

552

244

83

123

529

355

34

24

78

134

475

377

69

53

215

119

122

206

255

471

495

322

40

313

254

230

342

189

440

35

314

197

492

268

31

71

435

21

491

172

165

Laboratory code

Figure 67. Reported results and evaluation parameters for Sr-90. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 67. Evaluation results for Sr-90 in Sample 5, Rep. Value and Rep. Unc. in Bq/sample

| Reported value | Reported uncertainty | Relative bias | P-Test | Trueness | Precision | Final |
| --- | --- | --- | --- | --- | --- | --- |
| [Bq/sample ] | (k = 1) [Bq/sample ] | [%] | [%] | evaluation | evaluation | Score |
| 3.31 | 0.54 | 2.8 | 17.05 | A | A | A |
| 3.17 | 0.17 | -1.6 | 7.31 | A | A | A |
| 3.17 | 0.17 | -1.6 | 7.31 | A | A | A |

Labcode

2

5

7

| Reported value [Bq/sample ] 2.94 | Reported uncertainty (k = 1) [Bq/sample ] 0.34 | Relative bias [%] -8.7 | P-Test [%] 12.59 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 3.21 | 0.08 | -0.3 | 5.56 | A | A | A |
| 3.35 | 0.32 | 4.0 | 10.77 | A | A | A |
| 3.31 | 0.07 | 2.8 | 5.40 | A | A | A |
| 5.3 | 0.2 | 64.6 | 6.24 | N | N | N |
| 3.82 | 0.5 | 18.6 | 14.00 | A | A | A |
| 3.55 | 0.18 | 10.2 | 7.10 | A | A | A |
| 3.132 | 0.279 | -2.7 | 10.20 | A | A | A |
| 5.04 | 0.35 | 56.5 | 8.54 | N | N | N |
| 3.15 | 0.16 | -2.2 | 7.11 | A | A | A |
| 3.81 | 0.35 | 18.3 | 10.44 | A | A | A |
| 4.8 | 1.9 | 49.1 | 39.89 | N | N | N |
| 3.35 | 0.19 | 4.0 | 7.54 | A | A | A |
| 4.47 | 0.5 | 38.8 | 12.24 | N | N | N |
| 3.15 | 0.06 | -2.2 | 5.32 | A | A | A |
| 1.26 | 0.05 | -60.9 | 6.36 | N | N | N |
| 3.22 | 0.21 | 0.0 | 8.20 | A | A | A |
| 2.6 | 0.5 | -19.2 | 19.86 | A | A | A |
| 2.83 | 0.07 | -12.1 | 5.55 | A | A | A |
| 2.89 | 8.62 | -10.2 | 298.31 | A | N | W |
| 4.28 | 0.17 | 32.9 | 6.36 | N | N | N |
| 2.97 | 0.12 | -7.8 | 6.40 | A | A | A |
| 3.1 | 0.2 | -3.7 | 8.14 | A | A | A |
| 3.33 | 0.55 | 3.4 | 17.25 | A | A | A |
| 3.26 | 0.088 | 1.2 | 5.65 | A | A | A |
| 1.2 | 0.04 | -62.7 | 5.98 | N | N | N |
| 3.5 | 0.2 | 8.7 | 7.57 | A | A | A |
| 2.942 | 0.1851 | -8.6 | 8.02 | A | A | A |
| 4.26 | 0.22 | 32.3 | 7.17 | N | N | N |
| 3.4 | 0.33 | 5.6 | 10.90 | A | A | A |
| 5.1 | 0.3 | 58.4 | 7.70 | N | N | N |
| 3.18 | 0.11 | -1.2 | 6.05 | A | A | A |
| 3.36 | 0.06 | 4.3 | 5.28 | A | A | A |
| 3.96 | 0.45 | 23.0 | 12.40 | A | A | A |
| 3.2 | 0.13 | -0.6 | 6.42 | A | A | A |
| 3.17 | 0.19 | -1.6 | 7.79 | A | A | A |
| 3.63 | 0.08 | 12.7 | 5.44 | A | A | A |
| 3.71 | 0.72 | 15.2 | 20.03 | A | A | A |
| 3.31 | 0.538 | 2.8 | 17.00 | A | A | A |
| 3.27 | 0.18 | 1.6 | 7.42 | A | A | A |
| 3.34 | 0.18 | 3.7 | 7.33 | A | A | A |
| 2.96 | 1.28 | -8.1 | 43.53 | A | N | W |
| 2.88 | 0.12 | -10.6 | 6.48 | A | A | A |
| 3.2 | 0.3 | -0.6 | 10.61 | A | A | A |
| 2.84 | 0.11 | -11.8 | 6.30 | A | A | A |
| 3.09 | 0.3 | -4.0 | 10.91 | A | A | A |
| 1.48 | 0.15 | -54.0 | 11.29 | N | N | N |
| 4.37 | 0.32 | 35.7 | 8.85 | N | N | N |
| 4.4 | 0.5 | 36.6 | 12.40 | N | N | N |
| 3.71 | 0.21 | 15.2 | 7.53 | A | A | A |
| 2.96 | 0.24 | -8.1 | 9.51 | A | A | A |
| 3.19 | 0.11 | -0.9 | 6.05 | A | A | A |
| 3.15 | 0.27 | -2.2 | 9.91 | A | A | A |
| 3.5 | 0.04 | 8.7 | 5.10 | A | A | A |
| 4.03 | 0.11 | 25.2 | 5.67 | A | N | W |
| 3.44 | 0.31 | 6.8 | 10.29 | A | A | A |
| 2.2 | 0.44 | -31.7 | 20.61 | N | A | N |
| 3.05 | 0.188 | -5.3 | 7.92 | A | A | A |
| 3.15 | 0.35 | -2.2 | 12.17 | A | A | A |
| 2.82 | 0.1 | -12.4 | 6.10 | A | A | A |

Labcode

12

13

15

19

21

24

27

28

31

33

34

35

38

40

42

44

45

49

51

52

53

55

58

60

63

64

65

66

69

70

71

72

75

78

79

81

82

83

85

90

91

97

100

107

110

113

114

119

122

123

127

128

132

133

134

140

145

152

155

156

| Reported value [Bq/sample ] 3.6 | Reported uncertainty (k = 1) [Bq/sample ] 0.1 | Relative bias [%] 11.8 | P-Test [%] 5.69 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 3.11 | 0.16 | -3.4 | 7.15 | A | A | A |
| 9.47 | 0.086 | 194.1 | 5.05 | N | N | N |
| 3.25 | 0.21 | 0.9 | 8.15 | A | A | A |
| 3.6 | 0.4 | 11.8 | 12.17 | A | A | A |
| 3.29 | 0.13 | 2.2 | 6.35 | A | A | A |
| 6.74 | 0.4 | 109.3 | 7.74 | N | N | N |
| 3.18 | 0.16 | -1.2 | 7.07 | A | A | A |
| 3.01 | 0.34 | -6.5 | 12.34 | A | A | A |
| 2.67 | 0.04 | -17.1 | 5.19 | A | N | W |
| 3.5 | 0.2 | 8.7 | 7.57 | A | A | A |
| 4.71 | 0.09 | 46.3 | 5.32 | N | N | N |
| 2.93 | 0.26 | -9.0 | 10.17 | A | A | A |
| 2.27 | 0.162 | -29.5 | 8.70 | A | N | W |
| 2.88 | 0.12 | -10.6 | 6.48 | A | A | A |
| 4.98 | 0.99 | 54.7 | 20.49 | N | N | N |
| 2.7 | 0.2 | -16.1 | 8.92 | A | A | A |
| 4.4 | 1.7 | 36.6 | 38.95 | N | N | N |
| 3.37 | 0.08 | 4.7 | 5.51 | A | A | A |
| 3.28 | 0.12 | 1.9 | 6.17 | A | A | A |
| 3.4 | 0.01 | 5.6 | 4.98 | A | A | A |
| 4.32 | 0.54 | 34.2 | 13.45 | N | A | N |
| 3.36 | 0.28 | 4.3 | 9.70 | A | A | A |
| 3.13 | 0.23 | -2.8 | 8.87 | A | A | A |
| 3.4 | 0.4 | 5.6 | 12.77 | A | A | A |
| 4.7 | 0.1 | 46.0 | 5.41 | N | N | N |
| 3.1 | 0.4 | -3.7 | 13.83 | A | A | A |
| 3.2 | 0.16 | -0.6 | 7.05 | A | A | A |
| 3.1 | 0.22 | -3.7 | 8.66 | A | A | A |
| 2.66 | 0.4 | -17.4 | 15.84 | A | A | A |
| 3.7 | 0.4 | 14.9 | 11.90 | A | A | A |
| 3.11 | 0.18 | -3.4 | 7.63 | A | A | A |
| 3.18 | 0.16 | -1.2 | 7.07 | A | A | A |
| 3.03 | 0.2 | -5.9 | 8.26 | A | A | A |
| 4.6 | 0.08 | 42.9 | 5.26 | N | N | N |
| 4.4 | 0.22 | 36.6 | 7.05 | N | N | N |
| 3.43 | 0.68 | 6.5 | 20.44 | A | A | A |
| 3.05 | 0.12 | -5.3 | 6.34 | A | A | A |
| 2.84 | 0.31 | -11.8 | 11.99 | A | A | A |
| 3.23 | 0.31 | 0.3 | 10.81 | A | A | A |
| 3.03 | 0.15 | -5.9 | 7.01 | A | A | A |
| 3.53 | 0.53 | 9.6 | 15.82 | A | A | A |
| 3.26 | 0.482 | 1.2 | 15.60 | A | A | A |
| 5.02 | 0.25 | 55.9 | 7.04 | N | N | N |
| 3.16 | 0.23 | -1.9 | 8.81 | A | A | A |
| 2.49 | 0.14 | -22.7 | 7.50 | A | N | W |
| 3.53 | 0.07 | 9.6 | 5.35 | A | A | A |
| 3.13 | 0.12 | -2.8 | 6.28 | A | A | A |
| 3.42 | 0.18 | 6.2 | 7.24 | A | A | A |
| 3.5 | 0.2 | 8.7 | 7.57 | A | A | A |
| 3.07 | 0.15 | -4.7 | 6.97 | A | A | A |
| 3.11 | 0.16 | -3.4 | 7.15 | A | A | A |
| 2.69 | 0.51 | -16.5 | 19.60 | A | A | A |
| 3.15 | 0.39 | -2.2 | 13.34 | A | A | A |
| 2.5 | 1.1 | -22.4 | 44.28 | A | N | W |
| 3.34 | 0.4 | 3.7 | 12.97 | A | A | A |
| 3.56 | 0.15 | 10.6 | 6.51 | A | A | A |
| 3.22 | 0.18 | 0.0 | 7.48 | A | A | A |
| 3.21 | 0.0785 | -0.3 | 5.54 | A | A | A |
| 4.53 | 0.11 | 40.7 | 5.53 | N | N | N |

Labcode

158

159

165

166

168

170

172

178

181

183

187

189

191

192

194

197

204

206

207

208

214

215

220

222

223

230

231

240

241

242

244

249

250

252

254

255

257

258

260

262

264

265

267

268

276

278

281

283

285

286

287

288

291

293

299

302

303

304

311

313

| Reported value [Bq/sample ] 4.95 | Reported uncertainty (k = 1) [Bq/sample ] 0.17 | Relative bias [%] 53.7 | P-Test [%] 6.04 | Trueness evaluation N | Precision evaluation N | Final Score N |
| --- | --- | --- | --- | --- | --- | --- |
| 3.19 | 0.19 | -0.9 | 7.76 | A | A | A |
| 3.13 | 0.21 | -2.8 | 8.35 | A | A | A |
| 4.46 | 0.45 | 38.5 | 11.25 | N | N | N |
| 3.14 | 0.19 | -2.5 | 7.83 | A | A | A |
| 4.7 | 0.2 | 46.0 | 6.54 | N | N | N |
| 2.9 | 0.3 | -9.9 | 11.48 | A | A | A |
| 3.22 | 0.5 | 0.0 | 16.30 | A | A | A |
| 3.13 | 0.41 | -2.8 | 14.01 | A | A | A |
| 3.8 | 0.13 | 18.0 | 6.03 | A | N | W |
| 3.07 | 0.1 | -4.7 | 5.94 | A | A | A |
| 3.6 | 0.5 | 11.8 | 14.75 | A | A | A |
| 3.1 | 0.2 | -3.7 | 8.14 | A | A | A |
| 4.23 | 0.012 | 31.4 | 4.98 | N | N | N |
| 3.15 | 0.4 | -2.2 | 13.64 | A | A | A |
| 2.96 | 0.28 | -8.1 | 10.69 | A | A | A |
| 3.35 | 0.33 | 4.0 | 11.03 | A | A | A |
| 3 | 0.6 | -6.8 | 20.61 | A | A | A |
| 3.22 | 0.06 | 0.0 | 5.31 | A | A | A |
| 3.04 | 0.34 | -5.6 | 12.24 | A | A | A |
| 3.5 | 0.36 | 8.7 | 11.42 | A | A | A |
| 3.22 | 0.18 | 0.0 | 7.48 | A | A | A |
| 3.2 | 0.3 | -0.6 | 10.61 | A | A | A |
| 3.16 | 0.22 | -1.9 | 8.55 | A | A | A |
| 3.44 | 0.07 | 6.8 | 5.37 | A | A | A |
| 3.24 | 0.7 | 0.6 | 22.17 | A | A | A |
| 2.98 | 0.187 | -7.4 | 8.00 | A | A | A |
| 3.61 | 0.22 | 12.1 | 7.86 | A | A | A |
| 5.2 | 1.2 | 61.5 | 23.61 | N | N | N |
| 3.14 | 0.16 | -2.5 | 7.12 | A | A | A |
| 3.06 | 0.12 | -5.0 | 6.33 | A | A | A |
| 4.751 | 0.013 | 47.5 | 4.98 | N | N | N |
| 3.15 | 0.25 | -2.2 | 9.36 | A | A | A |
| 3.39 | 0.15 | 5.3 | 6.65 | A | A | A |
| 3.29 | 0.21 | 2.2 | 8.09 | A | A | A |
| 3.32 | 0.17 | 3.1 | 7.14 | A | A | A |
| 3.165 | 0.09 | -1.7 | 5.73 | A | A | A |
| 3.38 | 0.68 | 5.0 | 20.72 | A | A | A |
| 4.43 | 0.06 | 37.6 | 5.15 | N | N | N |
| 3.156 | 0.219 | -2.0 | 8.53 | A | A | A |
| 4.1 | 0.3 | 27.3 | 8.84 | A | N | W |
| 3.6 | 0.3 | 11.8 | 9.70 | A | A | A |
| 6.43 | 0.174 | 99.7 | 5.66 | N | N | N |
| 4.98 | 0.4 | 54.7 | 9.44 | N | N | N |
| 4.43 | 0.09 | 37.6 | 5.37 | N | N | N |
| 3.29 | 0.07 | 2.2 | 5.41 | A | A | A |
| 3.25 | 0.18 | 0.9 | 7.44 | A | A | A |
| 2.92 | 0.06 | -9.3 | 5.38 | A | A | A |
| 3 | 0.17 | -6.8 | 7.54 | A | A | A |
| 3.71 | 0.11 | 15.2 | 5.79 | A | N | W |
| 3.28 | 0.06 | 1.9 | 5.29 | A | A | A |
| 3.16 | 0.21 | -1.9 | 8.30 | A | A | A |
| 3.67 | 0.04 | 14.0 | 5.09 | A | N | W |
| 3.4 | 0.2 | 5.6 | 7.70 | A | A | A |
| 3.34 | 0.4 | 3.7 | 12.97 | A | A | A |

Labcode

314

317

319

322

341

342

344

346

347

355

369

370

371

377

379

381

382

384

391

393

395

405

410

414

420

428

432

433

435

437

439

440

443

445

454

463

466

470

471

474

475

484

491

492

495

508

516

523

527

529

533

536

552

553

555

Ba-133 determination of activity per sample in Sample 5

Table 68. Target values

| Parameter | Value |
| --- | --- |
| Massic activity [ Bq/sample ] | 10.1 |
| Uncertainty [ Bq/sample ] | 0.51 |
| Maximum Acceptable Relative Bias [%] | 30 |

Accepted 66%	Not Accepted 2%	Warning 2%	Not Reported 30%

14

16

Figure 68. Percentage of Accepted vs. Warnings vs. Not Accepted vs. Not Reported results

4

6

8

158

218

544

271

555

34

277

540

412

204

35

85

44

373

398

66

508

227

326

286

154

100

420

304

83

284

357

408

453

200

439

356

381

50

328

159

364

184

291

41

27

220

290

514

478

155

97

479

156

9

69

62

250

256

371

454

247

107

465

252

288

15

192

313

426

19

42

206

233

241

303

331

45

128

90

79

318

267

474

6

121

133

264

279

332

354

384

433

450

527

8

165

441

553

500

134

25

170

238

344

443

71

40

189

261

529

293

57

52

145

432

487

445

302

516

122

168

405

2

166

123

60

65

311

492

491

319

51

414

338

240

273

217

172

370

12

393

520

333

Laboratory code

Figure 69. Reported results and evaluation parameters for Ba-133. Blue points represent accepted scores, whereas yellow and red represent warning and not accepted scores, respectively.

Table 69. Evaluation results for Ba-133 in Sample 5, Rep. Value and Rep. Unc. in Bq/sample

Labcode

Reported value

[Bq/sample ]

Reported uncertainty

(k = 1) [Bq/sample ]

Relative bias

[%]

P-Test

[%]

Trueness

evaluation

Precision

evaluation

Final

Score

Massic activity  [ Bq/sample ]

10

12

2	10.53	0.77	4.3	8.89	A	A	A

5	10.1	0.4	0.0	6.42	A	A	A

6	10.1	0.34	0.0	6.07	A	A	A

| Reported value [Bq/sample ] 10.1 | Reported uncertainty (k = 1) [Bq/sample ] 0.38 | Relative bias [%] 0.0 | P-Test [%] 6.30 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 10.11 | 0.49 | 0.1 | 7.00 | A | A | A |
| 9.9 | 0.4 | -2.0 | 6.47 | A | A | A |
| 10 | 0.631 | -1.0 | 8.08 | A | A | A |
| 11.35 | 1.45 | 12.4 | 13.74 | A | A | A |
| 9.73 | 0.36 | -3.7 | 6.26 | A | A | A |
| 9.98 | 0.44 | -1.2 | 6.70 | A | A | A |
| 9.98 | 0.52 | -1.2 | 7.26 | A | A | A |
| 10 | 0.034 | -1.0 | 5.06 | A | A | A |
| 10 | 0.4 | -1.0 | 6.44 | A | A | A |
| 8.05 | 0.77 | -20.3 | 10.82 | A | A | A |
| 10.2 | 0.4 | 1.0 | 6.39 | A | A | A |
| 9.8 | 0.62 | -3.0 | 8.09 | A | A | A |
| 9.95 | 0.79 | -1.5 | 9.41 | A | A | A |
| 8.1 | 0.7 | -19.8 | 10.01 | A | A | A |
| 9.1 | 0.6 | -9.9 | 8.30 | A | A | A |
| 9.79 | 1.16 | -3.1 | 12.88 | A | A | A |
| 10.22 | 0.77 | 1.2 | 9.07 | A | A | A |
| 9.79 | 0.133 | -3.1 | 5.23 | A | A | A |
| 10 | 0.6 | -1.0 | 7.84 | A | A | A |
| 9.2 | 0.3 | -8.9 | 6.01 | A | A | A |
| 10.01 | 0.65 | -0.9 | 8.23 | A | A | A |
| 9.6 | 0.6 | -5.0 | 8.03 | A | A | A |
| 9.9 | 1.3 | -2.0 | 14.07 | A | A | A |
| 9.7 | 0.9 | -4.0 | 10.56 | A | A | A |
| 10.79 | 0.13 | 6.8 | 5.19 | A | A | A |
| 10.3 | 0.54 | 2.0 | 7.28 | A | A | A |
| 10.3 | 0.8 | 2.0 | 9.26 | A | A | A |
| 8.88 | 0.39 | -12.1 | 6.69 | A | A | A |
| 10.26 | 1 | 1.6 | 10.98 | A | A | A |
| 9.5 | 0.4 | -5.9 | 6.57 | A | A | A |
| 10.57 | 0.78 | 4.7 | 8.94 | A | A | A |
| 9.91 | 0.9 | -1.9 | 10.39 | A | A | A |
| 10.01 | 0.79 | -0.9 | 9.37 | A | A | A |
| 8.2 | 0.8 | -18.8 | 10.99 | A | A | A |
| 10.6 | 0.8 | 5.0 | 9.08 | A | A | A |
| 9.207 | 0.4802 | -8.8 | 7.26 | A | A | A |
| 9.9 | 0.4 | -2.0 | 6.47 | A | A | A |
| 9.86 | 0.38 | -2.4 | 6.35 | A | A | A |
| 10.21 | 0.51 | 1.1 | 7.10 | A | A | A |
| 9.9 | 0.35 | -2.0 | 6.16 | A | A | A |
| 10.27 | 0.35 | 1.7 | 6.09 | A | A | A |
| 10.22 | 0.63 | 1.2 | 7.97 | A | A | A |
| 10.03 | 0.3 | -0.7 | 5.87 | A | A | A |
| 9.8961 | 0.060094 | -2.0 | 5.09 | A | A | A |
| 10.19 | 0.22 | 0.9 | 5.49 | A | A | A |
| 9.96 | 0.22 | -1.4 | 5.51 | A | A | A |
| 9.5 | 1.4 | -5.9 | 15.58 | A | A | A |
| 6.78 | 0.09 | -32.9 | 5.22 | N | N | N |
| 9.11 | 0.476 | -9.8 | 7.27 | A | A | A |
| 10.02 | 0.45 | -0.8 | 6.76 | A | A | A |
| 10.14 | 0.31 | 0.4 | 5.90 | A | A | A |
| 9.86 | 0.5 | -2.4 | 7.16 | A | A | A |
| 9.46 | 0.62 | -6.3 | 8.27 | A | A | A |
| 11.3 | 0.9 | 11.9 | 9.43 | A | A | A |
| 9.96 | 0.7 | -1.4 | 8.65 | A | A | A |
| 9.53 | 0.62 | -5.6 | 8.24 | A | A | A |
| 10.53 | 0.82 | 4.3 | 9.28 | A | A | A |
| 9.97 | 0.6 | -1.3 | 7.86 | A | A | A |
| 9.82 | 0.68 | -2.8 | 8.57 | A | A | A |

Labcode

7

8

9

11

12

13

15

17

19

21

24

25

27

31

34

35

38

40

41

42

44

45

48

49

50

51

52

53

55

57

58

60

62

63

64

65

66

69

70

71

72

75

78

79

80

81

82

83

84

85

90

94

97

100

102

107

110

111

113

114

| Reported value [Bq/sample ] 10 | Reported uncertainty (k = 1) [Bq/sample ] 0.9 | Relative bias [%] -1.0 | P-Test [%] 10.32 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 10.1 | 0.7 | 0.0 | 8.58 | A | A | A |
| 10.5 | 0.9 | 4.0 | 9.95 | A | A | A |
| 10.56 | 0.53 | 4.6 | 7.12 | A | A | A |
| 10.2 | 1.1 | 1.0 | 11.91 | A | A | A |
| 10.1 | 0.7 | 0.0 | 8.58 | A | A | A |
| 10.01 | 0.36 | -0.9 | 6.20 | A | A | A |
| 10.5 | 0.9 | 4.0 | 9.95 | A | A | A |
| 10.1 | 0.1 | 0.0 | 5.15 | A | A | A |
| 10.17 | 0.04 | 0.7 | 5.06 | A | A | A |
| 10.88 | 0.68 | 7.7 | 8.03 | A | A | A |
| 10.3 | 1.64 | 2.0 | 16.70 | A | A | A |
| 9.41 | 0.47 | -6.8 | 7.10 | A | A | A |
| 9.8 | 0.6 | -3.0 | 7.94 | A | A | A |
| 9.41 | 0.48 | -6.8 | 7.18 | A | A | A |
| 9.84 | 0.74 | -2.6 | 9.06 | A | A | A |
| 9.89 | 0.16 | -2.1 | 5.30 | A | A | A |
| 0.38 | 0.04 | -96.2 | 11.67 | N | N | N |
| 9.71 | 0.15 | -3.9 | 5.28 | A | A | A |
| 10.12 | 0.29 | 0.2 | 5.81 | A | A | A |
| 10.55 | 0.35 | 4.5 | 6.04 | A | A | A |
| 10.5 | 0.9 | 4.0 | 9.95 | A | A | A |
| 10.2 | 0.3 | 1.0 | 5.84 | A | A | A |
| 11 | 0.8 | 8.9 | 8.85 | A | A | A |
| 9.32 | 0.52 | -7.7 | 7.53 | A | A | A |
| 10.1 | 0.5 | 0.0 | 7.07 | A | A | A |
| 9.17 | 0.103 | -9.2 | 5.17 | A | A | A |
| 10.52 | 0.74 | 4.2 | 8.66 | A | A | A |
| 9.4 | 0.5 | -6.9 | 7.33 | A | A | A |
| 9.765 | 1.25 | -3.3 | 13.76 | A | A | A |
| 10.6 | 0.8 | 5.0 | 9.08 | A | A | A |
| 10.22 | 0.12 | 1.2 | 5.18 | A | A | A |
| 9.67 | 0.8 | -4.3 | 9.69 | A | A | A |
| 9.98 | 0.998 | -1.2 | 11.20 | A | A | A |
| 9.78 | 0.28 | -3.2 | 5.80 | A | A | A |
| 10.22 | 0.87 | 1.2 | 9.90 | A | A | A |
| 9.6 | 0.4 | -5.0 | 6.55 | A | A | A |
| 10.12 | 0.38 | 0.2 | 6.29 | A | A | A |
| 9 | 0.3 | -10.9 | 6.05 | A | A | A |
| 9.98 | 0.638 | -1.2 | 8.15 | A | A | A |
| 10 | 1.2 | -1.0 | 13.02 | A | A | A |
| 10.07 | 0.35 | -0.3 | 6.13 | A | A | A |
| 9.96 | 0.8 | -1.4 | 9.49 | A | A | A |
| 9.88 | 0.92 | -2.2 | 10.59 | A | A | A |
| 11 | 0.6 | 8.9 | 7.43 | A | A | A |
| 10.2 | 0.9 | 1.0 | 10.17 | A | A | A |
| 10.98 | 2 | 8.7 | 18.90 | A | A | A |
| 6.76588 | 0.2869 | -33.0 | 6.59 | N | N | N |
| 9.8 | 0.8 | -3.0 | 9.60 | A | A | A |
| 9.8 | 0.4 | -3.0 | 6.49 | A | A | A |
| 7.56 | 0.33 | -25.1 | 6.67 | A | N | W |
| 9.32 | 0.29 | -7.7 | 5.93 | A | A | A |
| 10 | 0.3 | -1.0 | 5.87 | A | A | A |
| 10.9 | 0.5 | 7.9 | 6.82 | A | A | A |
| 10 | 0.7 | -1.0 | 8.63 | A | A | A |
| 10 | 0.8 | -1.0 | 9.46 | A | A | A |
| 10.2 | 0.311 | 1.0 | 5.90 | A | A | A |
| 10.9 | 0.53 | 7.9 | 7.01 | A | A | A |
| 10 | 0.5 | -1.0 | 7.11 | A | A | A |
| 9.94 | 0.99 | -1.6 | 11.17 | A | A | A |

Labcode

119

121

122

123

124

127

128

132

133

134

140

145

150

152

154

155

156

158

159

165

166

168

170

172

175

178

179

181

183

184

187

189

191

192

194

197

200

203

204

205

206

207

208

213

214

215

217

218

220

223

226

227

230

232

233

235

238

240

241

242

| Reported value [Bq/sample ] 9.95 | Reported uncertainty (k = 1) [Bq/sample ] 0.729 | Relative bias [%] -1.5 | P-Test [%] 8.90 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 19.62 | 1.08 | 94.3 | 7.47 | N | N | N |
| 9.92 | 0.62 | -1.8 | 8.03 | A | A | A |
| 9.97 | 0.42 | -1.3 | 6.58 | A | A | A |
| 9.1 | 0.5 | -9.9 | 7.46 | A | A | A |
| 9.92 | 0.35 | -1.8 | 6.16 | A | A | A |
| 9.92 | 0.4 | -1.8 | 6.46 | A | A | A |
| 9.37 | 0.94 | -7.2 | 11.23 | A | A | A |
| 10.43 | 0.31 | 3.3 | 5.86 | A | A | A |
| 10.9 | 0.83 | 7.9 | 9.14 | A | A | A |
| 10.22 | 0.23 | 1.2 | 5.53 | A | A | A |
| 10.55 | 0.22 | 4.5 | 5.46 | A | A | A |
| 10.1 | 0.8 | 0.0 | 9.39 | A | A | A |
| 10 | 1.2 | -1.0 | 13.02 | A | A | A |
| 10.08 | 0.97 | -0.2 | 10.87 | A | A | A |
| 9.42 | 0.47 | -6.7 | 7.10 | A | A | A |
| 9.7 | 0.9 | -4.0 | 10.56 | A | A | A |
| 7.7 | 0.7 | -23.8 | 10.40 | A | A | A |
| 10.1 | 0.3 | 0.0 | 5.86 | A | A | A |
| 10.9 | 0.6 | 7.9 | 7.47 | A | A | A |
| 9.08 | 0.48 | -10.1 | 7.31 | A | A | A |
| 8.2 | 0.4 | -18.8 | 7.02 | A | N | W |
| 8.35 | 0.1 | -17.3 | 5.19 | A | N | W |
| 10.1 | 0.5 | 0.0 | 7.07 | A | A | A |
| 10.25 | 0.14 | 1.5 | 5.23 | A | A | A |
| 9.97 | 0.33 | -1.3 | 6.04 | A | A | A |
| 9.52 | 0.78 | -5.7 | 9.62 | A | A | A |
| 10.02 | 0.54 | -0.8 | 7.39 | A | A | A |
| 9.4 | 0.4 | -6.9 | 6.60 | A | A | A |
| 10.1 | 0.3 | 0.0 | 5.86 | A | A | A |
| 9.97 | 0.6 | -1.3 | 7.86 | A | A | A |
| 9.8 | 0.8 | -3.0 | 9.60 | A | A | A |
| 9.78 | 1.174 | -3.2 | 13.02 | A | A | A |
| 10.2 | 0.2 | 1.0 | 5.42 | A | A | A |
| 10.24 | 0.73 | 1.4 | 8.74 | A | A | A |
| 10.11 | 0.76 | 0.1 | 9.06 | A | A | A |
| 10.4 | 0.4 | 3.0 | 6.35 | A | A | A |
| 32.32 | 0.44 | 220.0 | 5.23 | N | N | N |
| 10.4 | 1.1 | 3.0 | 11.72 | A | A | A |
| 10 | 0.7 | -1.0 | 8.63 | A | A | A |
| 9.49 | 0.92 | -6.0 | 10.93 | A | A | A |
| 10.7 | 1.3 | 5.9 | 13.16 | A | A | A |
| 9.6 | 1.5 | -5.0 | 16.42 | A | A | A |
| 10.3 | 0.36 | 2.0 | 6.14 | A | A | A |
| 10.6 | 1.38 | 5.0 | 13.96 | A | A | A |
| 9.99 | 0.51 | -1.1 | 7.18 | A | A | A |
| 10.23 | 0.35 | 1.3 | 6.10 | A | A | A |
| 10.567 | 0.492 | 4.6 | 6.87 | A | A | A |
| 9.65 | 0.7 | -4.5 | 8.84 | A | A | A |
| 10.06 | 0.27 | -0.4 | 5.72 | A | A | A |
| 10.7 | 0.62 | 5.9 | 7.69 | A | A | A |
| 10 | 0.3 | -1.0 | 5.87 | A | A | A |
| 9.48 | 0.6 | -6.1 | 8.10 | A | A | A |
| 9.38 | 0.8 | -7.1 | 9.91 | A | A | A |
| 9.7 | 1.4 | -4.0 | 15.29 | A | A | A |
| 9.92 | 0.354 | -1.8 | 6.18 | A | A | A |
| 10 | 0.5 | -1.0 | 7.11 | A | A | A |
| 10.1 | 0.7 | 0.0 | 8.58 | A | A | A |
| 26.7 | 0.259 | 164.4 | 5.14 | N | N | N |
| 10.88 | 0.93 | 7.7 | 9.93 | A | A | A |

Labcode

247

249

250

252

253

254

256

257

258

260

261

262

264

265

267

268

270

271

272

273

276

277

278

279

281

283

284

285

286

287

288

290

291

292

293

295

299

300

302

303

304

306

307

310

311

313

314

315

317

318

319

322

325

326

328

330

331

332

333

338

| Reported value [Bq/sample ] 10.1 | Reported uncertainty (k = 1) [Bq/sample ] 0.6 | Relative bias [%] 0.0 | P-Test [%] 7.80 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 10.01 | 0.26 | -0.9 | 5.68 | A | A | A |
| 10.2 | 0.7 | 1.0 | 8.52 | A | A | A |
| 10.22 | 0.45 | 1.2 | 6.70 | A | A | A |
| 11.38 | 0.55 | 12.7 | 6.99 | A | A | A |
| 9.7 | 0.4 | -4.0 | 6.52 | A | A | A |
| 9.2 | 0.4 | -8.9 | 6.66 | A | A | A |
| 10.1 | 0.5 | 0.0 | 7.07 | A | A | A |
| 9.97 | 0.05 | -1.3 | 5.07 | A | A | A |
| 9.65 | 0.5 | -4.5 | 7.23 | A | A | A |
| 9.53 | 0.48 | -5.6 | 7.13 | A | A | A |
| 9.74 | 0.91 | -3.6 | 10.62 | A | A | A |
| 10.1 | 0.31 | 0.0 | 5.91 | A | A | A |
| 11.2 | 1.5 | 10.9 | 14.31 | A | A | A |
| 9.93 | 0.49 | -1.7 | 7.06 | A | A | A |
| 9.2 | 0.35 | -8.9 | 6.32 | A | A | A |
| 7.72 | 0.31 | -23.6 | 6.45 | A | N | W |
| 10.5 | 1.6 | 4.0 | 16.05 | A | A | A |
| 9.66 | 0.5 | -4.4 | 7.23 | A | A | A |
| 9.65 | 0.66 | -4.5 | 8.50 | A | A | A |
| 10.1 | 1.3 | 0.0 | 13.83 | A | A | A |
| 8.57 | 1.15 | -15.2 | 14.34 | A | A | A |
| 11.62 | 0.95 | 15.1 | 9.61 | A | A | A |
| 9.2 | 0.16 | -8.9 | 5.34 | A | A | A |
| 11.4 | 0.6 | 12.9 | 7.29 | A | A | A |
| 10.21 | 1 | 1.1 | 11.02 | A | A | A |
| 9.8 | 0.9 | -3.0 | 10.48 | A | A | A |
| 9.2 | 1.6 | -8.9 | 18.11 | A | A | A |
| 10.84 | 0.39 | 7.3 | 6.20 | A | A | A |
| 9.55 | 0.54 | -5.4 | 7.58 | A | A | A |
| 10.58 | 0.73 | 4.8 | 8.55 | A | A | A |
| 10.08 | 0.51 | -0.2 | 7.15 | A | A | A |
| 10.5 | 0.55 | 4.0 | 7.28 | A | A | A |
| 10.7 | 0.6 | 5.9 | 7.55 | A | A | A |
| 9.55 | 0.23 | -5.4 | 5.59 | A | A | A |
| 9.3 | 0.6 | -7.9 | 8.19 | A | A | A |
| 9.86 | 0.34 | -2.4 | 6.11 | A | A | A |
| 8.7 | 0.21 | -13.9 | 5.60 | A | A | A |
| 10.87 | 0.68 | 7.6 | 8.04 | A | A | A |
| 10.1 | 0.4 | 0.0 | 6.42 | A | A | A |
| 0.43 | 0.04 | -95.7 | 10.58 | N | N | N |
| 9.47 | 0.7 | -6.2 | 8.95 | A | A | A |
| 10 | 0.07 | -1.0 | 5.10 | A | A | A |
| 9.91 | 0.5 | -1.9 | 7.14 | A | A | A |
| 9.9958 | 0.0801 | -1.0 | 5.11 | A | A | A |
| 9.2 | 0.6 | -8.9 | 8.25 | A | A | A |
| 10.2 | 0.245 | 1.0 | 5.59 | A | A | A |
| 10.3 | 0.632 | 2.0 | 7.95 | A | A | A |
| 10.1 | 0.19 | 0.0 | 5.39 | A | A | A |
| 9.5 | 3 | -5.9 | 31.98 | A | N | W |
| 10.1 | 0.7 | 0.0 | 8.58 | A | A | A |
| 9.6 | 0.4 | -5.0 | 6.55 | A | A | A |
| 10.164 | 0.061 | 0.6 | 5.09 | A | A | A |
| 10.12 | 0.34 | 0.2 | 6.07 | A | A | A |
| 10.2 | 0.7 | 1.0 | 8.52 | A | A | A |
| 10.31 | 0.84 | 2.1 | 9.59 | A | A | A |
| 10.3 | 0.5 | 2.0 | 7.00 | A | A | A |
| 10.1 | 0.3 | 0.0 | 5.86 | A | A | A |
| 9.79 | 0.55 | -3.1 | 7.55 | A | A | A |
| 9.57 | 0.21 | -5.2 | 5.51 | A | A | A |

Labcode

339

341

344

346

347

348

350

354

355

356

357

364

369

370

371

373

376

379

381

382

384

385

386

391

393

395

396

398

401

402

403

404

405

407

408

410

411

412

414

417

419

420

423

425

426

428

429

432

433

435

437

439

440

441

443

445

446

450

452

453

| Reported value [Bq/sample ] 9.94 | Reported uncertainty (k = 1) [Bq/sample ] 0.49 | Relative bias [%] -1.6 | P-Test [%] 7.06 | Trueness evaluation A | Precision evaluation A | Final Score A |
| --- | --- | --- | --- | --- | --- | --- |
| 9.76 | 0.78 | -3.4 | 9.45 | A | A | A |
| 10.05 | 0.4 | -0.5 | 6.43 | A | A | A |
| 9.96 | 0.9 | -1.4 | 10.35 | A | A | A |
| 10.6 | 0.36 | 5.0 | 6.09 | A | A | A |
| 10.1 | 0.29 | 0.0 | 5.81 | A | A | A |
| 10.089 | 0.958 | -0.1 | 10.75 | A | A | A |
| 10.4 | 0.5 | 3.0 | 6.97 | A | A | A |
| 9.82 | 0.44 | -2.8 | 6.75 | A | A | A |
| 9.86 | 0.79 | -2.4 | 9.47 | A | A | A |
| 10.12 | 0.34 | 0.2 | 6.07 | A | A | A |
| 10.2 | 0.4 | 1.0 | 6.39 | A | A | A |
| 10.3 | 1 | 2.0 | 10.94 | A | A | A |
| 10.681 | 0.187 | 5.8 | 5.34 | A | A | A |
| 10.65 | 0.9 | 5.4 | 9.84 | A | A | A |
| 10.99 | 0.35 | 8.8 | 5.97 | A | A | A |
| 10.91 | 1 | 8.0 | 10.46 | A | A | A |
| 10.16 | 1.036 | 0.6 | 11.38 | A | A | A |
| 9.3 | 0.7 | -7.9 | 9.06 | A | A | A |
| 9.8 | 0.7 | -3.0 | 8.75 | A | A | A |
| 10.4 | 0.59 | 3.0 | 7.59 | A | A | A |
| 14.35 | 0.61 | 42.1 | 6.60 | N | N | N |
| 9.82 | 0.38 | -2.8 | 6.36 | A | A | A |
| 9.56 | 0.59 | -5.3 | 7.97 | A | A | A |
| 10.1 | 0.58 | 0.0 | 7.65 | A | A | A |
| 10.1 | 0.5 | 0.0 | 7.07 | A | A | A |
| 10.22 | 0.36 | 1.2 | 6.16 | A | A | A |
| 9.46 | 0.34 | -6.3 | 6.20 | A | A | A |
| 10.3 | 0.17 | 2.0 | 5.31 | A | A | A |
| 10.65 | 0.24 | 5.4 | 5.53 | A | A | A |
| 9.99 | 0.35 | -1.1 | 6.15 | A | A | A |
| 8.53 | 1.18 | -15.5 | 14.73 | A | A | A |
| 7.438 | 0.099 | -26.4 | 5.22 | A | N | W |
| 10.13 | 0.34 | 0.3 | 6.06 | A | A | A |
| 7.8 | 0.5 | -22.8 | 8.16 | A | N | W |

Labcode

454

460

463

465

466

471

474

475

478

479

483

484

487

491

492

495

498

500

508

514

516

520

523

525

527

528

529

532

533

536

537

540

544

553

555

0.4

0.6

Display of intercomparison values

Reported results for Gross Alpha Determination in Sample 2

Table 70. Target values

Parameter	Value

0.8

Robust Mean, Bq/kg	0.319 Robust SD, Bq/kg		0.188

Activity concentration[ Bq/kg ]

−0.2

0.0

0.2

377

252

453

532

68

104

268

457

278

355

283

383

209

361

503

215

193

122

53

58

119

50

34

223

295

315

333

113

408

488

418

428

401

97

440

441

483

52

133

219

314

429

19

128

95

55

170

344

81

105

276

370

311

371

302

516

32

133

410

28

180

402

506

15

364

508

514

72

60

231

528

63

49

208

270

234

91

90

187

230

242

392

420

537

65

470

421

527

166

291

232

21

38

535

161

191

454

5

319

299

443

7

31

207

300

205

290

140

20

412

114

66

6

156

281

307

427

404

368

136

376

398

279

292

303

152

42

286

326

379

482

433

313

175

261

397

200

382

341

446

256

496

163

110

253

220

395

40

312

523

466

507

309

492

179

495

227

405

69

348

381

346

435

73

447

79

181

393

347

287

214

323

255

310

445

159

533

165

266

75

71

94

244

78

530

64

338

491

277

13

Labcode

Figure 70. Reported results and evaluation parameters for Gross Alpha in Sample 2.

Table 71. IC Evaluation results for Gross Alpha in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 5 | 0.285 | 0.048 | -0.18 | A |
| 6 | 0.313 | 0.062 | -0.03 | A |
| 7 | 0.293 | 0.051 | -0.14 | A |
| 13 | 98 | 6 | 519.58 | N |
| 15 | 0.23 | 0.06 | -0.47 | A |
| 19 | 0.183 | 0.037 | -0.72 | A |
| 20 | 0.31 | 0.01 | -0.05 | A |
| 21 | 0.28 | 0.07 | -0.21 | A |
| 28 | 0.22 | 0.071 | -0.53 | A |
| 31 | 0.295 | 0.0397 | -0.13 | A |
| 32 | 0.21 | 0.03 | -0.58 | A |
| 34 | 0.12 | 0.04 | -1.06 | A |
| 38 | 0.28 | 0.037 | -0.21 | A |
| 40 | 0.48 | 0.05 | 0.86 | A |
| 42 | 0.358 | 0.113 | 0.21 | A |
| 49 | 0.25 | 0.05 | -0.37 | A |
| 50 | 0.11 | 0.04 | -1.11 | A |
| 52 | 0.18 | 0.0095 | -0.74 | A |
| 53 | 0.106 | 0.029 | -1.13 | A |
| 55 | 0.191 | 0.048 | -0.68 | A |
| 58 | 0.108 | 0.018 | -1.12 | A |
| 60 | 0.234 | 0.037 | -0.45 | A |
| 63 | 0.247 | 0.018 | -0.38 | A |
| 64 | 21.06 | 0.1 | 110.32 | N |

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 65 | 0.27 | 0.07 | -0.26 | A |
| 66 | 0.312 | 0.01243 | -0.04 | A |
| 68 | 0.073 | 0.027 | -1.31 | A |
| 69 | 0.55 | 0.03 | 1.23 | A |
| 71 | 3.48 | 0.17 | 16.81 | N |
| 72 | 0.231 | 0.062 | -0.47 | A |
| 73 | 0.64 | 0.06 | 1.71 | A |
| 75 | 3.32 | 0.45 | 15.96 | N |
| 78 | 8.74 | 0.18 | 44.79 | N |
| 79 | 0.7 | 0.3 | 2.03 | Q |
| 81 | 0.2 | 0.066 | -0.63 | A |
| 90 | 0.26 | 0.03 | -0.31 | A |
| 91 | 0.258 | 0.03 | -0.32 | A |
| 94 | 4.3 | 0.27 | 21.18 | N |
| 95 | 0.19 | 0.05 | -0.69 | A |
| 97 | 0.173 | 0.031 | -0.78 | A |
| 104 | 0.08 | 0.03 | -1.27 | A |
| 105 | 0.2 | 0.08 | -0.63 | A |
| 110 | 0.458 | 0.028 | 0.74 | A |
| 113 | 0.152 | 0.038 | -0.89 | A |
| 114 | 0.311 | 0.069 | -0.04 | A |
| 119 | 0.109 | 0.015 | -1.12 | A |
| 122 | 0.105 | 0.015 | -1.14 | A |
| 128 | 0.187 | 0.053 | -0.70 | A |
| 133 | 0.21 | 0.01 | -0.58 | A |
| 133 | 0.18 | 0.02 | -0.74 | A |
| 136 | 0.3332 | 0.1709 | 0.08 | A |
| 140 | 0.306 | 0.042 | -0.07 | A |
| 152 | 0.351 | 0.037 | 0.17 | A |
| 156 | 0.316 | 0.044 | -0.02 | A |
| 159 | 1.077 | 0.666 | 4.03 | N |
| 161 | 0.283 | 0.02 | -0.19 | A |
| 163 | 0.451 | 0.09 | 0.70 | A |
| 165 | 1.666 | 0.15 | 7.16 | N |
| 166 | 0.278 | 0.085 | -0.22 | A |
| 170 | 0.193 | 0.018 | -0.67 | A |
| 175 | 0.376 | 0.037 | 0.30 | A |
| 179 | 0.519 | 0.103 | 1.06 | A |
| 180 | 0.22 | 0.11 | -0.53 | A |
| 181 | 0.74 | 0.15 | 2.24 | Q |
| 187 | 0.26 | 0.07 | -0.31 | A |
| 191 | 0.284 | 0.026 | -0.19 | A |
| 193 | 0.104 | 0.0021 | -1.14 | A |
| 200 | 0.4 | 0.07 | 0.43 | A |
| 205 | 0.303 | 0.234 | -0.09 | A |
| 207 | 0.3 | 0.05 | -0.10 | A |
| 208 | 0.25 | 0.04 | -0.37 | A |
| 209 | 0.1 | 0.04 | -1.16 | A |
| 214 | 0.83 | 0.06 | 2.72 | Q |
| 215 | 0.103 | 0.015 | -1.15 | A |
| 219 | 0.18 | 0.01 | -0.74 | A |
| 220 | 0.47 | 0.14 | 0.80 | A |
| 223 | 0.13 | 0.06 | -1.01 | A |
| 227 | 0.54 | 0.05 | 1.18 | A |
| 230 | 0.26 | 0.06 | -0.31 | A |
| 231 | 0.24 | 0.05 | -0.42 | A |
| 232 | 0.279 | 0.052 | -0.21 | A |
| 234 | 0.256 | 0.036 | -0.34 | A |
| 242 | 0.26 | 0.04 | -0.31 | A |
| 244 | 6.9 | 4.1 | 35.01 | N |
| 252 | 0.0288 | 0.055 | -1.54 | A |

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 253 | 0.46 | 0.12 | 0.75 | A |
| 255 | 0.869 | 0.04 | 2.93 | Q |
| 256 | 0.43 | 0.1 | 0.59 | A |
| 261 | 0.38 | 0.06 | 0.32 | A |
| 266 | 1.84 | 0.34 | 8.09 | N |
| 268 | 0.084 | 0.014 | -1.25 | A |
| 270 | 0.25 | 0.04 | -0.37 | A |
| 276 | 0.2 | 0.03 | -0.63 | A |
| 277 | 90.5 | 25 | 479.69 | N |
| 278 | 0.0855 | 0.0095 | -1.24 | A |
| 279 | 0.35 | 0.03 | 0.16 | A |
| 281 | 0.32 | 0.08 | 0.01 | A |
| 283 | 0.088 | 0.01 | -1.23 | A |
| 286 | 0.36 | 0.09 | 0.22 | A |
| 287 | 0.81 | 0.07 | 2.61 | Q |
| 290 | 0.304 | 0.477 | -0.08 | A |
| 291 | 0.278 | 0.044 | -0.22 | A |
| 292 | 0.35 | 0.05 | 0.16 | A |
| 295 | 0.137 | 0.022 | -0.97 | A |
| 299 | 0.29 | 0.04 | -0.15 | A |
| 300 | 0.3 | 0.02 | -0.10 | A |
| 302 | 0.205 | 0.03 | -0.61 | A |
| 303 | 0.35 | 0.11 | 0.16 | A |
| 307 | 0.32 | 0.03 | 0.01 | A |
| 309 | 0.51 | 0.04 | 1.02 | A |
| 310 | 1.01 | 0.13 | 3.68 | N |
| 311 | 0.202 | 0.04 | -0.62 | A |
| 312 | 0.48 | 0.05 | 0.86 | A |
| 313 | 0.373 | 0.17 | 0.29 | A |
| 314 | 0.18 | 0.01 | -0.74 | A |
| 315 | 0.14 | 0.039 | -0.95 | A |
| 319 | 0.289 | 0.049 | -0.16 | A |
| 323 | 0.85 | 0.21 | 2.82 | Q |
| 326 | 0.36 | 0.05 | 0.22 | A |
| 333 | 0.149 | 0.041 | -0.90 | A |
| 338 | 41.15 | 5.53 | 217.19 | N |
| 341 | 0.426 | 0.031 | 0.57 | A |
| 344 | 0.195 | 0.051 | -0.66 | A |
| 346 | 0.586 | 0.105 | 1.42 | A |
| 347 | 0.75 | 0.2 | 2.29 | Q |
| 348 | 0.56 | 0.13 | 1.28 | A |
| 355 | 0.087 | 0.02 | -1.23 | A |
| 361 | 0.1 | 0.1 | -1.16 | A |
| 364 | 0.23 | 0.023 | -0.47 | A |
| 368 | 0.332 | 0.05 | 0.07 | A |
| 370 | 0.2 | 0.05 | -0.63 | A |
| 371 | 0.203 | 0.037 | -0.62 | A |
| 376 | 0.34 | 0.02 | 0.11 | A |
| 377 | 0.000015 | 0 | -1.70 | A |
| 379 | 0.36 | 0.072 | 0.22 | A |
| 381 | 0.57 | 0.06 | 1.34 | A |
| 382 | 0.42 | 0.04 | 0.54 | A |
| 383 | 0.09 | 0.02 | -1.22 | A |
| 392 | 0.26 | 0.06 | -0.31 | A |
| 393 | 0.74 | 0.15 | 2.24 | Q |
| 395 | 0.47 | 0.05 | 0.80 | A |
| 397 | 0.3804 | 0.0479 | 0.33 | A |
| 398 | 0.34 | 0.08 | 0.11 | A |
| 401 | 0.172 | 0.063 | -0.78 | A |
| 402 | 0.22 | 0.03 | -0.53 | A |
| 404 | 0.33 | 0.26 | 0.06 | A |

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 405 | 0.54 | 0.03 | 1.18 | A |
| 408 | 0.1576 | 0.014 | -0.86 | A |
| 410 | 0.21 | 0.05 | -0.58 | A |
| 412 | 0.31 | 0.21 | -0.05 | A |
| 418 | 0.168 | 0.03 | -0.80 | A |
| 420 | 0.26 | 0.085 | -0.31 | A |
| 421 | 0.274 | 0.032 | -0.24 | A |
| 427 | 0.328333 | 0.016667 | 0.05 | A |
| 428 | 0.17 | 0.07 | -0.79 | A |
| 429 | 0.181 | 0.0153 | -0.73 | A |
| 433 | 0.37 | 0.25 | 0.27 | A |
| 435 | 0.6 | 0.1 | 1.49 | A |
| 440 | 0.1735 | 0.0531 | -0.77 | A |
| 441 | 0.174 | 0.048 | -0.77 | A |
| 443 | 0.29 | 0.04 | -0.15 | A |
| 445 | 1.03 | 0.15 | 3.78 | N |
| 446 | 0.426 | 0.058 | 0.57 | A |
| 447 | 0.641439 | 0.00063 | 1.72 | A |
| 453 | 0.05 | 33.78 | -1.43 | A |
| 454 | 0.284 | 0.036 | -0.19 | A |
| 457 | 0.085 | 0.00059 | -1.24 | A |
| 466 | 0.489 | 0.0564 | 0.90 | A |
| 470 | 0.27 | 0.05 | -0.26 | A |
| 482 | 0.36 | 0.08 | 0.22 | A |
| 483 | 0.174 | 0.048 | -0.77 | A |
| 488 | 0.16 | 0.01 | -0.85 | A |
| 491 | 84.368 | 3.205 | 447.07 | N |
| 492 | 0.51 | 0.035 | 1.02 | A |
| 495 | 0.5369 | 0.0669 | 1.16 | A |
| 496 | 0.44 | 0.02 | 0.64 | A |
| 503 | 0.1 | 0.1 | -1.16 | A |
| 506 | 0.22 | 0.065 | -0.53 | A |
| 507 | 0.497 | 0.095 | 0.95 | A |
| 508 | 0.23 | 0.04 | -0.47 | A |
| 514 | 0.23 | 0.043 | -0.47 | A |
| 516 | 0.206 | 0.059 | -0.60 | A |
| 523 | 0.48 | 0.1 | 0.86 | A |
| 527 | 0.275 | 0.062 | -0.23 | A |
| 528 | 0.24 | 0.03 | -0.42 | A |
| 530 | 13.729 | 0.211 | 71.33 | N |
| 532 | 0.07 | 0.017 | -1.32 | A |
| 533 | 1.43 | 0.33 | 5.91 | N |
| 535 | 0.28 | 0.07 | -0.21 | A |
| 537 | 0.26 | 0.03 | -0.31 | A |

0.5

Reported results for Gross Beta Determination in Sample 2

Table 72. Target values

Parameter	Value

1.0

Robust Mean, Bq/kg	0.417 Robust SD, Bq/kg		0.294

Activity concentration[ Bq/kg ]

−0.5

0.0

377

220

242

532

369

537

371

446

95

136

315

133

412

114

453

429

75

223

66

58

32

312

133

286

215

302

311

270

122

333

53

97

488

503

193

454

119

19

63

208

364

440

15

231

348

234

516

90

428

457

496

507

6

5

355

368

7

9

38

91

42

21

156

410

506

527

264

113

72

344

207

250

376

420

492

523

295

421

81

191

433

397

466

281

501

163

232

319

187

279

300

323

402

166

105

386

68

205

65

140

408

341

52

535

278

34

230

276

255

252

73

370

508

128

40

382

443

401

495

405

309

361

383

514

170

441

483

292

283

200

219

104

261

482

290

50

266

209

398

427

404

175

214

379

253

227

69

346

470

110

381

395

313

31

55

314

244

287

533

447

94

79

165

310

445

179

268

71

159

498

491

338

78

64

530

20

13

Labcode

Figure 71. Reported results and evaluation parameters for Gross Beta in Sample 2.

Table 73. IC Evaluation results for Gross Beta in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 5 | 0.234 | 0.023 | -0.62 | A |
| 6 | 0.232 | 0.028 | -0.63 | A |
| 7 | 0.24 | 0.022 | -0.60 | A |
| 9 | 0.24 | 0.027 | -0.60 | A |
| 13 | 521 | 31 | 1770.69 | N |
| 15 | 0.22 | 0.04 | -0.67 | A |
| 19 | 0.21 | 0.051 | -0.70 | A |
| 20 | 222.1 | 20.67 | 754.02 | N |
| 21 | 0.25 | 0.03 | -0.57 | A |
| 31 | 1.45 | 0.303 | 3.51 | N |
| 32 | 0.175 | 0.051 | -0.82 | A |
| 34 | 0.39 | 0.12 | -0.09 | A |
| 38 | 0.24 | 0.035 | -0.60 | A |
| 40 | 0.46 | 0.08 | 0.15 | A |
| 42 | 0.248 | 0.043 | -0.57 | A |
| 50 | 0.67 | 0.13 | 0.86 | A |
| 52 | 0.35 | 0.026 | -0.23 | A |
| 53 | 0.199 | 0.033 | -0.74 | A |
| 55 | 1.49 | 0.09 | 3.65 | N |
| 58 | 0.174 | 0.024 | -0.83 | A |
| 63 | 0.21 | 0.013 | -0.70 | A |
| 64 | 84.3 | 0.3 | 285.32 | N |
| 65 | 0.32 | 0.06 | -0.33 | A |
| 66 | 0.1717 | 0.0128 | -0.83 | A |
| 68 | 0.315 | 0.051 | -0.35 | A |
| 69 | 0.83 | 0.06 | 1.40 | A |

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 71 | 11.6 | 0.5 | 38.04 | N |
| 72 | 0.257 | 0.059 | -0.54 | A |
| 73 | 0.43 | 0.07 | 0.04 | A |
| 75 | 0.17 | 0.02 | -0.84 | A |
| 78 | 69.7 | 4.36 | 235.66 | N |
| 79 | 3.6 | 1.4 | 10.83 | N |
| 81 | 0.274 | 0.046 | -0.49 | A |
| 90 | 0.23 | 0.025 | -0.64 | A |
| 91 | 0.24 | 0.023 | -0.60 | A |
| 94 | 3.25 | 0.38 | 9.64 | N |
| 95 | 0.12 | 0.04 | -1.01 | A |
| 97 | 0.2 | 0.079 | -0.74 | A |
| 104 | 0.65 | 0.13 | 0.79 | A |
| 105 | 0.31 | 0.06 | -0.36 | A |
| 110 | 1.01 | 0.18 | 2.02 | Q |
| 113 | 0.255 | 0.038 | -0.55 | A |
| 114 | 0.158 | 0.035 | -0.88 | A |
| 119 | 0.207 | 0.03 | -0.71 | A |
| 122 | 0.197 | 0.029 | -0.75 | A |
| 128 | 0.441 | 0.06 | 0.08 | A |
| 133 | 0.14 | 0.01 | -0.94 | A |
| 133 | 0.19 | 0.02 | -0.77 | A |
| 136 | 0.1338 | 0.0679 | -0.96 | A |
| 140 | 0.32 | 0.091 | -0.33 | A |
| 156 | 0.25 | 0.04 | -0.57 | A |
| 159 | 21.6 | 2.21 | 72.05 | N |
| 163 | 0.291 | 0.055 | -0.43 | A |
| 165 | 4.521 | 0.13 | 13.96 | N |
| 166 | 0.308 | 0.11 | -0.37 | A |
| 170 | 0.53 | 0.05 | 0.38 | A |
| 175 | 0.74 | 0.16 | 1.10 | A |
| 179 | 7.572 | 0.396 | 24.34 | N |
| 187 | 0.3 | 0.06 | -0.40 | A |
| 191 | 0.277 | 0.067 | -0.48 | A |
| 193 | 0.201 | 0.004 | -0.73 | A |
| 200 | 0.64 | 0.11 | 0.76 | A |
| 205 | 0.319 | 0.105 | -0.33 | A |
| 207 | 0.26 | 0.09 | -0.53 | A |
| 208 | 0.21 | 0.02 | -0.70 | A |
| 209 | 0.7 | 0.14 | 0.96 | A |
| 214 | 0.75 | 0.06 | 1.13 | A |
| 215 | 0.191 | 0.028 | -0.77 | A |
| 219 | 0.64 | 0.06 | 0.76 | A |
| 220 | 0.06 | 0.02 | -1.21 | A |
| 223 | 0.17 | 0.06 | -0.84 | A |
| 227 | 0.82 | 0.23 | 1.37 | A |
| 230 | 0.41 | 0.06 | -0.02 | A |
| 231 | 0.22 | 0.04 | -0.67 | A |
| 232 | 0.296 | 0.042 | -0.41 | A |
| 234 | 0.222 | 0.02 | -0.66 | A |
| 242 | 0.06 | 0.01 | -1.21 | A |
| 244 | 1.7 | 1.9 | 4.36 | N |
| 250 | 0.26 | 0.04 | -0.53 | A |
| 252 | 0.428 | 0.02 | 0.04 | A |
| 253 | 0.78 | 0.56 | 1.23 | A |
| 255 | 0.417 | 0.019 | 0.00 | A |
| 261 | 0.65 | 0.09 | 0.79 | A |
| 264 | 0.253 | 0.015 | -0.56 | A |
| 266 | 0.67 | 0.2 | 0.86 | A |
| 268 | 11.3 | 0.57 | 37.02 | N |
| 270 | 0.195 | 0.029 | -0.76 | A |

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 276 | 0.41 | 0.1 | -0.02 | A |
| 278 | 0.368 | 0.03 | -0.17 | A |
| 279 | 0.3 | 0.08 | -0.40 | A |
| 281 | 0.283 | 0.102 | -0.46 | A |
| 283 | 0.6 | 0.03 | 0.62 | A |
| 286 | 0.19 | 0.1 | -0.77 | A |
| 287 | 1.7 | 0.29 | 4.36 | N |
| 290 | 0.653 | 0.052 | 0.80 | A |
| 292 | 0.58 | 0.09 | 0.55 | A |
| 295 | 0.262 | 0.05 | -0.53 | A |
| 300 | 0.3 | 0.04 | -0.40 | A |
| 302 | 0.193 | 0.03 | -0.76 | A |
| 309 | 0.5 | 0.05 | 0.28 | A |
| 310 | 6.1 | 0.69 | 19.33 | N |
| 311 | 0.193 | 0.0282 | -0.76 | A |
| 312 | 0.18 | 0.03 | -0.81 | A |
| 313 | 1.36 | 0.65 | 3.21 | N |
| 314 | 1.65 | 0.07 | 4.19 | N |
| 315 | 0.137 | 0.061 | -0.95 | A |
| 319 | 0.299 | 0.037 | -0.40 | A |
| 323 | 0.3 | 0.07 | -0.40 | A |
| 333 | 0.197 | 0.072 | -0.75 | A |
| 338 | 39.72 | 9.53 | 133.68 | N |
| 341 | 0.344 | 0.023 | -0.25 | A |
| 344 | 0.258 | 0.063 | -0.54 | A |
| 346 | 0.863 | 0.109 | 1.52 | A |
| 348 | 0.22 | 0.05 | -0.67 | A |
| 355 | 0.238 | 0.099 | -0.61 | A |
| 361 | 0.5 | 0.5 | 0.28 | A |
| 364 | 0.214 | 0.013 | -0.69 | A |
| 368 | 0.239 | 0.043 | -0.61 | A |
| 369 | 0.09 | 0.03 | -1.11 | A |
| 370 | 0.44 | 0.06 | 0.08 | A |
| 371 | 0.116 | 0.017 | -1.02 | A |
| 376 | 0.26 | 0.01 | -0.53 | A |
| 377 | 0.000025 | 0.000003 | -1.42 | A |
| 379 | 0.75 | 0.16 | 1.13 | A |
| 381 | 1.1 | 0.11 | 2.32 | Q |
| 382 | 0.46 | 0.06 | 0.15 | A |
| 383 | 0.52 | 0.08 | 0.35 | A |
| 386 | 0.31 | 0.03 | -0.36 | A |
| 395 | 1.13 | 0.13 | 2.43 | Q |
| 397 | 0.2806 | 0.0149 | -0.46 | A |
| 398 | 0.7 | 0.16 | 0.96 | A |
| 401 | 0.461 | 0.108 | 0.15 | A |
| 402 | 0.3 | 0.04 | -0.40 | A |
| 404 | 0.72 | 0.32 | 1.03 | A |
| 405 | 0.49 | 0.04 | 0.25 | A |
| 408 | 0.3266 | 0.05 | -0.31 | A |
| 410 | 0.25 | 0.06 | -0.57 | A |
| 412 | 0.14 | 0.12 | -0.94 | A |
| 420 | 0.26 | 0.06 | -0.53 | A |
| 421 | 0.27 | 0.037 | -0.50 | A |
| 427 | 0.7125 | 0.070833 | 1.01 | A |
| 428 | 0.23 | 0.07 | -0.64 | A |
| 429 | 0.169 | 0.0162 | -0.84 | A |
| 433 | 0.28 | 0.16 | -0.47 | A |
| 440 | 0.2193 | 0.0196 | -0.67 | A |
| 441 | 0.561 | 0.12 | 0.49 | A |
| 443 | 0.46 | 0.04 | 0.15 | A |
| 445 | 6.13 | 0.83 | 19.43 | N |

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 446 | 0.1186 | 0.0104 | -1.01 | A |
| 447 | 2.88529 | 0.0169 | 8.40 | N |
| 453 | 0.16 | 18.37 | -0.87 | A |
| 454 | 0.203 | 0.014 | -0.73 | A |
| 457 | 0.23 | 0.0091 | -0.64 | A |
| 466 | 0.281 | 0.0243 | -0.46 | A |
| 470 | 0.97 | 0.1 | 1.88 | A |
| 482 | 0.65 | 0.09 | 0.79 | A |
| 483 | 0.561 | 0.12 | 0.49 | A |
| 488 | 0.2 | 0.01 | -0.74 | A |
| 491 | 31.53 | 3.178 | 105.83 | N |
| 492 | 0.26 | 0.015 | -0.53 | A |
| 495 | 0.4611 | 0.0497 | 0.15 | A |
| 496 | 0.23 | 0.03 | -0.64 | A |
| 498 | 25.7 | 1.66 | 86.00 | N |
| 501 | 0.287 | 0.047 | -0.44 | A |
| 503 | 0.2 | 0.2 | -0.74 | A |
| 506 | 0.25 | 0.08 | -0.57 | A |
| 507 | 0.23 | 0.032 | -0.64 | A |
| 508 | 0.44 | 0.08 | 0.08 | A |
| 514 | 0.52 | 0.08 | 0.35 | A |
| 516 | 0.229 | 0.042 | -0.64 | A |
| 523 | 0.26 | 0.039 | -0.53 | A |
| 527 | 0.252 | 0.043 | -0.56 | A |
| 530 | 157.222 | 1.046 | 533.35 | N |
| 532 | 0.082 | 0.011 | -1.14 | A |
| 533 | 1.89 | 0.29 | 5.01 | N |
| 535 | 0.35 | 0.03 | -0.23 | A |
| 537 | 0.11 | 0.04 | -1.04 | A |

1

2

Activity concentration[ Bq/kg ]

−1

0

434

368

533

87

222

255

443

410

433

383

484

347

181

393

377

228

258

52

71

491

165

58

Labcode

Figure 72. Reported results and evaluation parameters for Ra-228 in Sample 2.

Table 75. IC Evaluation results for Ra-228 in Sample 2, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 52 | 1.78 | 0.099 | 1.60 | A |
| 58 |  | 0.5 |  |  |
| 71 | 3.6 | 0.18 | 4.19 | N |
| 87 | 0.05 | 0.04 | -0.86 | A |
| 165 | 34.8 | 2.3 | 48.57 | N |
| 181 | 0.65 | 0.12 | -0.01 | A |
| 222 | 0.075 | 0.02 | -0.83 | A |
| 228 | 0.85 | 0.1 | 0.28 | A |
| 255 | 0.083 | 0.012 | -0.82 | A |
| 258 | 1.1 | 0.1 | 0.63 | A |
| 347 | 0.64 | 0.16 | -0.02 | A |
| 368 | 0.023 | 0.008 | -0.90 | A |
| 377 | 0.71 | 0.0435 | 0.08 | A |
| 383 | 0.3 | 0.13 | -0.51 | A |
| 393 | 0.68 | 0.12 | 0.03 | A |
| 410 | 0.18 | 0.03 | -0.68 | A |
| 433 | 0.245 | 0.092 | -0.58 | A |
| 434 | 0 | 0 | -0.93 | A |
| 443 | 0.153 | 0.015 | -0.72 | A |
| 484 | 0.5 | 0.3 | -0.22 | A |
| 491 | 5.17 | 0.244 | 6.42 | N |
| 533 | 0.035 | 0.018 | -0.88 | A |

0.5

1.0

1.5

Activity concentration[ Bq/kg ]

−0.5

0.0

280

19

222

152

344

393

110

65

187

347

537

509

150

391

433

158

357

242

128

370

484

342

536

355

145

322

223

82

179

442

Labcode

Figure 73. Reported results and evaluation parameters for Cs-137 in Sample 4.

Table 77. IC Evaluation results for Cs-137 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 19 | 0.159 | 0.035 | -0.89 | A |
| 65 | 0.23 | 0.05 | -0.71 | A |
| 82 | 2.57 | 0.46 | 5.42 | N |
| 110 | 0.226 | 0.039 | -0.72 | A |
| 128 | 0.45 | 0.04 | -0.13 | A |
| 145 | 0.93 | 0.16 | 1.13 | A |
| 150 | 0.3 | 0.1 | -0.52 | A |
| 152 | 0.22 | 0.02 | -0.73 | A |
| 158 | 0.4 | 0.2 | -0.26 | A |
| 179 | 2.96 | 1.05 | 6.44 | N |
| 187 | 0.23 | 0.05 | -0.71 | A |
| 222 | 0.171 | 0.038 | -0.86 | A |
| 223 | 1.3 | 0.2 | 2.09 | Q |
| 242 | 0.42 | 0.03 | -0.21 | A |
| 280 | 0.059 | 0.0057 | -1.15 | A |
| 322 | 1.09 | 0.83 | 1.54 | A |
| 342 | 0.73 | 0.31 | 0.60 | A |
| 344 | 0.22 | 0.05 | -0.73 | A |
| 347 | 0.25 | 0.04 | -0.65 | A |
| 355 | 0.85 | 0.08 | 0.92 | A |
| 357 | 0.41 | 0.12 | -0.24 | A |
| 370 | 0.5 | 0.1 | 0.00 | A |
| 391 | 0.32 | 0.07 | -0.47 | A |
| 393 | 0.22 | 0.04 | -0.73 | A |
| 433 | 0.35 | 0.31 | -0.39 | A |
| 442 | 15.61951 | 1.27171 | 39.58 | N |

Table 77. IC Evaluation results for Cs-137 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 484 | 0.6 | 0.29 | 0.26 | A |
| 509 | 0.29 | 0.04 | -0.55 | A |
| 536 | 0.84 | 0.23 | 0.89 | A |
| 537 | 0.25 | 0.01 | -0.65 | A |

5

10

15

Activity concentration[ Bq/kg ]

−5

0

280

150

64

128

370

124

84

249

217

228

258

377

133

211

179

333

491

Labcode

Figure 74. Reported results and evaluation parameters for Ac-228 in Sample 4.

Table 79. IC Evaluation results for Ac-228 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 64 | 1.2 | 0.3 | -0.88 | A |
| 84 | 3.07 | 0.96 | -0.43 | A |
| 124 | 2.9 | 1.1 | -0.47 | A |
| 128 | 1.41 | 0.18 | -0.83 | A |
| 133 | 7.97 | 0.6 | 0.74 | A |
| 150 | 0.79 | 0.17 | -0.97 | A |
| 179 | 9.73 | 4.79 | 1.17 | A |
| 211 | 9.06 | 0.6 | 1.00 | A |
| 217 | 3.9 | 0.9 | -0.23 | A |
| 228 | 4.3 | 0.5 | -0.13 | A |
| 249 | 3.369288 | 0.095688 | -0.36 | A |
| 258 | 5.2 | 1 | 0.08 | A |
| 280 | 0.081 | 0.021 | -1.14 | A |
| 333 | 25.181 | 4.95 | 4.86 | N |
| 370 | 1.5 | 0.3 | -0.80 | A |
| 377 | 5.858558 | 5.1582 | 0.24 | A |
| 491 | 36.493 | 2.445 | 7.57 | N |

5

10

15

Activity concentration[ Bq/kg ]

−5

0

434

276

452

150

537

217

344

414

341

462

319

408

25

165

218

156

249

128

370

338

131

34

64

444

198

228

287

155

124

496

71

333

417

84

384

211

226

179

52

94

491

Labcode

Figure 75. Reported results and evaluation parameters for Bi-214 in Sample 4.

Table 81. IC Evaluation results for Bi-214 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 25 | 2.2 | 0.18 | -0.60 | A |
| 34 | 3.2 | 0.5 | -0.32 | A |
| 52 | 21.61 | 2.01 | 4.75 | N |
| 64 | 3.3 | 0.6 | -0.29 | A |
| 71 | 6.61 | 0.33 | 0.62 | A |
| 84 | 9.24 | 3.79 | 1.34 | A |
| 94 | 39.5 | 4.15 | 9.68 | N |
| 124 | 5.1 | 1 | 0.20 | A |
| 128 | 2.56 | 0.25 | -0.50 | A |
| 131 | 2.691456 | 0.590264 | -0.46 | A |
| 150 | 1.24 | 0.18 | -0.86 | A |
| 155 | 4.66 | 0.64 | 0.08 | A |
| 156 | 2.35 | 0.22 | -0.56 | A |
| 165 | 2.2 | 0.066 | -0.60 | A |
| 179 | 18.92 | 2.54 | 4.01 | N |
| 198 | 3.43 | 0.68 | -0.26 | A |
| 211 | 11 | 0.48 | 1.83 | A |
| 217 | 1.6 | 0.23 | -0.76 | A |
| 218 | 2.29652 | 0.35504 | -0.57 | A |
| 226 | 16.78 | 1.25 | 3.42 | N |
| 228 | 3.7 | 0.4 | -0.18 | A |
| 249 | 2.495709 | 0.074257 | -0.52 | A |
| 276 | 0.86 | 0.17 | -0.97 | A |
| 287 | 4.1 | 1.3 | -0.07 | A |
| 319 | 2.06 | 0.15 | -0.64 | A |
| 333 | 8.299 | 1.97 | 1.08 | A |

Table 81. IC Evaluation results for Bi-214 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 338 | 2.68 | 0.29 | -0.47 | A |
| 341 | 1.94 | 0.15 | -0.67 | A |
| 344 | 1.68 | 0.24 | -0.74 | A |
| 370 | 2.6 | 0.5 | -0.49 | A |
| 384 | 10.5 | 1.3 | 1.69 | A |
| 408 | 2.16 | 0.19 | -0.61 | A |
| 414 | 1.69 | 0.18 | -0.74 | A |
| 417 | 8.8 | 1.8 | 1.22 | A |
| 434 | 0 | 0 | -1.20 | A |
| 444 | 3.37 | 0.8 | -0.28 | A |
| 452 | 0.9894 | 0.207774 | -0.93 | A |
| 462 | 2 | 1 | -0.65 | A |
| 491 | 58.345 | 3.505 | 14.87 | N |
| 496 | 5.87 | 1.03 | 0.41 | A |
| 537 | 1.3 | 0.09 | -0.85 | A |

4

6

Reported results for Pb-214 determination in Sample 4

Table 82. Target values

Parameter	Value

8

10

Robust Mean, Bq/kg	3.50 Robust SD, Bq/kg		2.33

Activity concentration[ Bq/kg ]

−4

−2

0

2

276

537

452

150

414

34

217

175

344

218

341

462

408

319

25

165

228

69

156

128

370

437

293

249

338

464

444

64

198

555

131

155

287

124

496

71

333

84

211

384

417

226

491

52

179

94

Labcode

Figure 76. Reported results and evaluation parameters for Pb-214 in Sample 4.

Table 83. IC Evaluation results for Pb-214 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 25 | 2.18 | 0.22 | -0.57 | A |
| 34 | 1.6 | 0.3 | -0.82 | A |
| 52 | 21.61 | 2.01 | 7.77 | N |
| 64 | 3.3 | 0.5 | -0.09 | A |
| 69 | 2.4 | 0.6 | -0.47 | A |
| 71 | 5.66 | 0.28 | 0.93 | A |
| 84 | 7.24 | 2.34 | 1.61 | A |
| 94 | 39.5 | 4.15 | 15.45 | N |
| 124 | 4.6 | 0.5 | 0.47 | A |
| 128 | 2.45 | 0.24 | -0.45 | A |
| 131 | 3.785761 | 0.335417 | 0.12 | A |
| 150 | 1.52 | 0.26 | -0.85 | A |
| 155 | 4.03 | 0.44 | 0.23 | A |
| 156 | 2.42 | 0.37 | -0.46 | A |
| 165 | 2.2 | 0.066 | -0.56 | A |
| 175 | 1.62 | 0.19 | -0.81 | A |
| 179 | 33.39 | 3.72 | 12.83 | N |
| 198 | 3.43 | 0.37 | -0.03 | A |
| 211 | 7.61 | 0.4 | 1.76 | A |
| 217 | 1.6 | 0.23 | -0.82 | A |
| 218 | 1.9196 | 0.28843 | -0.68 | A |
| 226 | 14.83 | 1.32 | 4.86 | N |
| 228 | 2.3 | 0.7 | -0.52 | A |
| 249 | 2.601051 | 0.071483 | -0.39 | A |
| 276 | 0.83 | 0.13 | -1.15 | A |
| 287 | 4.5 | 1.4 | 0.43 | A |

Table 83. IC Evaluation results for Pb-214 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 293 | 2.6 | 0.6 | -0.39 | A |
| 319 | 2.06 | 0.15 | -0.62 | A |
| 333 | 6.071 | 2.093 | 1.10 | A |
| 338 | 2.68 | 0.29 | -0.35 | A |
| 341 | 1.94 | 0.15 | -0.67 | A |
| 344 | 1.68 | 0.24 | -0.78 | A |
| 370 | 2.5 | 0.5 | -0.43 | A |
| 384 | 8.6 | 1.1 | 2.19 | Q |
| 408 | 2.05 | 0.3 | -0.62 | A |
| 414 | 1.55 | 0.16 | -0.84 | A |
| 417 | 8.7 | 1.7 | 2.23 | Q |
| 437 | 2.5 | 0.4 | -0.43 | A |
| 444 | 2.83 | 0.7 | -0.29 | A |
| 452 | 1.384 | 0.2076 | -0.91 | A |
| 462 | 2 | 0.18 | -0.64 | A |
| 464 | 2.81 | 0.44 | -0.30 | A |
| 491 | 18.106 | 1.795 | 6.27 | N |
| 496 | 5.41 | 0.63 | 0.82 | A |
| 537 | 1.28 | 0.09 | -0.95 | A |
| 555 | 3.6 | 1.4 | 0.04 | A |

5

Reported results for Ra-226 determination in Sample 4

Table 84. Target values

Parameter	Value

10

15

Robust Mean, Bq/kg	4.40 Robust SD, Bq/kg		3.59

Activity concentration[ Bq/kg ]

−5

0

537

280

217

414

344

57

147

523

341

553

319

218

408

529

15

165

528

110

509

181

207

249

159

338

411

475

525

198

533

276

214

462

322

124

417

84

484

34

401

520

52

397

491

94

555

179

Labcode

Figure 77. Reported results and evaluation parameters for Ra-226 in Sample 4.

Table 85. IC Evaluation results for Ra-226 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 15 | 2.2 | 0.65 | -0.61 | A |
| 34 | 10.6 | 7 | 1.73 | A |
| 52 | 21.61 | 2.01 | 4.79 | N |
| 57 | 1.7 | 0.5 | -0.75 | A |
| 84 | 8.43 | 3.35 | 1.12 | A |
| 94 | 39.5 | 4.15 | 9.78 | N |
| 110 | 2.32 | 0.21 | -0.58 | A |
| 124 | 4.8 | 0.6 | 0.11 | A |
| 147 | 1.7046 | 0.1971 | -0.75 | A |
| 159 | 2.62 | 0.71 | -0.50 | A |
| 165 | 2.2 | 0.066 | -0.61 | A |
| 179 | 124.83 | 18.33 | 33.55 | N |
| 181 | 2.42 | 0.35 | -0.55 | A |
| 198 | 3.43 | 0.37 | -0.27 | A |
| 207 | 2.45 | 0.37 | -0.54 | A |
| 214 | 4 | 0.09 | -0.11 | A |
| 217 | 1.6 | 0.23 | -0.78 | A |
| 218 | 2.10806 | 0.32174 | -0.64 | A |
| 249 | 2.54838 | 0.076967 | -0.52 | A |
| 276 | 3.6 | 1.19 | -0.22 | A |
| 280 | 1.49 | 0.22 | -0.81 | A |
| 319 | 2.06 | 0.15 | -0.65 | A |
| 322 | 4.5 | 0.67 | 0.03 | A |
| 338 | 2.68 | 0.29 | -0.48 | A |
| 341 | 1.94 | 0.15 | -0.69 | A |
| 344 | 1.68 | 0.24 | -0.76 | A |

Table 85. IC Evaluation results for Ra-226 in Sample 4, Rep. Value and Rep. Unc. in Bq/kg

| Labcode | Rep. Value | Rep. Unc. | Z-Score | Z-Score Evaluation |
| --- | --- | --- | --- | --- |
| 397 | 32.72 | 1.12 | 7.89 | N |
| 401 | 14.1 | 4.6 | 2.70 | Q |
| 408 | 2.12 | 0.45 | -0.64 | A |
| 411 | 2.9 | 0.2 | -0.42 | A |
| 414 | 1.61 | 0.12 | -0.78 | A |
| 417 | 8.4 | 1.6 | 1.11 | A |
| 462 | 4 | 2 | -0.11 | A |
| 475 | 3.1 | 0.2 | -0.36 | A |
| 484 | 8.7 | 3.2 | 1.20 | A |
| 491 | 38.226 | 1.597 | 9.42 | N |
| 509 | 2.33 | 0.06 | -0.58 | A |
| 520 | 14.17 | 0.99 | 2.72 | Q |
| 523 | 1.91 | 0.17 | -0.69 | A |
| 525 | 3.22 | 0.41 | -0.33 | A |
| 528 | 2.2 | 0.2 | -0.61 | A |
| 529 | 2.13 | 0.88 | -0.63 | A |
| 533 | 3.49 | 0.39 | -0.25 | A |
| 537 | 1.3 | 0.09 | -0.86 | A |
| 553 | 1.98 | 0.14 | -0.67 | A |
| 555 | 48.9 | 6.8 | 12.40 | N |

# Contributors to Evaluation and Report

I. Davies, M. Horsky, B. Seslak, A. Trinkl, I. Vukanac: Terrestrial Environmental Radiochemistry Laboratory, Division of Physical and Chemical Sciences, Department of Nuclear Sciences and Applications, International Atomic Energy Agency

# Disclaimer

This report is generated automatically and intended solely for the recipient’s personal use. While every effort is made to en- sure accuracy, we acknowledge that discrepancies may occur. Should you identify any incorrect information, please notify us promptly so we can address the issue. Your vigilance helps us maintain the highest standard of correctness.